## Table 1
Solution/Solvent | Reagent(s)
Reconstitution solution | Acetonitrile:water (20:80, v/v) with 0.1% v/v formic acid
Extraction solvent | Tert-butyl methyl ether
Injection solvent | 2-propanol
Mobile phase A | Acetonitrile:Water (5:95, v/v) with 0.1% v/v formic acid
Mobile phase B | Acetonitrile:Water (95:5, v/v) with 0.1% v/v formic acid
Standard curve | Spiked docetaxel (dissolved in ethanol) in human plasma to 0.084, 0.21, 0.524, 1.311, 3.277, 8.192, 20.48, 51.2, 128, 288, 320 ng/mL
Internal standard | Paclitaxel dissolved in ethanol–water (50:50, v/v) to 360 ng/mL
Test solution | Docetaxel at 82 ng/mL and paclitaxel at 96 ng/mL dissolved in ethanol.
Quality control samples | Docetaxel spiked in human plasma to 0.084 (LLOQ), 0.252 (Low QC), 8 (Med 1 QC), 160 (Med 2 QC) and 256 (High QC) ng/mL.

## Table 2
0 | 1 | 2
(A) | (A) | (A)
Parameters
Curtain gas (psi) | 30 | 30
Collision gas (psi) | 30 | 30
Ion spray voltage (volts) | 5500 | 5500
Temperature (°C) | 300 | 300
Ion source gas 1 (psi) | 65 | 65
Ion source gas 2 (psi) | 30 | 30
(B) | (B) | (B)
Parameters | Docetaxel | Paclitaxel
Q1 mass (amu) | 808.4 | 854.5
Q3 mass (amu) | 527.2 | 285.0
Declustering potential (volts) | 45 | 60
Entrance potential (volts) | 4 | 5
Collision energy (volts) | 13 | 14
Collision cell exit potential (volts) | 14 | 8

## Table 3
Unnamed: 0_level_0_Unnamed: 0_level_1_Unnamed: 0_level_2 | Nominal Concentration (ng/mL)_7.5_Actual Concentration Day 131 (ng/mL) | Nominal Concentration (ng/mL)_32_Actual Concentration Day 131 (ng/mL) | Nominal Concentration (ng/mL)_140_Actual Concentration Day 131 (ng/mL) | Nominal Concentration (ng/mL)_2000_Actual Concentration Day 131 (ng/mL)
| N = 6 | N = 6 | N = 6 | N = 6
Mean | 8.0 | 32.5 | 144.8 | 1984.1
Standard deviation | 0.41 | 2.9 | 9.4 | 100.5
Observed change (%) | 6.9% | 1.7% | 3.4% | 0.8%

## Table 4
Unnamed: 0_level_0_Unnamed: 0_level_1 | Nominal Concentration (ng/mL)_LLOQ | Nominal Concentration (ng/mL)_Low QC | Nominal Concentration (ng/mL)_Med 1 QC | Nominal Concentration (ng/mL)_Med 2 QC | Nominal Concentration (ng/mL)_High QC
QC | 0.084 | 0.252 | 8 | 160 | 256
| Actual Concentration (ng/mL) | Actual Concentration (ng/mL) | Actual Concentration (ng/mL) | Actual Concentration (ng/mL) | Actual Concentration (ng/mL)
| LLOQ | Low QC | Med 1 QC | Med 2 QC | High QC
N | 18 | 18 | 18 | 18 | 18
Mean | 0.087 | 0.253 | 8.0 | 158.4 | 254.6
SD | 0.011 | 0.024 | 0.64 | 12.1 | 13.8
CV | 13.0% | 9.4% | 8.1% | 7.6% | 5.4%
Accuracy | 10.3% | 7.0% | 6.9% | 6.0% | 4.4%

## Table 5
0 | 1 | 2
Concentration | 82 ng/mL | 82 ng/mL
Filtration | No | Yes
N | 3 | 3
Mean PAR | 2.9235 | 2.5736
SD PAR | 0.1185 | 0.1290
Percentage (%) |  | 88.03

## Freetext Data
Data | Value | Relevance
Average observed change after long-term storage | Between 0.8% and 6.9% | Relevant for assessing long-term stability and ensuring the reliability of the assay over extended storage duration.
Number of quality control levels analyzed for precision and accuracy | Five | Relevant for validating that the assay complies with FDA acceptance criteria for precision and accuracy.
Lowest level of quantification (LLOQ) | 0.084 ng/mL | Relevant for demonstrating improved sensitivity compared to previously reported assays, aiding the detection of low unbound docetaxel concentrations.
Unbound docetaxel concentration after IV administration | Average 0.7% | Relevant for comparing unbound drug levels between intravenous and oral formulations, which informs pharmacokinetic profiling.
Unbound docetaxel fraction after oral administration (oDox+E) | 1.3% | Relevant to assess the pharmacokinetic differences and therapeutic potential of the oral formulation.
Recovery rate of docetaxel | 88% | Relevant for evaluating efficiency and the extent of nonspecific binding that could affect quantification accuracy during ultrafiltration.
Volume of plasma used for the assay | 485 µL | Relevant for methodological details that determine the feasibility and applicability of the assay for patient samples.
Correlation coefficient of standard curves | >0.99 | Relevant to confirm the linearity and reliability of the standard curve used to calculate concentrations of unbound docetaxel.
Number of patients analyzed for unbound docetaxel concentrations | Nine | Relevant to validate assay implementation and data reliability in real-world clinical studies involving patients.
Acceptance criteria for precision and accuracy | ≤20% for LLOQ, ≤15% for QC concentrations | Relevant adherence to FDA guidelines ensuring assay performance validation.
Sensitivity improvement compared to previous assays | 0.084 ng/mL vs 0.1–0.4 ng/mL in earlier reports | Relevant for proving the assay's advancement in sensitivity for measuring unbound docetaxel levels.
Fraction unbound for IV vs oral docetaxel formulations | 0.7% vs 1.3% | Relevant to highlight differences in drug binding and therapeutic implications between the two formulations.
Patient sample size | 9 patients | The study involved nine patients with metastatic prostate cancer to gather pharmacokinetic data on unbound plasma docetaxel levels after IV and oral administration, ensuring diverse and representative pharmacokinetic profile data.
Patient population | Metastatic prostate cancer patients | Patients with metastatic prostate cancer were selected to evaluate the pharmacokinetics of docetaxel in a population likely to benefit from oncology treatments.
Blood sampling points | 24 and 23 plasma samples for IV and oral administration respectively | The study collected a comprehensive set of plasma samples to compare pharmacokinetics of IV docetaxel against oral docetaxel plus encequidar (oDox+E).
Sampling design | 7–8 plasma samples selected using optimal design for each route | The optimal sampling design allowed efficient capture of pharmacokinetic data across different time points for robust modeling.
Sampling coverage | Time course sampling post-dose | The study employs detailed time-course sample collection to assess the pharmacokinetic changes and calculate unbound drug fractions effectively.
Model type used in docetaxel pharmacokinetics | Not explicitly mentioned in freetext whether the model is 1-compartment, 2-compartment, physiologically based model, or IVIVE | The paper discusses the pharmacokinetic modeling but does not provide explicit details about the specific type of model used for unbound drug quantification, limiting definitive classification.
Docetaxel storage temperature | −60 °C | Ensures long-term stability of docetaxel over 131 days, crucial for accurate quantification in experiments.
Long-term stability change | 0.8–6.9% | Validates the assay's capacity to maintain docetaxel stability within acceptable limits over extended storage periods.
Precision criteria | ≤20% for LLOQ, ≤15% for other QC levels | Ensures reliability of unbound docetaxel quantification based on FDA guidelines.
Accuracy criteria | ≤20% variation for LLOQ, ≤15% variation for other QC concentrations | Validates the accuracy of the unbound docetaxel measurement assay.
Recovery rate | 88% | Indicates efficiency of ultracentrifugation for isolating docetaxel from plasma samples.
Retention time for docetaxel | 1.82 min | Parameter estimated from HPLC-MS/MS analysis, ensures repeatable identification of docetaxel in chromatography.
Retention time for paclitaxel | 1.84 min | Supports reliable internal standard analysis for docetaxel quantification using HPLC-MS/MS.
Observed CV for docetaxel retention time | 0.06% to 0.31% | Confirms stability and reproducibility of docetaxel measurements in chromatography methods.
Observed CV for paclitaxel retention time | 0.03% to 0.33% | Validates reproducibility of the internal standard within the developed assay.
Observed CV for peak area ratio (PAR) | 0.8% to 2.8% | Ensures consistent quantitative measurements between test runs during patient sample analysis.
Correlation coefficient for standard curves | >0.99 | Demonstrates linearity and robustness of standard curve models for back-calculating docetaxel concentrations.
Unbound fraction after oDox+E | 1.3% | Represents pharmacologically active drug concentration post-oral administration; crucial for clinical comparison with IV formulations.
Unbound fraction after IV docetaxel | 0.7% | Indicates differences in drug availability to molecular targets compared to oral formulation.
Lowest detected unbound concentration | LLOQ: 0.084 ng/mL | Parameter relevant for detecting trace unbound docetaxel levels in pharmacokinetic study samples.
Incubation temperature | 37 °C | Maintains physiological conditions for ultrafiltration, enabling accurate separation of unbound docetaxel.
Incubation duration | 1 h | Optimized for ultrafiltration efficiency during sample preparation.
Centrifugation force | 6000× g | Ensures effective separation of filtrate for downstream analysis.
Centrifugation temperature | 37 °C | Prevents degradation or binding changes in docetaxel during isolation.
Injected sample volume | 20 µL | Optimized for detection sensitivity during HPLC-MS/MS assay of unbound docetaxel.
Extraction solvent volume | 600 µL | Ensures complete extraction of docetaxel and internal standard from plasma filtrate, enhancing assay reliability.
Mobile phase flow rate | 0.5 mL/min | Parameter relevant for efficient chromatographic separation during HPLC analysis.
Unbound fraction of docetaxel in plasma | 1.3% (oDox+E), 0.7% (IV) | Demonstrates that the unbound fraction of docetaxel differs between oral and IV administration, which can influence pharmacokinetics and pharmacological activity.
Protein-binding to albumin and AAG | High binding to albumin and alpha 1-acid glycoprotein | Strong protein binding affects the unbound concentration of docetaxel, directly impacting its pharmacokinetics.
Encequidar use | Intestine-specific P-glycoprotein inhibitor increases oral docetaxel absorption | Encequidar facilitates oral absorption, overcoming bioavailability barriers and altering pharmacokinetics.
Polysorbate 80 (Tween 80) in IV formulation | Required for IV formulation but absent in oral docetaxel | Tween 80, found in the IV formulation, affects systemic distribution but is not present in oral administration, influencing pharmacokinetic profiles.
Low bioavailability of unmodified oral docetaxel | Poor bioavailability due to P-glycoprotein-mediated efflux | Oral administration without encequidar leads to insufficient systemic exposure, markedly altering pharmacokinetic outcomes.
Retention time consistency for docetaxel during HPLC-MS/MS | CVs for retention times <5%, PAR CVs <10% | Stability and reproducibility in retention times support reliable pharmacokinetic assessment.
Long-term stability of stored plasma samples | Stability change between 0.8% and 6.9% at −60°C over 131 days | Preserved sample stability ensures integrity of pharmacokinetic analysis over extended storage durations.
Recovery efficiency in ultrafiltration | 88% recovery, 12% non-specific binding | Recovery metrics affect the accurate measurement of unbound docetaxel, critical for pharmacokinetics.
Dose-dependent unbound concentration detection | Insufficient sensitivity at 75 mg/m2 dose 7 hours post-administration | Highlights that early-phase sensitivity challenges can limit detection of pharmacokinetic variations at lower doses.
Variability in unbound drug exposure | Correlation of unbound concentration with therapeutic effect and toxicity | Reinforces the importance of evaluating unbound drug levels for pharmacokinetic modeling and clinical relevance.