## Freetext Data
Data | Value | Relevance
Objective | To develop a robust RP-HPLC method for simultaneous determination of Aspirin and Apixaban | Relevant as it outlines the study's primary purpose related to pharmaceutical analysis.
Column | Phenomenex Luna C18 | Critical for understanding the separation process relevant to chromatographic methodology development.
HPLC system | Shimadzu LC20AT | Provides details on the equipment used, essential for method reproducibility and validation.
Mobile Phase | 40:60 Acetonitrile: phosphate buffer (pH 4) modified by O-Phosphoric Acid | Identifies the solvent system, necessary for achieving effective compound separation.
Detection Wavelength | 227 nm | Specifies the detection parameter, critical for identifying ASP and API during testing.
Flow Rate | 1 ml/min | Essential for understanding the operational conditions influencing chromatographic resolution.
Column Temperature | 30°C | Indicates experimental conditions affecting analyte retention times.
Retention Times | Aspirin: 5.37 min, Apixaban: 7.10 min | Key for determining the separation efficiency of the developed method.
Linearity Range | Aspirin: 50-300 μg/mL; Apixaban: 5-15 μg/mL | Demonstrates the method’s capability to analyze different concentrations within pharmaceutical formulations.
Recovery Percentage | Aspirin: 90.02%-101%; Apixaban: 98.18%-101.18% | Validates the method’s accuracy in determining the compounds.
LOD and LOQ | Aspirin: 0.84 μg/mL and 2.55 μg/mL; Apixaban: 0.41 μg/mL and 1.24 μg/mL | Determines the method's sensitivity, vital for detecting low concentrations of the analytes.
Stability Testing | Conducted under various stress conditions following ICH guidelines | Ensures method robustness and suitability for routine quality control testing.
Flow Rate | 1 ml/min | Relevant as it specifies the rate of mobile phase movement in the RP-HPLC system, critical for consistent separation and detection of Aspirin and Apixaban.
Column Temperature | 30°C | Relevant because the temperature affects the retention time, peak shape, and reproducibility of the method for aspirin and apixaban separation.
Detection Wavelength | 227 nm | Relevant as the UV detection wavelength is optimized for capturing the maximum absorbance of both Aspirin and Apixaban in the RP-HPLC method.
Mobile Phase Composition | 40:60 Acetonitrile: phosphate buffer (pH 4) modified by O-Phosphoric Acid | Relevant because this ratio and pH ensure effective separation and peak resolution for the two compounds.
Retention Time for Aspirin | 5.37 min | Relevant as it confirms the reliable identification of Aspirin under the defined HPLC conditions.
Retention Time for Apixaban | 7.10 min | Relevant as it validates the method for the distinct detection and quantification of Apixaban.
Linearity Range for Aspirin | 50-300 μg/mL | Relevant because it defines the concentration range over which Aspirin exhibits a linear response under the validated method.
Linearity Range for Apixaban | 5-15 μg/mL | Relevant as it outlines the calibration curve range for Apixaban, critical for method validation.
Limit of Detection (LOD) for Aspirin | 0.84 μg/mL | Relevant to establish the minimum detectable concentration for Aspirin in the RP-HPLC method.
Limit of Detection (LOD) for Apixaban | 0.41 μg/mL | Relevant to determine the sensitivity of the method for detecting Apixaban at low levels.
Limit of Quantification (LOQ) for Aspirin | 2.55 μg/mL | Relevant as it defines the smallest quantifiable measurement of Aspirin using the RP-HPLC method.
Limit of Quantification (LOQ) for Apixaban | 1.24 μg/mL | Relevant because it specifies the lowest concentration of Apixaban that can be reliably quantified in the method.
Stability Testing Basis | ICH Guidelines | Relevant as it ensures the method’s robustness and reliability under stress conditions, in line with international standards.
Method Validation Parameters | Linearity, precision, accuracy, sensitivity | These factors affect the reliability and precision of pharmacokinetic studies as they ensure accurate quantification of drug concentrations.
Mobile Phase Composition | Acetonitrile and phosphate buffer (pH 4) modified with O-Phosphoric Acid | The pH and composition influence the solubility and retention behavior of the drugs, which are key to understanding their pharmacokinetics.
Column Conditions | Phenomenex Luna C18, 30°C column temperature | Column conditions impact separation efficiency, which is critical for accurate determination of pharmacokinetics.
Flow Rate | 1 mL/min | The flow rate affects resolution and retention time, influencing the accurate pharmacokinetic profiling of drugs.
Detection | UV at 227 nm | UV detection enables precise measurement of drug concentration, fundamental for pharmacokinetic analysis.
Linearity Range for Aspirin | 50-300 μg/mL | Establishing linearity helps ensure the drug concentration range is accurately detected, impacting pharmacokinetics.
Linearity Range for Apixaban | 5-15 μg/mL | Apixaban's quantifiable range is crucial for its pharmacokinetic assessments.
Stability Testing | Conducted under ICH guidelines | Stability assessments ensure the drugs maintain efficacy during pharmacokinetic evaluations under stress conditions.