## Freetext Data
Data | Value | Relevance
Quantitation Method | LC-MS/MS | The study developed an LC-MS/MS method, central to addressing the pharmacokinetics of tazemetostat.
Sample Volume | 20 μL | Specifies the small sample volume required for conducting the plasma quantitation assay, making it feasible for clinical use.
Analyte | Tazemetostat | Demonstrates the detection focus on tazemetostat, which is relevant for its pharmacokinetics in lymphoma patients.
Chromatographic Column | Kinetex C18 | Relevant for achieving chromatographic separation of tazemetostat in plasma samples.
Gradient Composition | 0.1% formic acid in water and acetonitrile | Details the gradient used for chromatographic separation, critical to the assay development.
Run Time | 3 minutes | Demonstrates the efficiency and feasibility of the assay for high-throughput clinical analyses.
Detection Method | SCIEX 6500+ tandem-mass spectrometer with electrospray positive-mode ionization | Highlights the detection setup, ensuring high specificity and sensitivity for tazemetostat.
Validation Guidance | FDA latest guidance | Ensures the assay complies with regulatory standards, supporting its reliability in clinical studies.
Linear Range | 10–5,000 ng/mL | Relevant for defining the quantitation limits of tazemetostat in plasma samples.
Accuracy Range | 91.9–103.7% | Indicates the accuracy of the validated method, giving confidence in the data generated.
Imprecision | <4.4% | Demonstrates the precision of the assay, critical for reproducibility in clinical trials.
Recovery Range | 93.3–121.1% | Supports the efficiency of sample preparation and analyte extraction, key to accurate quantitation.
Matrix Effect | −25.5% to −4.9% | Indicates robustness of the assay to matrix-related variability, relevant for patient sample analysis.
Impact of Hemolysis, Lipemia, and Dilution | No impact | Ensures the method is robust across different plasma conditions, a common challenge in clinical samples.
Freeze-Thaw Stability | Stable after 3 cycles | Shows the method's reliability across repeated sample handling, relevant for long-term sample analyses.
Room Temperature Stability | Stable for 24 h | Indicates the flexibility of sample processing under varying conditions in clinical workflows.
Long-term Storage Stability | Stable for 4 months at −80°C | Essential for storage of biobanked patient samples over trial duration.
Incurred Sample Reanalysis Agreement | 94.4% within 20% difference (n=36) | Validates the method's reliability and reproducibility when analyzing real patient samples.
External Validation Bias | −11.1% | Demonstrates the accuracy and reliability of the assay when tested externally.
Pharmacokinetic Findings | Variable concentration-time profiles observed in 3 patients | Relevant to understanding tazemetostat pharmacokinetics in lymphoma patients, driving further study needs.
Assay Suitability | Suitable for tazemetostat quantitation in human plasma | Confirms the assay’s capability to support clinical pharmacokinetic studies.
Incurred sample reanalysis yielded | 94.4% samples within 20% difference (n=36) | Indicates sample size used in the pharmacokinetic (PK) study, which affects the reliability of the PK analysis for patient data.
Patients in PK study | Three patients | Identifies the actual patient sample size for initial PK data collection, critical for understanding the variability in pharmacokinetics.
Patient population | Patients with lymphomas | Specifies the target population for the study, providing context for how the PK data will apply in clinical and therapeutic settings.
Pharmacokinetic data variability | Variable concentration-time profiles | Highlights the observation of variability in PK profiles, supporting the rationale for further data collection to understand pharmacokinetics in the patient population.
Not mentioned |  | The paper does not specify a pharmacokinetic model, such as 1-compartment, 2-compartment, physiologically based model, or IVIVE. It provides quantitative plasma levels and pharmacokinetic data but does not directly reference or employ a specific compartmental or physiological modeling approach.
Protein precipitation method | Used | Relevant as it describes the sample preparation parameter estimated from the developed method.
Chromatographic column | Kinetex C18 | Relevant as it refers to the specific parameter used for separation in LC-MS/MS.
Chromatographic run time | 3 minutes | Relevant as it defines the duration of separation parameter estimated from experimental setup.
Mobile phase composition | 0.1% formic acid in water and acetonitrile | Relevant as it specifies gradient conditions, an important parameter for chromatographic separation estimated from this method.
Detection method | SCIEX 6500+ tandem-mass spectrometer with electrospray positive-mode ionization | Relevant as it specifies the detection instrumentation and ionization mode, crucial for parameter determination.
Linearity range | 10–5,000 ng/mL | Relevant as this parameter represents the assay's capability to predict analyte concentration, estimated from validation results.
Accuracy range | 91.9–103.7% | Relevant as this parameter evaluates assay performance, estimated through validation against known concentrations.
Precision (imprecision) | <4.4% | Relevant as it indicates repeatability of the assay, a parameter determined during validation.
Recovery range | 93.3–121.1% | Relevant as it demonstrates analyte extraction efficiency, an estimated parameter via recovery tests.
Matrix effect range | −25.5% to −4.9% | Relevant as it assesses the influence of the plasma matrix, determined through experimental validation.
Plasma stability conditions | 3 freeze-thaw cycles; 24 h at room temperature; 4 months at −80°C | Relevant as it represents parameters estimating analyte stability under different handling conditions.
Incurred sample reanalysis | 94.4% samples within 20% difference (n=36) | Relevant as it serves as a confirmation of method reproducibility, parameter estimated from actual patient samples.
External validation bias | −11.1% | Relevant as it establishes the reliability of assay performance in external testing, a calculated parameter.
Pharmacokinetic variability | Variable concentration-time profiles in 3 patients | Relevant as it refers to parameters describing drug behavior in plasma, estimated from patient pharmacokinetic data.
Hemolysis | Did not impact quantitation | Hemolysis was assessed as a potential factor but was found to have no influence on tazemetostat pharmacokinetics based on the study's plasma quantitation approach.
Lipemia | Did not impact quantitation | Lipemia was evaluated and determined not to affect the analytical quantitation of tazemetostat, indicating it does not influence pharmacokinetics in this context.
Dilution | Did not impact quantitation | The study confirmed that dilution of plasma samples does not interfere with tazemetostat quantitation, suggesting this factor does not contribute significantly to pharmacokinetics variability.
Matrix effect | Ranged from −25.5% to −4.9% | Matrix effect differences were observed, which could potentially explain some variability in pharmacokinetics by influencing LC-MS/MS method sensitivity and accuracy.
Freeze-thaw cycles | No impact after 3 cycles | Stability studies showed no degradation of tazemetostat after multiple freeze-thaw cycles, ensuring accurate pharmacokinetics assessment under routine sample handling conditions.
Room temperature | Stable for 24 hours | Stability at room temperature confirms tazemetostat is not affected during short-term handling, which supports reliable pharmacokinetics measurements.
Long-term storage | Stable for 4 months at −80°C | The drug's long-term stability supports its consistent pharmacokinetic evaluation over extended periods.
Patient variability | Variable concentration-time profiles observed | Observed inter-patient variability in drug plasma concentration-time profiles suggests that individual patient differences may influence pharmacokinetics.
External validation bias | Mean bias of −11.1% | External validation bias highlights analytical discrepancies that could affect pharmacokinetics interpretation.