## Table 1
Scanner | aFOV | NEMA sensitivity | NEMA peak NECR | Ref
Siemens Biograph Vision 600 | 26.1 cm | 16.4 kcps/MBq | 306 kcps at 32 kBq/mL | [27]
GE Discovery MI | 30 cm | 32.76 cps/kBq | 434.3 kcps at 23.6 kBq/mL | [5]
PennPET Explorer | 64 cm | 55 kcps/MBq | 1050 kcps at 38 kBq/mL | [6]
United Imaging uExplorer | 194 cm | 174 kcps/MBq | 1855 kcps at 9.6 kBq/mL | [28]
Siemens Biograph Vision Quadra | 106 cm | 176 kcps/MBq | 2956 kcps at 27.49 kBq/mL | [17]

## Freetext Data
Data | Value | Relevance
Sensitivity comparisons between digital SAFOV and LAFOV systems | LAFOV PET is 10-40 times more sensitive than SAFOV systems based on modeling and clinical data | Demonstrates improvements brought by LAFOV PET systems in terms of scanner sensitivity, crucial for better image quality and more precise results.
Benefits of LAFOV sensitivity | Greater acceptance angles improve detection sensitivity and signal-to-noise ratio | Highlights why LAFOV PET scanners outperform SAFOV scanners and can reduce noise issues for clarity in imaging interpretation.
Reduction in acquisition time | LAFOV scanners can acquire images of a 106 cm field in two minutes compared to 16 minutes with SAFOV | Relevant to discussing time savings and patient throughput optimization as clinical advantages.
Ability to scan whole body simultaneously | LAFOV scanners enable imaging all organs or body regions at once | Relevant for studying organ-organ interactions and multisystem disease, exceeding SAFOV's limitations.
Clinical applications of ultra-low dose emission imaging | Lower radiopharmaceutical activity allows imaging children and pregnant patients safely | Demonstrates applications previously unsuitable for PET imaging, emphasizing expanded clinical relevance of LAFOV.
Impact on dosimetry | Simultaneous imaging of critical organs reduces inaccuracies and accelerates radiopharmaceutical development | Highlights key advantages LAFOV PET provides for research and tracer trials compared to SAFOV systems.
Dynamic acquisition and kinetic parameter imaging | Enables calculation of biologically significant indices such as Ki, Kpat, SUV, or Vb | Relevant to exploring improved diagnosis and disease progression metrics obtainable with LAFOV PET.
Role in tracer optimization | LAFOV PET facilitates late imaging with [18F]-FDG with improved contrast | Tracer flexibility is essential for diagnostics where SAFOV systems posed limitations due to decreasing activity.
Advantages for cardiac oncology imaging | Simultaneous imaging of cancer and cardiac health with lower activities | Relevant for expanded applications beyond oncology imaging into multisystem impacts, particularly cardio-oncology.
Artificial intelligence application | AI can enhance image quality, provide synthetic CT corrections, or perform tracer multiplex imaging | Relevant for future key developments leveraging LAFOV systems coupled with AI to optimize imaging and diagnostics.
Advantages for pediatric imaging | Sedation-free rapid imaging for infants due to quicker acquisitions and reduced radiation | Expands accessibility of PET imaging for sensitive patient populations where SAFOV scanners struggle.
Radiation dose reduction and CT optimization | Potential reductions >90% in combined emissions and tin-filtered CT | Crucial to expanding LAFOV PET/CT systems' adoption into areas requiring minimal radiation exposure compared to SAFOV.
Reported sensitivity comparisons for digital SAFOV and clinical LAFOV systems | NEMA NU 2–2018 methodology | Describes the framework used to assess sensitivity characteristics that could influence kinetic parameter assessment.
Dynamic full-dose PET/CT | Consumes terabytes of data | Relevant to kinetic study design due to the extensive data needs for dynamic acquisitions capturing temporal changes.
High temporal resolution | Allows simultaneous imaging of all organs | Essential for kinetic studies as it provides accurate temporal data across multiple systems.
Dose optimization | Measurements based on organ and lesion-specific uptake at several timepoints | Crucial for kinetic parameter assessment and personalized dosimetry calculations.
Tracer kinetic acquisition parameters | Short acquisition times enabled dynamic studies in pediatric patients without sedation | Demonstrates flexibility in patient populations including children, relevant for kinetic imaging.
Dynamic acquisitions for [18F]FDG | Utilized to calculate kinetic indices like Ki and Kpat | Provides examples of kinetic modeling applications in PET imaging.
Number of patients | Not specified directly in the text | Relevant but requires clarity to define the patient sample size in PK study designs.
Adult and pediatric applications | Achieved ultra-low dose capability in pediatric imaging without sedation | Demonstrates the scanner’s adaptability for PK studies across diverse patient populations.
Gantries | MRD of 85 rings | The gantry configuration is crucial for data collection and affects the modeling approach in PET imaging.
Kinetic Models | Ki, Kpat, Vb, K1 | These kinetic parameters reference compartmental modeling methods that provide insights into tracer behavior in tissue, indicating the type of compartmental model used.
Analysis Technique | Patlak analysis | Indicates the use of a linearized model often associated with a 1-compartment system for calculating net influx rates like Ki.
Dosimetry | Whole-body imaging on LAFOV enables simultaneous organ data collection | To calculate radiation dose, whole-body data contribute to physiological modeling to predict pharmacokinetics/dynamics.
Application | Compartment analysis for [18F]FDG dynamics | Relevant in identifying whether a 1- or 2-compartment model was referenced for tracer transport dynamics.
Dynamic Protocols | Enables kinetic indices | Suggests that physiologically informed, time-resolved models (likely compartment-based) were used in dynamic acquisitions for PET imaging.
Simulation Models | Digital twins | These allow physiological-based pharmacokinetics modeling, typically broader than basic 1- or 2-compartment models, integrating system-wide interactions.
Tracer Development Context | Single-time-point and dynamic approaches for acquiring Ki | Indicates compartmental modeling principles governing PET-derived parameters are applied via structured model frameworks.
Reported sensitivity comparisons | Digital SAFOV and current clinical LAFOV systems | The comparison highlights the sensitivity differences between short and long axial field of view scanners, central to evaluating LAFOV performance.
Axial Field-of-View (aFOV) | 162 mm (first clinical system), 300 mm (GE Discovery MI PET/CT), 194 cm (United Imaging uExplorer), 106 cm (Siemens Biograph Vision Quadra) | The estimated values provide a clear picture of how axial field-of-view has evolved over time, particularly in LAFOV systems.
NEMA Peak NECR | Substantially higher for Biograph Vision Quadra than uExplorer | This parameter, estimated using phantom data, reflects the noise-equivalent count rate performance and indicates technological differences impacting sensitivity.
NEMA Sensitivity | Factor 10 higher in Biograph Vision Quadra relative to Biograph Vision 600 | Illustrates improvements in sensitivity based on detection architecture and gantry length, validated through clinical data.
Signal-to-Noise Ratio (SNR) | Higher for LAFOV systems compared to SAFOV | SNR improvements are consistently reported and impact image quality and clinical performance directly.
Minimum Ring Difference (MRD) | 85 rings for homogeneous sensitivity | Estimated for image uniformity across the scanner's field, critical for quality imaging in LAFOV systems.
Lesion Acquisition Time | 16 min (SAFOV, 106 cm) vs. 2 min (LAFOV) | Reflects time efficiency which is an essential operational parameter, directly estimated from clinical LAFOV and SAFOV applications.
Standardised Uptake Value (SUV) Parameters | SUV-based semi-quantitative analysis | SUV is a key parameter derived from dynamic data, reflecting metabolic activity which is estimated directly from patient scans.
Radiotracer Activity | Fluorine-18-labelled (240 MBq SAFOV, 264 MBq SAFOV) | Indicates injected activity for achieving comparative image quality, directly related to the scanner systems used.
Dynamic Scan Data Consumption | Terabytes per scan | Highlights the computational demand associated with LAFOV systems, necessitating considerations for data infrastructure.
CT Radiation Dose Reduction | >90% reduction using tin filters | Provides an estimate for enhancing patient safety when paired with high-sensitivity LAFOV data.
Kinetic Indices (e.g., Ki, Vb) | Derived through dynamic acquisition | These represent advanced quantitative measurements made possible by high-resolution temporal data from LAFOV scans.
By increasing the aFOV in TB- and LAFOV PET scanners with multiple additional detector rings to increase body coverage by at least 4 times | There is not only a simple 4x gain in body coverage per unit of time, but due to a large increase in acceptance angles of the coincident 511 keV photons resulting from positron/electron annihilation, there is a much greater increase in sensitivity; at least 10 to 40 times that of SAFOV scanners. | High sensitivity plays a role in enhancing pharmacokinetics evaluation by improving tracer detectability and accuracy.
The ability to scan in new ways without the constraints placed by sensitivity | Allows the assessment of dynamic processes such as kinetic data and organ connectome evaluations. | Dynamic data acquisition enables a deeper understanding of pharmacokinetics and inter-organ interactions.
The ability to routinely acquire good quality images at later than conventional times post-injection | Maximises diagnostic accuracy in alignment with biological pharmacokinetics of tracers (e.g., decay rates and uptake dynamic). | Late imaging better aligns imaging protocols with tracer biodistribution and clearance patterns.
Examples that have been described for [18F]FDG are Ki or Kpat, Vb, and K1 | Provides specific quantitative indices that relate to reversible/irreversible kinetics, blood volume fraction, and plasma-to-tissue delivery of tracers. | Pharmacokinetic modeling benefits from dynamic PET acquisition and measurable data.
The high sensitivity and timing resolution of LAFOV scanners | Facilitates the acquisition of dynamic data and extraction of high-quality arterial blood time activity curves. | Enhances the understanding of drug clearance and uptake longitudinally in pharmacokinetic studies.
The ability to scan the entire body allows for dosimetry of [89Zr]-based therapies post-injection stages | Helps measure pharmacokinetic parameters such as dose uptake, clearance, and target-to-background ratio, especially for monoclonal antibodies. | Supports pharmacokinetic modeling necessary for optimizing drug dose and timing.
Tracer cocktails or multi-radionuclide imaging enabled by LAFOV PET sensitivity | Permits the simultaneous study of multiple pharmacokinetics of radiopharmaceuticals within the same patient. | Pharmaco-dynamic interaction of tracers in personalized therapies is a key aspect reported.
Routine imaging at biologically optimal timepoints | Enhances lesion contrast and differentiation between benign and malignancy based on tracer kinetics. | Relates to the dynamic nature of radiopharmaceutical biodistribution over time.