## Table 1
Unnamed: 0 | LM1010 (Method A)
Instrument | LM1010 (certified medical device),Hitachi High-Tech Analytical Science (Tokyo, Japan)
Mobile phase | Mobile phase A, mobile phase B, Hitachi High-Tech Analytical Science
Column | LaChrome LM Type A, Hitachi High-Tech Analytical Science
Extraction method | Spin column set, Hitachi High-Tech Analytical Science
Plasma volume (μL) | 150
Fresh plasma sample | 1 μg/mL spiked serum recovery: 99.6%
Frozen (−30 °C) plasma sample | 1 μg/mL spiked serum recovery: 97.3%
Lower limits of quantitation (μg/mL) | 0.276
Calibration curve range (μg/mL) | 1–5
CV (%) | 0.630–1.02
Accuracy (%) | 99.6–104.2
Retention time of VRCZ blood concentration (min) | 1.76

## Table 2
Unnamed: 0 | Unnamed: 1 | p-Value
Number | 7
Age | 64 (55–76)
Sex (male) (%) | 2/7 (28.6)
Weight (kg) | 52.0 (36.5–67.0)
ALT (U/L) | 17.0 (5.0–71.0)
AST (U/L) | 25.5 (13.0–69.0)
Scr (mg/dL) | 0.78 (0.36–1.02)
BUN (mg/dL) | 18.0 (6.0–40.9)
Total bilirubin (mg/dL) | 0.30 (0.20–1.4)
VRCZ initial dosage (mg/day) | 600 (300–600)
VRCZ maintenance dosage (mg/day) | 300 (150–400)
VRCZ blood concentration (μg/mL)
Method A | 1.78 (0.33–5.61) | 0.750
Method B | 1.53 (0.31–5.47) | 0.750
Time required for measuring (h)
Method A | 0.433 (0.400–0.467) | <0.001
Method B | 74.3 (71.2–94.2) | <0.001

## Table 3
Patients No. | No.1 | No.2 | No.3
Age (years) | 71 | 70 | 65
Sex | Female | Female | Male
Weight (kg) | 36.5 | 63.9 | 50.5
Diagnosis | Chronic necrotizing pulmonary aspergillosis | Fungal sinusitis | Chronic necrotizing pulmonary aspergillosis
Pathogen | Aspergillus niger | Acrophialophora spp. | Aspergillus fumigatus
Concomitant medication(dosage/day) | Levothyroxine Sodium Hydrate 62.5 μg | Butyric acid-producing bacillus 3 g | L-Carbocisteine 1500 mg
Concomitant medication(dosage/day) | Ambroxol Hydrochloride 45 mg | Lansoprazole 30 mg | Dimemorfan Phosphate 30 mg
Concomitant medication(dosage/day) | Montelukast Sodium 10 mg | Amlodipine Besilate 5 mg | Entecavir Hydrate 0.5 mg
Concomitant medication(dosage/day) | Gefapixant Citrate 90 mg | Lemborexant 2.5 mg | Eszopiclone 1 mg
Concomitant medication(dosage/day) |  | Ramelteon 8 mg
VRCZ dosage (/day) | 400 mg → 280 mg → 200 mg | 600 mg → 300 mg | 600 mg → 400 mg → 300 mg
VRCZ blood concentration (μg/mL) | 4.52 → 2.40 → 1.99 → 2.74 | 2.29 → 2.51 | 5.26 → 4.37 → 2.17
VRCZ treatment period | 9 months~ | 4.5 months | 5 months~
Hepatic function (before treatment) | ALT: 9 U/L, AST: 30 U/L | ALT: 33 U/L, AST: 17 U/L | ALT: 16 U/L, AST: 19 U/L
Hepatic function (after treatment) | ALT: 10 U/L, AST: 22 U/L | ALT: 17 U/L, AST: 22 U/L | ALT: 15 U/L, AST: 26 U/L
Renal function (before treatment) | Scr: 0.39 mg/dL, BUN: 8.5 mg/dL | Scr: 0.43 mg/dL, BUN: 18.4 mg/dL | Scr: 0.68 mg/dL,BUN: 7.3 mg/dL
Renal function (after treatment) | Scr: 0.53 mg/dL, BUN: 15.6 mg/dL | Scr: 0.54 mg/dL, BUN: 18.0 mg/dL | Scr: 0.57 mg/dL,BUN: 7.1 mg/dL
Hepatic dysfunction
Renal dysfunction
Visual impairment
Treatment outcome | Remission | Death | Remission

## Freetext Data
Data | Value | Relevance
Treatment Alternatives | HPLC-UV vs. LC-MS/MS methods | Relevant as the study compares the effectiveness of rapid TDM approaches for voriconazole in outpatient care to conventional methods.
Timing Difference | 0.433 hours (HPLC-UV) vs. 74.3 hours (LC-MS/MS) | Relevant in assessing the speed improvement of rapid TDM for clinical decision-making.
Correlation Strength | Pearson’s r = 0.988 | Relevant in evaluating the accuracy and agreement of the new HPLC-UV method compared to LC-MS/MS for blood concentration measurement.
Safety Monitoring | Absence of renal, hepatic, visual toxicities in monitored outpatients | Relevant to understanding how rapid TDM supports patient safety during long-term treatment with voriconazole.
Resistance Prevention | Maintains VRCZ target levels between 1–4 μg/mL | Relevant in addressing fungal resistance risks, a key clinical concern outlined in the study.
Implementation Feasibility | Optimized workflow requiring one click and 7-min measurement | Relevant to practical adoption in clinical settings and enhancing outpatient care processes.
Economic Impact | Lower cost compared to LC-MS/MS and avoids outsourcing | Relevant to examining cost advantages that facilitate integration of TDM in outpatient care settings.
Study Scope | Outpatient setting in Japan | Relevant to exploring the expansion of infrastructure and its global applicability, as outpatient facilities in Japan have TDM barriers.
Patient Sample Size | 7 inpatients and 3 outpatients | The limited sample size may affect the statistical power and generalizability of results relevant to the PK study design.
Blood Sampling Time Points | Blood sampling performed before VRCZ administration; measurement cycle of 7 minutes for HPLC-UV method | Sampling times are critical for capturing accurate pharmacokinetic data before drug administration and steady-state concentrations.
Patient Population | Two males and one female outpatient over 65 years with chronic necrotizing pulmonary aspergillosis and fungal sinusitis; inpatients with VRCZ treatment | The demographic and clinical background of patients helps contextualize variability in pharmacokinetics due to age and underlying conditions.
Blood Concentrations Measurement Methods | HPLC-UV (method A) and LC-MS/MS (method B) used; repetitive measurements included to account for intra-individual variability | Accurate measurement methods and inclusion of intra-individual variability are essential to evaluating PK study reliability.
Exclusion Criteria | Patients who opted out of the study | Exclusion criteria ensure the population studied is consistent, promoting valid PK data analysis.
Outpatient Measurement Workflow | Sampling performed on the day of outpatient treatment using HPLC-UV, reporting within 0.433 hours | The outpatient workflow impacts the feasibility and rapid monitoring of PK data for TDM.
Outpatient Follow-Up | Prospective monitoring of treatment duration, adverse events, and dose optimization | Long-term outpatient follow-up provides data on sustained PK parameters and safety profiles over time.
Model Used for TDM | Not explicitly detailed (no specific pharmacokinetic model mentioned like 1-compartment, 2-compartment, physiologically based model, or IVIVE) | The study focused on comparing analytical methodologies (HPLC-UV vs LC-MS/MS) rather than implementing or validating a pharmacokinetic model; relevant to determining whether a model was applied directly but absent evidence suggests reliance on observational data and target ranges.
Optimization Attempt | Use of TDM software provided by Pfizer Japan Inc. (version 1.2) combined with pharmacometrics | This attempts to optimize VRCZ dosing, indicating potential analytical oversight with models or software, yet not specifying the pharmacokinetic compartmental model used.
Concentration Reporting Method | Rapid measurement techniques (HPLC-UV workflow) for TDM | This is indirectly relevant as the method supports therapeutic dose adjustments but isn't tied to a concrete pharmacokinetic modeling framework.
Target Therapeutic Range | 1–4 μg/mL | Therapeutic ranges are often determined using pharmacokinetic simulations/models, but the paper does not employ modeling terminology explicitly, showing a gap in direct compartmental model use.
Deming’s regression slope | 1.03 | Shows high correlation between the HPLC-UV and LC-MS/MS methods, confirming validity of direct blood concentration measurements.
Deming’s regression intercept | -0.08 | Indicates the offset between the two techniques, demonstrating consistency in measurement outcomes.
Pearson’s correlation coefficient | 0.988 | Reflects strong agreement in measured VRCZ blood concentrations between method A and method B.
Mean quantitative bias in Bland–Altman analysis | 0.025 μg/mL | Confirms minimal bias and high reliability of HPLC-UV for therapeutic purposes.
Variance range in Bland–Altman analysis | -0.5059 to 0.5559 μg/mL | Demonstrates acceptable agreement between methods, critical for ensuring accurate dosing.
Lower Limit of Quantification (LLOQ) in HPLC-UV | 0.276 μg/mL | High sensitivity ensures identification of sub-therapeutic levels, directly impacting patient outcomes.
Target therapeutic range for VRCZ | 1–4 μg/mL | Guides dosage optimization, essential for safety and efficacy in fungal infection treatment.
Time required for measurement using HPLC-UV | 0.433 h | Significantly faster processing supports immediate therapeutic decision-making.
Time required for measurement using LC-MS/MS | 74.3 h | Highlights delays impacting outpatient monitoring and treatment adjustments.
Blood sample volume post-filtration | >150 μL | Required for consistency in concentration determination, ensuring reliable patient-level data.
Outpatient age group monitored | >65 years | Represents patient demographic parameters for treatment validation of chronic fungal infections.
Underlying conditions in patients | Chronic necrotizing pulmonary aspergillosis, fungal sinusitis | Clinical contexts where VRCZ efficacy and accurate monitoring are vital.
VRCZ dosage adjustment instances | 2 out of 3 patients | Demonstrates variability in pharmacokinetics and importance of precise monitoring.
Absence of renal dysfunction, hepatic dysfunction, or visual impairment | Monitored outcomes | Validates safety of rapid HPLC-UV testing in outpatient treatment plans.
Inter-patient variability | Reported as a factor in VRCZ pharmacokinetics | Demonstrates the challenge in maintaining optimal drug levels and avoiding adverse effects.
Nonlinear pharmacokinetics of VRCZ | Confirmed | Highlights why TDM is necessary for accurate dosing and safety.
Genetic polymorphisms | Inter-racial differences noted | Contributes to variability in VRCZ levels and underlines pharmacogenetic considerations in treatment.
Drug–drug interactions | Numerous interactions noted | Essential for understanding variations in VRCZ pharmacokinetics and potential adverse effects.
Patient age | All monitored patients over 65 years old | Age may influence drug metabolism, necessitating individualized dosing strategies.
Underlying diseases | Chronic necrotizing pulmonary aspergillosis and fungal sinusitis observed | Disease state impacts pharmacokinetics and adjusts target therapy.
Combination drugs | No adverse interactions reported | Validates the importance of considering co-administered medications in therapy design.
Renal dysfunction | Serum creatinine increase criteria established | Impacts drug clearance and necessitates adjustments in dosing.
Hepatic dysfunction | Diagnostic criteria include ALT/alkaline phosphatase thresholds | Liver dysfunction affects drug metabolism, emphasizing TDM importance.
Visual impairment | Symptoms monitored | Recognition of neurotoxicity as a pharmacokinetic-related side effect.