## Freetext Data
Data | Value | Relevance
R-squared correlation | 0.998 | Demonstrates strong quantitative correlation between microsampling device and neat plasma, relevant for assessing device accuracy in longitudinal monitoring.
Median CV for microsampled samples | 1.6% | Indicates high precision in measurements, vital for establishing the reliability of the microsampling device for clinical monitoring.
Protein stability at RT for 232 days | High stability observed except for APOB and APOC1 | Relevant for evaluating the long-term usability of dried plasma samples in clinical trials.
Quantification differences at protein level | Median difference of -1.6% between neat plasma and microsampling device | Important to address calibration issues for accurate clinical biomarker measurement.
Normalization method | Albumin (ALB) normalization used | Ensures consistency of results despite volumetric imprecision, crucial for robust longitudinal studies.
Impact on FGB quantification | Significant decrease observed | Highlights a specific artifact in microsampling that must be addressed to ensure accurate biomarker monitoring.
Advantages of dried plasma format | Improved sample stability and reduced handling/storage issues | Relevant for optimizing sample collection for decentralized clinical trials.
Storage conditions | Stable when stored at room temperature for up to seven months | Critical for reducing storage costs and ensuring sample integrity during extended studies.
Elution buffer optimization | 50 mM AmBic with SDC | Provides better recovery for quantified proteins, improving assay robustness for targeted proteomics.
Carbon emission reduction | Achieved by avoiding cold-chain transportation | Highlights environmental benefits and cost savings in clinical trial workflows.
Patient sample size | Nine healthy donors | Relevant as the study aimed to benchmark the quantitative performance of microsampling devices across a controlled group of individuals, representing a typical sample size for baseline proteomics evaluations.
Sampling method | Fingerprick capillary blood and venous blood samples | Relevant because comparing capillary and venous blood sampling methods ensures robustness and applicability for various clinical settings.
Blood sample volume | 30–60 μL of capillary blood per fingerprick | Relevant to the study design as this ensures minimal patient burden while maintaining sufficient sample volume for analysis.
Neat plasma processing | Blood centrifuged at 2,000 rcf for 10 min | Relevant for establishing a reference workflow and comparing it with microsampling techniques.
Dry sample preparation time | 24 hours drying at room temperature | Relevant as it defines the sample preparation protocol for microsampling device comparison.
Storage conditions | −20 °C after drying for 24 hours | Relevant to examining long-term stability and method compliance during population-level monitoring.
Sampling device | Microsampling device (Capitainer Plasma) evaluated for volumetric precision below 3% CV | Relevant to validating precision and consistency in patient-centric clinical trials.
Patient population | Healthy adults with informed consent under ethical regulations | Relevant because it establishes the population characteristics and ensures compliance with ethical standards for future clinical trial applicability.
Duration of stability study | Up to 232 days storage at room temperature | Relevant for validating the dried plasma microsampling device’s long-term viability for PK studies in decentralized trials.
Model Type | Not explicitly stated | The paper does not describe a specific pharmacokinetic or compartmental model, focusing instead on the evaluation of microsampling device performance for proteomic quantification.
Study Design | Quantitative proteomics approach | While not a traditional pharmacokinetic model, the study uses a targeted quantification method to assess the accuracy and precision of the microsampling device relative to conventional plasma sampling, which relates to model validation but is not inherently compartmental modeling.
Normalization Technique | Albumin normalization | This technique addresses variance in input materials and indirectly supports longitudinal studies but doesn't define a compartmental model often used in drug disposition studies.
Evaluation Method | LC-MS/MS analysis | The lack of references to compartmental models or IVIVE frameworks suggests a focus on robust methodological assessment rather than pharmacokinetic modeling.
Protein concentration range | 1300 μM to 30 nM | This range reflects the diversity in protein concentration levels targeted in plasma samples and benchmarks the device's accuracy across varying abundances, relevant for evaluating the robustness of quantification.
Quantitative correlation coefficient | R > 0.99 | Indicates high accuracy of the microsampling device compared to neat plasma, crucial for validating the device's quantitative performance.
Quantitative precision in healthy donors | CV below 10% for 80% of peptides | Demonstrates reproducibility and reliability of the device in capturing consistent proteomic data across individuals.
Long-term stability storage temperature | Room temperature (RT) up to 232 days | Represents storage conditions critical for assessing the durability and stability of dried plasma obtained using the novel device.
Fibrinogen beta chain outlier deviation | Lower peptide levels in microsampled data compared to neat plasma | Highlights specific biases introduced by the device, relevant for target validation.
Median CV for microsampling samples | 1.6% | Shows high precision in the normalized dataset, relevant for identifying low variability in proteomic samples.
Normalization target | Albumin (ALB) | Albumin normalization compensates for sampling inconsistencies and is relevant for ensuring consistent proteomic quantification.
Quantitative molar difference between plasma and device | Median −1.6% | Reflects minor deviations in protein quantification, showcasing the device's accuracy in producing comparable plasma data.
Protein degradation trends over time | APOB and APOC1 decrease at RT storage | Relevant for stability analysis, which is key for long-term sample usability in clinical proteomics.
Peptide sequence-level concentration | YYCFQGNQFLR higher due to repeat sequence | Validates peptide-level assessment for quantification consistency and its stability within the dynamic range of proteins.
Device volume precision | Below 3% CV | Ensures consistent volumetric sampling necessary for reliable proteomics, aligning device reliability with clinical trial needs.
Collision energy optimization | Table S2 | Relevant for enhancing LC-MS/MS methodology, optimizing mass spectrometry measurements for precise protein quantification.
The device was assessed quantitatively | Quantitative assessment of microsampling device using stable isotope-labeled protein fragments | Relevant as microsampling affects sample integrity and structure, potentially influencing pharmacokinetics analyses.
Variable spot sizes and spot homogeneity in DBS | DBS associated with inconsistencies due to hematocrit levels | Relevant because variability in sample concentration can directly affect pharmacokinetics measurement accuracy.
Contamination with intracellular components in DBS | Leukocyte and erythrocyte-derived matrix interferences | Relevant as contamination impacts protein measurement sensitivity, which is critical for pharmacokinetics studies.
Lateral flow device impact | Produces concentration gradient due to chromatographic effects | Relevant because non-uniform analyte concentration may skew pharmacokinetics data.
Volumetric control of sampling | Ensures precise and reproducible sampling | Relevant as consistency in sample volume improves pharmacokinetics measurement rigor.
Impact of elution buffers | Buffer composition influences protein recovery during elution | Relevant as variations in protein recovery can lead to pharmacokinetic inconsistencies.
Normalization technique | Normalization to albumin due to sampling imprecision | Relevant since normalization corrects sample variability, enhancing pharmacokinetics reliability.
Protein degradation over time | Stability of proteins like APOB and APOC1 affected by storage conditions | Relevant because pharmacokinetics studies depend on accurate protein-level data over longitudinal periods.
Fibrinogen beta chain (FGB) discrepancy | Lower and variable levels in microsampling versus neat plasma | Relevant because bias in protein recovery can alter pharmacokinetics interpretation linked to coagulation-related biomarkers.
Microsampling approach to patient-centric trials | Enables decentralized trials by reducing patient and sampling burden | Relevant since accessibility impacts sample collection frequency and pharmacokinetics profiling depth.