## Table 1
Unnamed: 0 | Calibrationequation | Relative standard deviation (%) | Coefficient ofdetermination | Concentration range of the calibration standard (μmol/L)
Carbamazepine | y = 6614 x | 0.2 | 1.00 | 0.9–84.6
Phenytoin | y = 6067 x | 1 | 1.00 | 0.7–79.3
Voriconazole | y = 4404 x | 1.3 | 1.00 | 0.6–79.3
Lamotrigine | y = 16,148 x | 1.1 | 1.00 | 0.7–79.3
Meropenem | y = 5824 x | 0.5 | 1.00 | 0.5–261.0
Mycophenolic acid | y = 9579 x | 0.2 | 1.00 | 0.6–62.4
Linezolid | y = 10,539 x | 0.2 | 1.00 | 0.6-148.2
Vancomycin | y = 5460 x | 1.6 | 1.00 | 0.1–69.0
Caffeine | y = 1400 x | 0.6 | 0.999-1.00 | 1.0-515.0

## Table 2
Unnamed: 0_level_0_Unnamed: 0_level_1 | Non-analyte reference material_Carbamazepine | Non-analyte reference material_Caffeine
Carbamazepine | 1 | 4.72
Phenytoin | 0.917 | 4.33
Voriconazole | 0.666 | 3.15
Lamotrigine | 2.44 | 11.5
Meropenem | 0.881 | 4.16
Mycophenolic acid | 1.45 | 6.84
Linezolid | 1.59 | 7.53
Vancomycin | 0.826 | 3.9
Caffeine | 0.212 | 1

## Table 3
Unnamed: 0_level_0_Unnamed: 0_level_1 | Sample name_Sample name | Spikedconcentration(μg/mL)_Spikedconcentration(μg/mL) | RMS method(Calibrant: Carbamazepine)_Concentration (μg/mL) | RMS method(Calibrant: Carbamazepine)_RSD(%) | RMS method(Calibrant: Caffeine)_Concentration (μg/mL) | RMS method(Calibrant: Caffeine)_RSD(%) | Conventional method_Concentration (μg/mL) | Conventional method_RSD(%)
Carbamazepine spiked | Sample 1 | 20 | 21.1 | 1.9 | 21.1 | 1.9 | 21.1 | 1.9
Phenytoin spiked | Sample 2 | 2 | 2.1 | 0.9 | 2.1 | 0.9 | 2.1 | 0.9
Phenytoin spiked | Sample 3 | 20 | 18 | 2.5 | 18.1 | 2.5 | 18.1 | 2.5
Voriconazole spiked | Sample 4 | 1 | 1.1 | 1 | 1 | 1.1 | 1.1 | 1
Voriconazole spiked | Sample 5 | 5 | 5.1 | 2 | 5.2 | 1.9 | 5.2 | 1.9
Meropenem spiked | Sample 6 | 0.5 | 0.4 | 1.4 | 0.4 | 1.4 | 0.4 | 1.4
Meropenem spiked | Sample 7 | 100 | 85.5 | 2.8 | 85.9 | 2.8 | 85.5 | 2.8
Mycophenolic acid spiked | Sample 8 | 1 | 1.1 | 0.5 | 1.1 | 0.5 | 1.1 | 0.5
Mycophenolic acid spiked | Sample 9 | 20 | 17.5 | 1.3 | 17.5 | 1.3 | 17.5 | 1.3
Linezolid spiked | Sample 10 | 2 | 2.3 | 0.7 | 2.3 | 0.7 | 2.3 | 0.7
Linezolid spiked | Sample 11 | 50 | 54.9 | 5.5 | 54.8 | 5.5 | 54.7 | 5.5

## Freetext Data
Data | Value | Relevance
Blood drug concentrations measured by the proposed RMS method | Accurate and consistent with conventional methods | Demonstrates the efficacy and validity of the proposed RMS-based HPLC method in determining blood drug concentrations effectively.
Analytes quantified using RMS method | Carbamazepine, phenytoin, voriconazole, lamotrigine, meropenem, mycophenolic acid, linezolid, vancomycin, caffeine | Validates the range of drugs that can be measured using the new RMS technique, expanding its utility for clinical applications.
Non-analyte reference materials used | Carbamazepine and caffeine | These stable and readily available materials serve as alternative references, enabling the quantification of analytes without identical reference materials.
Primary advantage of using RMS for quantification | Does not require identical authentic reference materials for each analyte | Reduces the need, cost, and complexity of acquiring analyte-specific reference standards while maintaining accuracy.
Limitations addressed by RMS method | Difficulty in obtaining stable and pure analyte-specific reference materials | Highlights the method's ability to overcome challenges in conventional HPLC quantification methods.
HPLC method attributes | Fast, accurate, and reproducible measurements | Emphasizes the method's capability to deliver reliable results quickly, essential for therapeutic drug monitoring.
Use of certified reference materials | Traceable to international system of units (SI) | Ensures high reliability and standardization of results obtained via RMS method.
Analytical parameters validated | Accuracy, precision, linearity | Confirms the methodological robustness and reproducibility of the proposed RMS-based HPLC system.
Correlation coefficient (R2) of calibration curves | >0.999 | Indicates excellent linearity of calibration curves, ensuring precise quantification within the validated range.
Drugs currently lacking bioassay kits | Voriconazole, lamotrigine | Demonstrates the clinical need addressed by the RMS method for measuring drugs not supported by existing immunoassay kits.
Measurement time for each HPLC analysis | 7 minutes | Highlights the efficiency of the method, important for routine clinical and emergency applications such as therapeutic drug monitoring.
Application areas | Routine clinical analysis, therapeutic drug monitoring (TDM), toxicological emergencies | Indicates the broad-based clinical relevance and practical implementation of the method.
Support for unstable drug quantification | Meropenem and mycophenolic acid measured without decomposition issues | Demonstrates the method's practical advantage in preserving drug stability, a notable issue with conventional techniques.
Number of calibration points for analytes | 6–10 depending on the drug | Ensures reliable construction of calibration curves for accurate quantitative analysis.
Funding acknowledgement | Supported by Japan Society for the Promotion of Science (JSPS) KAKENHI (grant number 21K05430) | Validates the integrity of the research, supported by national scientific funding for unbiased execution.
Detector used in HPLC | UV-Vis detector | Establishes the methodology details crucial for replication and applicability in various settings.
Main application of HPLC method | Optimization of drug doses and administration based on therapeutic drug monitoring | Provides insight into the practical utility of the method in individualizing and improving patient care.
Targeted pharmacological areas | Broad-spectrum antimicrobials, anticonvulsants, antifungals, immunosuppressants, and toxicity screening | Showcases the versatile clinical applications and relevance of the validated method.
RSD range for calibration slopes | 0.2% to 1.6% | Demonstrates the method's precision and reproducibility in calibration.
Validation results comparison | Matching values with conventional HPLC method | Confirms the reliability and equivalence of the RMS method to standard calibration techniques in terms of drug quantification accuracy.
Patient sample size | Control serum samples spiked with each analyte used liquid control serum I, prepared by Hitachi High-Tech Science Co., Ltd. | Indicates the sample size design, leveraging control serum samples for spiked analytes.
Blood sampling method | 150 μL of control serum samples filtered through a 0.45 μm syringe filter treated with solid-phase extraction | Highlights the handling and pretreatment of serum samples to ensure precise drug quantification.
Tissue/sample type | Control serum samples | Identifies serum as the biological medium used, important for therapeutic drug monitoring applications.
Analyte preparation for test solutions | Analytes were spiked into control serum samples and pretreated with solid-phase extraction using spin columns optimized for each analyte. | Reveals the method of spiking and processing analytes, crucial to designing an accurate pharmacokinetic study.
Analyte concentration ranges | Concentration calibration for analytes varied between 0.15 to 515 μmol/L depending on the drug | Describes the concentration gradient investigated, essential for ensuring data covers appropriate dosing ranges.
Linear range verification | Eight calibration standard solutions for some analytes, six for voriconazole, ten for meropenem, linezolid, vancomycin, and caffeine | Details the systematic approach confirming the detectability, supporting the method's validity.
Patient population | Not specified; the study used control serum samples. | Indicates the study is conducted in a controlled experimental setup without specific patient diversity.
Runtime for HPLC analysis | Measurement completed in 7 minutes, including flushing and re-equilibration | Provides details about the efficiency and rapidity of drug detection, impacting therapeutic applications.
Non-analyte reference materials | Carbamazepine and caffeine used | Identifies controls used for calculating relative molar sensitivity, critical for precise drug concentration evaluation.
Model | Relative Molar Sensitivity (RMS)-Based Quantification Method | The paper focuses on developing a quantification technique using RMS, which serves as the core model for determining drug concentrations without requiring authentic reference materials.
Alternative Model | Conventional HPLC Method | The paper compares the RMS-based approach to the conventional HPLC calibration curve method, which relies on authentic reference materials. This provides context and validation for the model's application.
Calibration wavelengths | 220 nm (phenytoin, lamotrigine, mycophenolic acid), 235 nm (vancomycin), 250 nm (linezolid), 254 nm (voriconazole), 270 nm (caffeine), 280 nm (carbamazepine), 295 nm (meropenem) | Each analyte was monitored at its respective wavelength, indicating the sensitivity and specificity of the HPLC method.
HPLC System | Evaluation HPLC system for LM1010 (Hitachi High-Tech Science Co., Ltd.) | The system used determines the accuracy and reproducibility of drug quantification.
Mobile Phases | Mobile Phase A and Mobile Phase B (Hitachi High-Tech Science Co., Ltd.) | Optimized mobile phases ensure separation performance for different analytes.
Analytical Column | LaChrome LM TypeA analytical column (Hitachi High-Tech Science Co., Ltd.) | This column was critical for effective separation, ensuring accurate quantification.
Purity of Reference Materials | Certified reference materials with traceable purity to SI | Establishes the reliability of calibration and RMS values for accurate measurements.
Limit of Quantification (LoQ) | Carbamazepine: 0.9 μmol/L, Phenytoin: 0.7 μmol/L, Voriconazole: 0.15 μmol/L, Lamotrigine: 0.7 μmol/L, Meropenem: 0.5 μmol/L, Mycophenolic acid: 0.6 μmol/L, Linezolid: 0.6 μmol/L, Vancomycin: 0.1 μmol/L, Caffeine: 1.0 μmol/L | Indicates the minimum concentration detectable for each analyte—essential for therapeutic drug monitoring purposes.
Control Sample Preparation | 150 μL serum filtered with a 0.45 μm syringe filter, pre-treated via solid-phase extraction optimized for each analyte | Ensures sample matrix-related interferences were removed, providing accurate measurements.
Linearity Range | Carbamazepine: 0.9–85 μmol/L, Phenytoin: 0.7–79 μmol/L, Voriconazole: 0.15–5 μmol/L, Lamotrigine: 0.7–79 μmol/L, Meropenem: 0.5–261 μmol/L, Mycophenolic acid: 0.6–62 μmol/L, Linezolid: 0.6–148 μmol/L, Vancomycin: 0.1–69 μmol/L, Caffeine: 1.0–515 μmol/L | The concentration range where the method shows good linearity; crucial for quantifying analytes within therapeutic levels.
RMS Validation Reproducibility | RSD ranging from 0.2 to 1.6% for calibration slopes over replicates | The small relative standard deviation assures the precision of the slope values and RMS approach.
RMS Reference Materials | Carbamazepine, Caffeine | Two stable reference substances with well-defined purity were used to calculate analyte RMS accurately.
HPLC Run Time | 7 minutes per sample, including flushing and re-equilibration | Demonstrates the method's efficiency for routine clinical analyses.
Calibration Slope RSD | < 5.5% for analytes; slopes consistent across triplicates | Consistency supports the reproducibility and reliability of calculating RMS values.
Spin Column Extraction | Solid-phase extraction optimized for each drug | Ensures uniform sample handling and preparation for HPLC analysis.
Analyte Solution Concentrations | Carbamazepine, Lamotrigine, Phenytoin, Mycophenolic Acid: 20 μg/mL; Voriconazole: 5 μg/mL; Linezolid: 50 μg/mL; Vancomycin, Caffeine: 100 μg/mL | An essential parameter for constructing calibration solutions to derive RMS effectively.
Quantitative Agreement | RSD < 5.5% between RMS-based and conventional quantifications | Ensures the RMS method's accuracy relative to the standard calibration method.
Regression Coefficient (R²) for Linearity | > 0.999 for all analytes | Confirms excellent linearity of calibration curves, critical for accurate analyte quantification.
Weight | Differences between individuals' weight can affect drug efficacy and blood drug concentrations. | Weight is mentioned as a factor influencing pharmacokinetics and variability in individual drug response.
Age | Individual age differences can lead to variations in blood drug concentrations. | Age is acknowledged as a contributing factor to pharmacokinetics and therapeutic drug monitoring (TDM).
Pre-existing medical conditions | Conditions like renal impairment and other medical factors impact drug concentration. | Relevant as it shows how medical conditions influence pharmacokinetics and drug response.
Concomitant medications | Taking multiple medications concurrently can alter blood drug concentrations. | Highlights drug-drug interactions as a pharmacokinetics factor, affecting the therapeutic and adverse effects.
Renal impairment | Renal insufficiencies, such as those mentioned for linezolid, can increase drug levels and risk of side effects. | Demonstrates how organ function influences drug clearance and pharmacokinetics.
Dosage regimen | Adjusting dosages helps maintain effective blood concentrations of drugs with narrow therapeutic ranges. | Significantly influences pharmacokinetics by determining optimal drug concentrations.
Intoxication scenarios | Caffeine levels need quick measurement in cases of intoxication to determine pharmacokinetics. | Accounts for emergency situations where pharmacokinetics understanding is crucial.
Stability of drugs | Drugs like meropenem and mycophenolic acid are unstable, affecting their pharmacokinetics. | Impacts how these drugs are measured and stored to ensure accurate pharmacokinetics.