## Table 1
Metamizole concentration, g/L | Creatinine concentration with enzymatic method Median (IQR), µmol/L(%) | Creatinine concentration with Jaffe method Median (IQR), µmol/L(%)
0 | 126 (123, 130)(0.0 %) | 126 (124, 130)(0.0 %)
0.016 | 123 (123, 130)(2.4 %) | 124 (122, 127)(1.6 %)
0.031 | 118 (110, 121)(6.7 %) | 129 (127, 132)(2.4 %)
0.063 | 110 (108, 113)(12.7 %) | 123 (122, 126)(2.8 %)
0.125 | 96 (93, 100)(23.8 %) | 125 (123, 128)(1.2 %)
0.25 | 76 (74.78)(39.7 %) | 125 (123, 129)(0.8 %)
0.5 | 51 (49.52)(59.5 %) | 126 (124, 130)(0.4 %)
1 | 33 (32.34)(74.6 %) | 124 (123, 126)(2.0 %)
2 | 12 (11.12)(90.5 %) | 125 (123–126)(0.8 %)
10 | 5 (5.6)(96 %) | 134 (133, 135)(6.0 %)

## Table 2
Serum sample | Creatinine concentration with enzymatic method, µmol/L | Creatinine concentration with Jaffe method, µmol/L
Sample 1 | 15 | 63
Sample 2 | 42 | 41
Sample 3 | 8 | 93
Sample 4 | 27 | 28
Sample 5 | 3 | 110
Sample 6 | 2 | 141

## Freetext Data
Data | Value | Relevance
--- | --- | ---
Objective of the study | To examine the interference of metamizole in the measurement of serum creatinine concentration by the enzymatic method | Relevant as it directly addresses the primary aim of assessing metamizole interference in creatinine measurement.
Methods used | Enzymatic and Jaffe methods for measuring serum creatinine concentrations | These methods are central to understanding the observed interference of metamizole in laboratory assays.
Interference mechanism | The dipyrone molecule (metamizole) interferes in the final enzymatic reaction sequence due to its similarity with the 4-aminophenazone molecule | Essential for understanding how the interference occurs biologically and chemically in the enzymatic method.
Effect on enzymatic method | Serum creatinine concentrations decrease with increasing concentrations of metamizole | Evidence of interference with the enzymatic method, validating the study's focus.
Effect on Jaffe method | No interference observed with metamizole in creatinine measurement | Highlights the Jaffe method as interference-free and provides a basis for comparing the two methods.
Study location | Bellvitge University Hospital, Spain | Relevant geographic and institutional context for potential clinical extrapolation of findings.
Patient context | 21% of inpatients received intravenous metamizole and 18% received oral metamizole in the hospital | Demonstrates the clinical significance and prevalence of metamizole use in the hospital setting.
Quality control | CV (coefficient of variation) of 7.4% and systematic error under 8.8% were achieved for both methods | Ensures reliability and precision of the measurements conducted during the study.
Maximum therapeutic plasma concentration of metamizole | 0.03–0.1 g/L (IV administration) | Pertinent to establishing the concentration range where interference is observable in clinical cases.
Graphical result | Increased metamizole concentration results in reduced creatinine concentration in enzymatic methods (p<0.05) | Visual and statistical evidence supporting metamizole interference in enzymatic assays.
Role of Trinder reaction | Interference occurs due to processes specific to the Trinder reaction in the enzymatic creatinine measurement | Clarifies the biochemical point of interference that does not affect the Jaffe method.
Clinical recommendation | Reprocess samples with the Jaffe method or request a new sample if interference is suspected | Provides actionable guidance on handling potential errors in creatinine results due to metamizole.
Discrepancy in samples | Samples 1, 3, 5, and 6 showed significant variation in creatinine values between enzymatic and Jaffe methods | Concrete clinical evidence of methodological differences due to interference.
Study limitation | Main metabolites of metamizole were not included in the analysis | Affects the precision of identifying all sources of observed interference.
Pre-analytical advice | Blood sample extraction should precede intravenous metamizole administration | Relevant for preventing potential contamination and reducing errors in clinical practice.
Future research need | In vivo studies with serial sampling after metamizole administration | Calls for exploration to better understand the onset and scope of interference under real clinical conditions.
Ethics approval | Study approved by Research Ethics Committee, Bellvitge University Hospital (PR240/22) | Confirms compliance with research ethics, ensuring validity and ethical rigor of the study.
Laboratory-based percentage error threshold | Total allowable error: 20% | Relevant as it defines the cut-off for assessing clinical significance of interference.
Study Design | Interference study of creatinine concentration measurement with enzymatic and Jaffe methods | Provides the framework for assessing the interference caused by metamizole and its implications.
Patient Samples | Pool of 30 serum samples from patients used for interference testing | Shows the sample size for serum pool utilized in the PK study setup.
Control Group | Sodium chloride added to 180µL serum pool without metamizole as a negative control | Establishes baseline for comparison in interference study.
Sampling Method | Serum samples tested included six patient samples suspected of intravenous metamizole contamination | Indicates inclusion of real-world patient samples directly affected by metamizole exposure.
Therapeutic Range Studied | Metamizole concentrations ranging from 0.03 g/L to 0.15 g/L (therapeutic concentrations) | Demonstrates relevance to pharmacokinetic levels observed during metamizole administration.
Sampling Details | Serum samples thawed and processed using both enzymatic and Jaffe methods | Describes the laboratory handling of patient samples relevant to interference testing.
Patient Population | Inpatients at Bellvitge University Hospital from various departments (e.g., emergency, intensive care) | Highlights a diverse patient group relevant in studying drug-induced interference.
Post-administration PK Detail | Reference to metamizole Cmax after intravenous (0.03-0.015 g/L at 1 hour) and oral (0.15 g/L at 2 hours) administration | Contextualizes sampling within known pharmacokinetics of metamizole.
Study Objective | Examine metamizole interference in enzymatic creatinine measurement | Aligns the study's purpose with interference-driven pharmacokinetic impacts.
Statistical Analysis | Wilcoxon test performed to determine interference significance (p<0.05) | Relevant for validating the PK study results statistically related to method differences.
Limitations | Absence of serial measurements following intravenous administration | Acknowledges gaps in temporal sampling data, which affect PK result generalizability.
Ethical Approval | Research Ethics Committee approval (reference PR240/22) | Ensures adherence to ethical standards in the human study population.
Model Used | Not specified in the study; pharmacokinetic studies referenced report metamizole concentrations in plasma (e.g., 0.03–0.015 g/L after intravenous administration at specific time points) | The paper references pharmacokinetic studies to discuss metamizole concentrations but does not explicitly utilize or specify a pharmacokinetic model like 1-compartment or physiologically based models.
Coefficient of variation for imprecision | 7.4% | This parameter ensures that the measurement methods meet the laboratory's quality requirements by evaluating their analytical precision.
Systematic error | 8.8% | This parameter is crucial for confirming the accuracy of the enzymatic and Jaffe methods in creatinine measurement.
Serum creatinine concentration (baseline average) | 126 µmol/L | This is the starting concentration of creatinine in the serum pool used for the study, enabling the evaluation of metamizole interference.
Metamizole concentration range | 0.03-0.15 g/L | Derived from pharmacokinetic studies and represents therapeutic and intravenous administration levels for assessing potential interference.
Sampling methods | Pool of 30 serum samples | This standardized pooling allows the study to evaluate the impact of metamizole under controlled conditions.
Analyzed samples suspected of interference | 6 serum samples | These samples, flagged due to significant creatinine variability, were analyzed to investigate real-world interference.
Instrument | Cobas c702 analyzer (Roche Diagnostics, Switzerland) | This device was used to perform both enzymatic and Jaffe creatinine measurements, ensuring standardization in methodology.
Internal quality control | Liquid Assayed Multiqual Levels 1 and 3 | Used to validate methodological accuracy and precision over 16 days of the study.
Therapeutic concentration of metamizole (grey zone in graph) | up to 0.062 g/L | This concentration was used as a reference to evaluate analytical significance during interference testing.
来源 Bio-Rad External Quality Assurance Program | Control specifications for systematic error and coefficient of variation | Relevant for calibrating analytical performance of tested methods.
The absence of serial measurements following intravenous metamizole administration | Noted as a study limitation | This is relevant as pharmacokinetics involve the study of drug concentration changes over time, and the lack of serial measurements prevents tracking of metamizole concentrations post-administration.
Maximum metamizole concentrations in plasma | 0.03 to 0.015 g/L after 1 h (IV) and 0.15 g/L after 2 h (oral, 1,000 mg dose) | These values are relevant as they provide information on the peak plasma levels of metamizole based on the mode of administration, influencing pharmacokinetics.
Potential contamination during blood extraction | Elevated concentrations could result from direct venous access during intravenous administration | Highlighting contamination risks is relevant as it may artificially affect observed pharmacokinetics.
Exclusion of metamizole’s main metabolites from analysis | Measurement limitations in the study | Metabolite concentration affects pharmacokinetics, and the inability to measure them limits a complete understanding of interference mechanisms.
Unknown precise intravenous metamizole dose initiating interference | Recognized as a gap in current knowledge | This is relevant as determining doses that impact pharmacokinetics would help clarify at what levels interference becomes significant.