## Table 1
Sample | Clozapine(ng/mL) | Mycophenolic Acid(µg/mL) | Sunitinib(ng/mL) | N-Desethylsunitinib(ng/mL) | Voriconazole(µg/mL)
CS1 | 30 | 0.2 | 1 | 1 | 0.1
CS2 | 60 | 0.4 | 2 | 2 | 0.2
CS3 | 150 | 1 | 5 | 5 | 0.5
CS4 | 750 | 5 | 25 | 25 | 2.5
CS5 | 1500 | 10 | 50 | 50 | 5
CS6 | 3000 | 20 | 100 | 100 | 10
LLOQ | 30 | 0.2 | 1 | 1 | 0.1
LQC | 100 | 0.66 | 3.3 | 3.3 | 0.33
MQC | 300 | 2 | 10 | 10 | 1
HQC | 1000 | 6.6 | 33 | 33 | 3.3

## Table 2
Compound | Type | SRM Transition(m/z) | Q1(V) | CE(V) | Q3(V) | Q-Array Bias(V) | RT(min) | IS
Clozapine | Analyte | 328.70 > 272.10 | −23.0 | −23.0 | −19.0 | - | 2.03 | Clozapine-d8
Clozapine-d8 | IS | 336.90 > 276.10 | −24.0 | −24.0 | −30.0 | - | 2.02 | -
Mycophenolic acid | Analyte | 321.40 > 207.00 | −11.0 | −20.0 | −22.0 | - | 2.48 | Mycophenolic acid-d3
Mycophenolic acid-d3 | IS | 324.05 > 210.05 | −16.0 | −22.0 | −22.0 | - | 2.48 | -
Sunitinib | Analyte | 399.10 > 283.05 | −20.0 | −28.0 | −20.0 | - | 2.1 | Sunitinib-d10
N-desethylsunitinib | Analyte | 371.10 > 283.10 | −19.0 | −22.0 | −20.0 | - | 2.05 | Sunitinib-d10
Sunitinib-d10 | IS | 409.50 > 283.55 | −13.0 | −24.0 | −10.0 | - | 2.11 | -
Voriconazole | Analyte | 349.85 > 281.10 | −13.0 | −16.0 | −13.0 | 100 | 2.4 | Voriconazole-d3
Voriconazole-d3 | IS | 352.90 > 284.10 | −13.0 | −17.0 | −20.0 | - | 2.4 | -

## Table 3
Analyte | Range | Calibration Regression Curves | R 2
Clozapine | 30–3000 ng/mL | y = 0.0017x − 0.0017 | 0.9986
Mycophenolic acid | 0.2–20 µg/mL | y = 0.1741x − 0.0056 | 0.9993
Sunitinib | 1–100 ng/mL | y = 0.0668x + 0.0175 | 0.9989
N-desethylsunitinib | 1–100 ng/mL | y = 0.0833x + 0.0337 | 0.9976
Voriconazole | 0.1–10 µg/mL | y = 0.2430x + 0.0046 | 0.9978

## Table 4
Analyte_Analyte | Intra-Day (n = 5)_RE (%) | Intra-Day (n = 5)_CV (%) | Inter-Day (n = 15)_RE (%) | Inter-Day (n = 15)_CV (%)
Clozapine
LLOQ | 6.2 | 2.8 | 7.4 | 5.7
LQC | −1.6 | 1.2 | −0.4 | 3
MQC | 0.1 | 1.1 | −3.0 | 4.2
HQC | 1.2 | 0.7 | 1.3 | 3.7
Mycophenolic acid
LLOQ | 11.3 | 5.1 | 11.2 | 3.8
LQC | −2.8 | 2.7 | 1.6 | 4.6
MQC | −5.1 | 2.2 | −2.7 | 3.8
HQC | −5.7 | 2.8 | 0.7 | 5.9
Sunitinib
LLOQ | −10.7 | 5.8 | −1.2 | 9.9
LQC | −4.1 | 4.2 | 0.9 | 7.2
MQC | 4.3 | 2.2 | 0.8 | 6.7
HQC | −0.4 | 0.8 | −1.5 | 8.2
N-desethylsunitinib
LLOQ | −14.8 | 8.8 | −4.2 | 10.5
LQC | −2.9 | 2.2 | −0.5 | 6.2
MQC | 3.8 | 2.3 | 3.0 | 5.2
HQC | −0.1 | 1.4 | 0.8 | 5.7
Voriconazole
LLOQ | −2.6 | 3.1 | −5.4 | 8.2
LQC | 2.4 | 1.2 | 4.6 | 3.8
MQC | 4.6 | 1.9 | 5.9 | 5.1
HQC | 3.5 | 1.6 | 7.4 | 4.8

## Table 5
Condition_Condition | Clozapine_RE (%) | Clozapine_CV (%) | Mycophenolic Acid_RE (%) | Mycophenolic Acid_CV (%) | Sunitinib_RE (%) | Sunitinib_CV (%) | N-Desethylsunitinib_RE (%) | N-Desethylsunitinib_CV (%) | Voriconazole_RE (%) | Voriconazole_CV (%)
Room temperature for 24 h | Room temperature for 24 h | Room temperature for 24 h | Room temperature for 24 h | Room temperature for 24 h | Room temperature for 24 h | Room temperature for 24 h | Room temperature for 24 h | Room temperature for 24 h | Room temperature for 24 h | Room temperature for 24 h
LQC | −13.1 | 6.2 | −0.8 | 2.8 | 5.9 | 4.5 | 8.1 | 3.3 | −5.4 | 3.6
MQC | −5.6 | 3.2 | 6.4 | 6.7 | 7.2 | 5.4 | 10.9 | 4.5 | 8.5 | 5.2
HQC | −11.0 | 6.0 | −6.8 | 4.3 | 1.1 | 6.4 | 5.5 | 3.7 | −3.4 | 2.8
4 °C for 24 h | 4 °C for 24 h | 4 °C for 24 h | 4 °C for 24 h | 4 °C for 24 h | 4 °C for 24 h | 4 °C for 24 h | 4 °C for 24 h | 4 °C for 24 h | 4 °C for 24 h | 4 °C for 24 h
LQC | −3.1 | 1.6 | 0.3 | 1.9 | 1.2 | 3.2 | −1.2 | 7.4 | 1.6 | 0.9
MQC | −8.4 | 4.0 | −5.4 | 2.5 | −3.5 | 6.1 | 5.4 | 5.5 | 0.5 | 0.9
HQC | −3.1 | 1.8 | −0.7 | 0.8 | −1.8 | 7.4 | 3.4 | 7.2 | 4.5 | 1.9
4 °C for 28 days | 4 °C for 28 days | 4 °C for 28 days | 4 °C for 28 days | 4 °C for 28 days | 4 °C for 28 days | 4 °C for 28 days | 4 °C for 28 days | 4 °C for 28 days | 4 °C for 28 days | 4 °C for 28 days
LQC | −7.0 | 5.5 | −6.7 | 5.0 | −3.3 | 4.1 | 1.4 | 1.9 | −3.0 | 5.5
MQC | −12.3 | 6.5 | −10.6 | 5.8 | −9.6 | 5.1 | −0.2 | 2.2 | −4.3 | 3.9
HQC | −9.4 | 5.8 | −7.6 | 6.2 | −1.8 | 3.7 | 4.2 | 4.1 | −2.2 | 3.9

## Table 6
Analyte | Intercept (95% CI) | Slope (95% CI) | Significance(CUSUM Test for Linearity) | Correlation (r) | Significance(Spearman’s Correlation)
Clozapine | −5.49 (−49.68–23.59) | 1.04 (0.97–1.13) | 0.82 | 0.978 | <0.0001
Mycophenolic acid | −0.03 (−0.08–0.06) | 1.01 (0.97–1.03) | 0.99 | 0.997 | <0.0001
Sunitinib | −0.53 (−4.90–0.68) | 1.01 (0.98–1.11) | 0.82 | 0.937 | <0.0001
N-desethylsunitinib | −0.24 (−1.45–0.35) | 1.03 (0.99–1.11) | 0.99 | 0.969 | <0.0001
Voriconazole | −0.01 (−0.06–0.05) | 0.99 (0.95–1.04) | 0.82 | 0.995 | <0.0001

## Freetext Data
Data | Value | Relevance
Clinical Setting | Akita University Hospital, 1-1-1 Hondo, Akita 010-8543, Japan | The study was conducted in a clinical setting, illustrating the real-world application of the developed method.
Objective | Development of a method for simultaneous concentration determination of five analytes using LC-MS/MS for clinical TDM with a fully automated system | Key aim of the research to improve TDM in clinical practice by automating drug concentration determination.
Drugs Analyzed | Clozapine, Mycophenolic Acid, Sunitinib, N-Desethylsunitinib, Voriconazole | These drugs are commonly used in therapeutic settings and require monitoring due to their narrow therapeutic ranges and potential side effects.
Automated System Used | CLAM-2030 (Shimadzu Corporation, Kyoto, Japan) | Central to the research as the fully automated LC-MS/MS preprocessing system developed and tested in this study.
Validation Guidelines | US FDA and EMA guidelines for validation of bioanalytical methods | Ensures the developed method meets standard requirements for clinical integration.
Assay Accuracy | Intra-day RE was −14.8% to 11.3%; Inter-day RE was −5.4% to 11.2% | Confirmed that the method meets precision and accuracy requirements for clinical TDM applications.
Study Approval | Ethical Review Committee of Tohoku University Graduate School of Medicine, approval number 2023-1-291 | Validates ethical considerations and adherence to guidelines in clinical research.
Sample Stability | Stability RE values −13.6% to 10.9% and CV < 8.9%; stable up to 28 days at 4°C | Relevant for clinical settings where sample transport and storage might occur.
Comparison | Comparable performance between CLAM and manual methods as shown by Bland–Altman and Passing–Bablok analyses | Demonstrated that the automated method can reliably replace the manual method in clinical practice.
Time Efficiency | CLAM-2030 reduces pretreatment time from 15 min to 3 min per sample | Relevant to streamlining and increasing throughput in routine TDM.
Correlation Coefficients | 0.978–0.997 for the five analytes between automated and manual methods | Indicates a strong agreement and reliability of the automated system in analysis.
Clinical Utility | Fully automated method supports routine clinical TDM with high efficiency and reliability | Directly supports the relevance of the method to improve patient care in clinical pharmacy settings.
Patient Sample Collection | Plasma samples collected through EDTA tubes and stored at 4°C until analysis | Reflects practical procedures aligning with routine clinical workflows.
Patient Sample Size | Not explicitly stated; routine TDM patient samples were analyzed | Relevant because understanding the size of patient samples provides insight into the robustness and generalizability of the study findings.
Blood Sampling | Plasma collected using EDTA collection tubes, centrifuged at 1580×g for 10 minutes | Relevant because plasma collection and preparation are critical steps for analyzing drug concentrations in TDM studies.
Sample Storage Temperature | Plasma samples stored at 4 °C until analysis | Relevant because proper sample storage is necessary to maintain sample integrity for accurate pharmacokinetic measurements.
Ethical Approval | Ethical Review Committee of Tohoku University Graduate School of Medicine, approval number: 2023-1-291 | Relevant because ethical approval ensures the compliance of the study design with ethical standards, which is particularly critical in patient-involved research.
Patient Consent | Obtained using an opt-out method | Relevant because this detail emphasizes the ethical considerations in utilizing patient samples for the study.
Patient Population | Patients undergoing routine therapeutic drug monitoring | Relevant because the study population directly relates to the clinical applicability of the method for TDM.
Sample Handling | Blood centrifugation followed by plasma collection and refrigeration | Relevant because standardized handling impacts the reliability of drug concentration measurements, a pivotal component of TDM.
Therapeutic Drug Monitoring Approach | Personalized treatment optimizing drug dosages based on factors like plasma concentrations, therapeutic efficacy, or adverse reactions | Highlights the optimization framework but does not specify a model type such as compartmental or IVIVE.
Analytical Method Employed | Liquid Chromatography/Tandem Mass Spectrometry (LC-MS/MS) | Discusses the methodological approach for analysis, not a pharmacokinetic model.
Pretreatment System Used | Fully automated CLAM-2030 LC-MS/MS preprocessing system | Mentions an automation tool, but not a model related to drug distribution or kinetics.
Validation Standards | FDA and EMA guidelines for bioanalytical assays | Refers to guidelines for validation, not for pharmacokinetic modeling.
MS Conditions Setup | Retention times, collision energy, MS parameters for analytes | Method validation details, not a mention of a pharmacokinetic model.
Analyte Stability Data | Stable in plasma for up to 28 days at 4 °C | Discusses drug stability, which could feed into models but does not specify one.
Calibration and QC | Linearity over the working range with R2 >0.99 | Relevant to analytical aspects, but not specifying the utilization of a pharmacokinetic model.
Statistical Analyses | Bland–Altman and Passing–Bablok used for method comparisons | Describes statistical validation, unrelated to compartmental or IVIVE models.
Parameters | Interface voltage | Essential to MS analysis; used to generate ions in the LC-MS/MS system.
Value | 4.0 kV | Specifies the parameter setting for accurate LC-MS/MS detection in the CLAM method.
Parameters | Desolvation line temperature | Controls ion transfer by preventing moisture interference in MS analysis.
Value | 250 °C | Used to maintain stability of ion transfer in LC-MS/MS.
Parameters | Block heater temperature | Enhances the desolvation process in MS analysis.
Value | 400 °C | Ensures effective evaporation of solvents, critical for accurate mass spectrometry readings.
Parameters | Interface temperature | Improves ion generation stability in MS analysis.
Value | 300 °C | Directly impacts ion production efficiency for reliable LC-MS/MS analysis.
Parameters | Nebulizing gas flow | Optimizes spraying of the liquid sample to assist in ionization.
Value | 3 L/min | Sets flow rate for uniform sample nebulization in LC-MS/MS.
Parameters | Drying gas flow | Assists in solvent evaporation during ionization.
Value | 10 L/min | Critical for maintaining a dry ion environment in the mass spectrometer.
Parameters | Heating gas flow | Contributes to efficient solvent evaporation.
Value | 10 L/min | Optimizes the drying process to stabilize ions for accurate measurements.
Estimated source | Shimadzu LCMS-8050 system specifications and optimization experiments | Describes the parameters' origin through validation and hardware setup, directly tied to the study’s methodology and accuracy testing.
Therapeutic drug monitoring (TDM) | Personalized treatment approach | TDM is used to optimize drug dosages, which are influenced by individual pharmacokinetic factors such as plasma drug concentrations, therapeutic efficacy, and adverse reactions.
Sunitinib plasma concentrations | Above 50 ng/mL for efficacy and below 100 ng/mL to avoid toxicity | Demonstrates how the pharmacokinetics of sunitinib affect its therapeutic effects and toxicity.
Mycophenolic acid variability | High interindividual pharmacokinetic variability | Indicates the need for TDM owing to differing pharmacokinetics impacting drug concentration and efficacy.
Clozapine blood concentration | Linked to efficacy and side effects like sleepiness and seizures | Highlights the pharmacokinetic factor of clozapine plasma levels influencing its therapeutic and adverse outcomes.
Drug-specific antibodies in immunoassays | Not always commercially available | The reliability of pharmacokinetic data can be influenced by the availability of suitable assays.
Cross-reactions in immunological methods | May overestimate drug concentrations | Pharmacokinetic data accuracy is crucial and susceptible to immunoassay specificity.
Analyte stability | Stable within ±13.6% RE over 28 days | Stability impacts the reliability of pharmacokinetic assessments over time and transport.
Calibration linearity | R2 > 0.99 for all analytes | Ensures accuracy in the assessment of pharmacokinetic profiles across the measured ranges.
Automated CLAM system | Reduces inter-user variability | Influences pharmacokinetic consistency by providing reproducible drug concentration measurements through automation.