## Table 1
Unnamed: 0 | LLOQ0.05 ng/mL | QC A0.15 ng/mL | QC B18.7 ng/mL | QC C37.5 ng/mL
Intra-day
1 Mean [ng/mL] | 0.0536 | 0.141 | 16.7 | 32
Accuracy (%) | 107 | 93.9 | 89 | 85.4
Precision (% CV) | 9.97 | 3.49 | 3.09 | 3.85
2 Mean [ng/mL] | 0.0471 | 0.132 | 16.8 | 34
Accuracy (%) | 94.2 | 88.3 | 89.4 | 90.7
Precision (% CV) | 9.6 | 3.04 | 3.24 | 4.11
3 Mean [ng/mL]) | 0.0498 | 0.142 | 16.8 | 33.3
Accuracy (%) | 99.7 | 95 | 89.6 | 88.9
Precision (% CV) | 8.28 | 6.47 | 5.36 | 2.94
Inter-day
Mean [ng/mL] | 0.0512 | 0.139 | 16.8 | 33.3
Accuracy (%) | 102 | 92.7 | 89.4 | 88.9
Precision (% CV) | 11.2 | 5.4 | 3.87 | 3.22

## Table 2
Unnamed: 0 | QC A0.15 ng/mL | QC B18.7 ng/mL | QC C37.5 ng/mL
Autosampler (24 h at 10–15 °C)
Accuracy (%) | 100 | 102 | 103
Bench-top (4 h)Accuracy (%) | 101 | 90.4 | 87.5
Freeze and thaw
Accuracy (%) | 92.7 | 99.7 | 96.7
Four weeks at−20 °C
Accuracy (%) | 97.3 | 95.8 | 86.4

## Table 3
Unnamed: 0 | QC A0.15 ng/mL | QC B18.7 ng/mL | QC C37.5 ng/mL
Matrix effect (%) | 114 | 95.7 | 96.9
IS-normalised matrix effect (%) | 115 | 94.5 | 98.7
Recovery (%) | 94.2 | 100 | 96.6

## Table 4
Kinetic Parameter | OCT1 | OCT2 | OCT3
Km [mM] | 24.4 ± 10.1 | 6.3 ± 0.9 | 4.6 ± 1.1
Vmax [amol/min/cell] | 515 ± 233 | 1513 ± 706 | 1197 ± 318
IC50 of metformin transport inhibition by verapamil [µM] | 4.6 ± 1.4 | 23.3 ± 6.3 | 28.4 ± 3.0

## Freetext Data
Data | Value | Relevance
Metformin LLOQ | 0.05 ng/mL | Relevant because achieving this low LLOQ supports highly sensitive measurement of metformin, which is critical for in vitro analysis of OCT1-3 activity in cells.
Internal standard | Metformin D6 | Relevant because it ensures precision and accuracy in LC-MS/MS quantification, crucial for validating uptake kinetics and inhibition assays.
Chromatography method | Ion-pairing chromatography with HFBA | Relevant because this method overcomes metformin's high polarity and retention challenges, providing superior peak shape and sensitivity.
Peak width (HFBA) | 6 seconds at baseline | Relevant since the sharp peak improves the sensitivity, enabling accurate quantification of metformin at very low concentrations.
Validation guidelines | FDA and EMA | Relevant because compliance with these bioanalytical method guidelines ensures the reliability and acceptance of the developed assay.
Cell lines used | HEK293-Flp-In-hOCT1, HEK293-hOCT2, HEK293-hOCT3 | Relevant as these cell lines overexpressing OCT1-3 were essential for evaluating the metformin uptake kinetics and inhibition assays.
Inhibitor tested | Verapamil | Relevant as the study aimed to measure OCT1-3 inhibition, and verapamil served as a reference inhibitor to validate the method.
Observed Km for OCT2 | Higher compared to OCT1 and OCT3 | Relevant because it reflects the differences in transporter affinities for metformin, critical for determining OCT-specific transport kinetics.
Sample preparation method | 7% HFBA for protein precipitation and lysis | Relevant as the method eliminates interfering substances and aligns with HFBA-based chromatographic conditions, enhancing assay performance.
Ionisation mode | ESI positive ion mode | Relevant since the selected ionisation mode optimises metformin detection and quantification essential for reliable data.
Matrix effect | Negligible | Relevant because minimal matrix effects ensure the robustness of the assay in cell-based experiments.
Recovery rate | Approximately 100% | Relevant as it confirms efficient analyte extraction from cell pellets, critical for accurate quantification in inhibition and kinetic studies.
Interference from matrix | None detected | Relevant as this specificity ensures that metformin measurements are unaffected by cellular components or contaminations.
Uptake characteristics | Michaelis–Menten kinetics | Relevant as this model describes the transport properties of OCT1-3, vital for assessing the physiological and pharmacological relevance of metformin transport.
Calibration range | 0.05–50 ng/mL | Relevant because it is suitable for the low concentrations of metformin typically present in cellular uptake studies.
Incubation times for uptake | 5–10 minutes depending on OCT type | Relevant as optimising these conditions ensures reliable kinetic measurements for metformin uptake.
Post-column infusion technique | Aqueous ammonia | Relevant as it resolves ion suppression caused by HFBA, enabling sensitive mass spectrometry detection critical for low-concentration quantification.
Inhibitor IC50 values | Micromolar range for verapamil | Relevant because confirming inhibition ranges assures the method's ability to screen OCT inhibitors effectively.
Preparation system | Microtubes with exact cell count and volume | Relevant since precise sample handling ensures consistency and reproducibility in cellular uptake and inhibition assays.
Conceptualization | G.B.-S., J.W., and M.S. | Relevant because it indicates the planning of the study and provides context for how the PK study fits within the scope of the research.
Investigation Methodology | Utilized UPLC-MS/MS to quantify metformin uptake in cell pellets | Relevant because the method provides insight into the sampling and detection approaches for understanding metformin pharmacokinetics (PK) at the cellular level.
Cell Model | HEK293-Flp-In, HEK-hOCT2, and HEK-hOCT3 cells overexpressing organic cation transporters (OCTs) | Relevant as these cell models constitute the "patient population" in the in vitro PK design.
Sampling Time Points | Time-dependent uptake was measured at 0.5 to 30 minutes | Relevant because sampling intervals directly relate to establishing the PK profile in the cellular system.
Sample Preparation | Cell pellets were extracted using 7% HFBA for protein precipitation | Relevant because it describes the sample handling process for PK analysis.
Measurement Incubation Details | Metformin concentration used: 20 µM for OCT1 and OCT3, 2 µM for OCT2 | Relevant because initial drug concentrations are critical for understanding uptake and clearance in PK studies.
Method Sensitivity | Lower limit of quantification (LLOQ): 0.05 ng/mL (0.38 nM) | Relevant to ensure accurate PK profiling even at low metformin concentrations.
Validation Compliance | Method validated per FDA and EMA guidelines | Relevant as this ensures the PK data are reliable and meet international bioanalytical standards.
Experiment Volume | Initial cell sample volume: 0.5–2 µL | Relevant as it represents the miniaturized nature of the sampling process, impacting overall PK study sensitivity.
Technical Replicates | Experiments conducted with technical triplicates and at least three biological replicates | Relevant for ensuring validity and reproducibility in PK study results.
Model | Cellular OCT inhibition assay using HEK293 cells overexpressing OCTs 1–3 | This is relevant because the authors established an in vitro model system to study the uptake and inhibition kinetics of metformin, which can be seen as an experimental/logical basis for in vitro-in vivo extrapolation (IVIVE) for transporter-mediated effects.
Model Type | In vitro cellular assay | This specifies the experimental setup, which is foundational to understanding the pharmacokinetics of metformin in relation to organic cation transporters.
Model Parameters | Michaelis-Menten kinetics for OCTs | These parameters are crucial for defining the kinetic framework underlying the cellular transport model.
Transport Characteristic | Metformin uptake mediated by OCT1-3 | Relevant as it confirms the model's focus on active transporter-mediated processes without passive diffusion, guiding the specificity of cellular kinetic modeling.
Comparison | Kinetics of OCT2 > OCT1 > OCT3 | These transporter-specific kinetic insights help refine in vitro modeling for inhibitors at the cellular transporter level.
m/z transition for metformin | 130 → 60 | This parameter defines the primary ion transition used for metformin quantification in the developed UPLC-MS/MS method.
m/z transition for metformin D6 | 136 → 60 | This parameter represents the primary ion transition for the internal standard, ensuring accurate quantification by accounting for matrix effects.
Capillary voltage | 0.5 kV | This parameter was optimized to achieve sensitive ionization of the analytes during mass spectrometric detection.
Cone voltage | 20 V | This parameter was selected to optimize the ion focusing for accurate multiple reaction monitoring.
Collision energy | 12 V | This value was optimized for efficient fragmentation of metformin and metformin D6 during analysis.
Chromatographic column | Waters Acquity BEH C18 Peptide Column (300 Å, 1.7 μm, 2.1 × 150 mm) | This parameter specifies the column used for optimal chromatographic resolution and retention of metformin.
Flow rate | 0.5 mL/min | This parameter was set to achieve effective separation and peak sharpness during the HFBA-based ion-pairing chromatography.
Eluent A composition | 0.1% v aqueous HFBA | This value specifies the composition of the aqueous mobile phase, enabling hydrophobic ion pairing for metformin retention.
Eluent B composition | ACN + 0.1% v HFBA | This value defines the acetonitrile mobile phase, critical for gradient elution in reverse-phase chromatography.
Gradient starting condition | 98% A/2% B | This parameter initiated chromatographic separation, providing sufficient retention for metformin.
Gradient final condition | 5% A/95% B | This parameter ensured robust elution of highly retained analytes before column re-equilibration.
Lower limit of quantification (LLOQ) | 0.05 ng/mL | This parameter establishes the sensitivity limit of the assay, enabling accurate metformin measurement in cell pellet samples.
Signal-to-noise ratio at LLOQ | ~500 | This parameter supports the established sensitivity of the method, critical for low-concentration metformin detection.
Extraction recovery | ~100% | This parameter verifies the efficiency of the HFBA-based cell lysate extraction method, ensuring accurate quantification.
Matrix effect | Negligible | This parameter indicates method robustness in managing biological interference, ensuring reliable quantification.
Km value for metformin transport via OCT1 | 36.8 ± 2.8 µM | This parameter characterizes the affinity of metformin for OCT1, derived from Michaelis-Menten kinetics in HEK293-Flp-In-hOCT1 cells.
Km value for metformin transport via OCT2 | 16.1 ± 1.1 µM | This parameter reflects the higher affinity of OCT2 for metformin, measured in HEK293-hOCT2 cells.
Km value for metformin transport via OCT3 | 45.7 ± 4.0 µM | This parameter quantifies the metformin affinity in HEK293-hOCT3 cells, indicating a lower affinity compared to OCT2.
Vmax for OCT1 | 18.7 ± 2.1 nmol/10^6 cells/min | This parameter reflects the maximal transport velocity observed for OCT1 during metformin uptake.
Vmax for OCT2 | 60.7 ± 4.0 nmol/10^6 cells/min | This parameter illustrates OCT2's higher transport capacity for metformin in comparison to OCT1 and OCT3.
Vmax for OCT3 | 11.2 ± 1.5 nmol/10^6 cells/min | This parameter indicates OCT3's lower transport capacity for metformin relative to OCT1 and OCT2.
Inhibitor IC50 for OCT1 (verapamil) | 7.3 ± 0.6 µM | This parameter quantifies the potency of verapamil in inhibiting OCT1-mediated metformin transport, derived from inhibition assays.
Inhibitor IC50 for OCT2 (verapamil) | 41.6 ± 2.1 µM | This parameter reveals the comparative potency of verapamil in OCT2 inhibition, derived from IC50 curves in HEK293-hOCT2 cells.
Inhibitor IC50 for OCT3 (verapamil) | 67.2 ± 4.5 µM | This parameter illustrates the weakest inhibition effect of verapamil on OCT3 among the three transporters.
Elution peak width at baseline (HFBA) | 6 s | This parameter demonstrates the improved chromatographic sharpness and resolution achieved with HFBA in the separation process.
Metformin used as an organic cation transporter (OCT) substrate | Yes | Relevant as metformin pharmacokinetics is dependent on transporter-mediated uptake and efflux rather than passive diffusion, affecting absorption and excretion.
Effect of verapamil as an OCT1-3 inhibitor | Verapamil inhibits metformin uptake in a dose-dependent manner with IC50 values in the micromolar range, and inhibition is most potent in OCT1 | Relevant because inhibition of OCTs can alter the pharmacokinetics of metformin by reducing its cellular uptake.
Role of OCT2 in metformin transport | OCT2 has the highest affinity and transport velocity for metformin among OCT1-3 | Relevant as the differential transporter activity directly affects metformin pharmacokinetics, especially its renal excretion.
Chromatographic retention issues in metformin separation | High polarity and hydrophilicity of metformin cause challenges in chromatographic retention | Relevant as measurement techniques impact the evaluation of pharmacokinetic behaviors, including tissue interactions or transporter affinities.
Matrix effects on metformin quantification | Minimal matrix effects reported, and the assay was validated for high specificity in HEK293 and other cell matrices | Relevant as it ensures reliable pharmacokinetic measurement in specific cell environments without interference.
Michaelis-Menten kinetics observed in metformin uptake | Apparent Km values for OCT-mediated metformin uptake are consistent with previously reported values | Relevant as these parameters explain transporter-mediated uptake efficiency, a critical pharmacokinetic factor.
Effect of OCT expression on metformin accumulation | Different cell lines (HEK293-OCT1, HEK293-OCT2, HEK293-OCT3) yield varying metformin accumulation rates and affinity constants | Relevant since variations in transporter expression alter metformin's distribution and elimination, affecting pharmacokinetics.
Cell line-dependent metformin uptake | OCT2-mediated uptake was faster than OCT1 and OCT3 | Relevant because this affects metformin’s pharmacokinetics, particularly in tissue-specific drug absorption and distribution.
Stability of metformin in various conditions | Metformin is stable in storage conditions and after multiple freeze-thaw cycles | Relevant as it ensures that pharmacokinetic studies yield consistent data unaffected by sample degradation.
Impact of ion pairing in chromatography on metformin detection | Use of HFBA improved retention and peak shape compared to TFA for metformin detection | Relevant since accurate quantification is essential for studying pharmacokinetics under specialized experimental conditions.