## Table 1
Method | Parameter | Reference values | APAP | Unnamed: 4 | NAPQI | Unnamed: 6 | NAC
HPTLC | Retardation factor (Rf) | – | 0.25 |  | 0.40 |  | 0.55
HPTLC | Capacity factor (k0) | 1–10 | 1.01 |  | 1.52 |  | 2.34
HPTLC | Selectivity(α) | ≥ 1 |  | 2.04 |  | 1.48
HPTLC | Resolution (Rs) | > 1.5 |  | 1.6 |  | 2.78
HPTLC | Tailing factor(T) | ≤ 2 | 0.99 |  | 1.32 |  | 1.44
Method | Parameter |  | APAP |  | NAPQI |  | NAC
HPLC | Capacity factor (k0) | 1–10 | 2.34 |  | 5.48 |  | 7.08
HPLC | Selectivity(α) | ≥ 1 |  | 1.89 |  | 3.45
HPLC | Resolution (Rs) | > 1.5 |  | 4.43 |  | 6.25
HPLC | Tailing factor(T) | ≤ 2 | 1.45 |  | 1.56 |  | 1.27
HPLC | Column efficiency (N) | N > 2000 | 2358 |  | 4545 |  | 2747
HPLC | Height equivalent to theoretical plate(cm/plate) | The smaller the value the higher the column efficiency | 0.0106 |  | 0.0055 |  | 0.0091
HPLC | Retention time tR (min) | – | 2.435 |  | 3.812 |  | 5.093

## Table 2
Method parameter_Method parameter | HPTLC-densitometry method_APAP | HPTLC-densitometry method_NAPQI | HPTLC-densitometry method_NAC | HPLC–DAD method_APAP | HPLC–DAD method_NAPQI | HPLC–DAD method_NAC
Range | 0.5–6 μg/band | 1–7 μg/band | 0.3–18 μg/band | 1–30 μg/mL | 10–200 μg/mL | 3–100 μg/mL
Slope | 55.75 | 125.84 | 98.35 | 345.15 | 67.04 | 167.64
Intercept | − 43 | − 34.22 | 12.19 | − 280.09 | − 123.37 | 70.92
Correlation coefficient (r) | 0.9998 | 0.9999 | 0.9998 | 0.9999 | 0.9997 | 0.9998
Accuracy (mean ± SD) | 100.43 ± 0.43 | 101.85 ± 1.35 | 100.87 ± 1.01 | 101.94 ± 1.29 | 100.99 ± 1.45 | 99.51 ± 1.22
LOD | 0.126 | 0.211 | 0.083 | 0.314 | 2.659 | 0.558
LOQ | 0.358 | 0.614 | 0.233 | 0.739 | 7.386 | 1.542
Precision (RSD%) | Precision (RSD%) | Precision (RSD%) | Precision (RSD%) | Precision (RSD%) | Precision (RSD%) | Precision (RSD%)
Repeatability | 0.843 | 0.357 | 0.454 | 0.893 | 0.593 | 0.845
Intermediate precision | 1.393 | 1.427 | 1.945 | 1.468 | 1.359 | 1.375
Specificity (mean ± SD) | 100.39 ± 0.956 | 101. 93 ± 0.780 | 99.12 ± 0.748 | 101.38 ± 1.836 | 101.83 ± 0.859 | 100.83 ± 0.943

## Table 3
HPTLC-densitometry method_APAP_Added µg/band | HPTLC-densitometry method_APAP_%Recovery | HPTLC-densitometry method_NAPQI_Added µg/band | HPTLC-densitometry method_NAPQI_%Recovery | HPTLC-densitometry method_NAC_Added µg/band | HPTLC-densitometry method_NAC_%Recovery | HPLC–DAD method_APAP_Added µg/mL | HPLC–DAD method_APAP_%Recovery | HPLC–DAD method_NAPQI_Added µg/mL | HPLC–DAD method_NAPQI_%Recovery | HPLC–DAD method_NAC_Added µg/mL | HPLC–DAD method_NAC_%Recovery
1 | 102.23 | 1 | 103.09 | 1 | 107.46 | 1 | 112.43 | 15 | 109.53 | 5 | 105.44
2 | 106.23 | 4 | 105.73 | 4 | 103.87 | 5 | 106.66 | 30 | 105.87 | 20 | 108.88
3 | 112.34 | 5 | 109.34 | 10 | 104.73 | 20 | 104.65 | 100 | 109.76 | 70 | 110.53
5 | 110.34 | 6 | 102.85 | 16 | 103.89 | 25 | 103.89 | 150 | 104.97 | 90 | 103.97
Mean ± SD | 107.78 ± 2.033 | Mean ± SD | 105.25 ± 2.256 | Mean ± SD | 104.98 ± 2.313 | Mean ± SD | 106.90 ± 2.117 | Mean ± SD | 107.53 ± 2.284 | Mean ± SD | 107.20 ± 2.493

## Table 4
Method parameter_Method parameter | HPTLC-densitometry method_APAP | HPTLC-densitometry method_NAPQI | HPTLC-densitometry method_NAC | HPLC–DAD method_APAP | HPLC–DAD method_NAPQI | HPLC–DAD method_NAC | Reported method 26_APAP | Reported method 26_NAPQI | Reported method 19_APAP | Reported method 19_NAC
Range | 0.5–6.0 μg/band | 1.0–7.0 μg/band | 0.3–18.0 μg/band | 1.0–30.0 μg/mL | 10.0–200.0 μg/mL | 3.0–100.0 μg/mL | 1–8 μg/mL | 1–8 μg/mL | 0.2–1000 μg/mL | 0.02–100 μg/mL
Mobile Phase | Methanol: ethyl acetate: glacial acetic acid (8:2:0.2) | Methanol: ethyl acetate: glacial acetic acid (8:2:0.2) | Methanol: ethyl acetate: glacial acetic acid (8:2:0.2) | Water: methanol: formic acid (70: 30: 0.15) | Water: methanol: formic acid (70: 30: 0.15) | Water: methanol: formic acid (70: 30: 0.15) | Methanol, water, and buffer solution (20:40:40) | Methanol, water, and buffer solution (20:40:40) | Gradient elution:10 mM aqueous ammonium acetate, pH 3.5 (A) and methanol (B) | Gradient elution:10 mM aqueous ammonium acetate, pH 3.5 (A) and methanol (B)
GAPI
AGREE
BAGI

## Freetext Data
Data | Value | Relevance
The study goal | To find an easy, accurate, and selective way to measure APAP, NAPQI, and NAC | Relevant as it aligns with the research aim of detecting drugs, metabolites, and antidotes in a single run.
Methods developed | HPTLC and HPLC methodologies | Relevant as the methods are central to achieving simultaneous determination of the focus analytes.
Material tested | APAP, NAPQI, NAC in plasma samples | Relevant since the study quantifies drug, toxic metabolite, and antidote from blood plasma.
HPLC mobile phase | Water, methanol, and formic acid (70:30:0.15, v/v/v) | Relevant as the solvent system is critical for the chromatography process optimization.
HPTLC stationary phase | Silica gel 60 F254 | Relevant as the stationary phase facilitates analytes' separation.
HPLC stationary phase | Zorbax® SB C18 (4.6 × 250 mm, 5 μm) | Relevant since it specifies the column used for separation during HPLC analyses.
Wavelength for detection | 254 nm | Relevant because detection wavelength ensures accurate measurements of analytes.
Linear concentration ranges | APAP: 1–30 μg mL−1; NAPQI: 10–200 μg mL−1; NAC: 3–100 μg mL−1 | Relevant as it defines the sensitivity and dynamic range of the methods.
HPTLC mobile phase ratio | Methanol, ethyl acetate, and glacial acetic acid (8:2:0.2, v/v/v) | Relevant for its role in analyte separation optimization on the HPTLC system.
Validation protocols | ICH validation tests | Relevant because validation ensures methods meet international analytical standards.
Sensitivity (LOD/LOQ) | Low LOD and LOQ values | Relevant as it highlights the methods' potential in detecting low concentrations of analytes.
Plasma sample preparation | Vortex, centrifugation, evaporation under nitrogen | Relevant for depicting how analytes were extracted from human plasma for analysis.
Run time (HPLC) | 5 minutes | Relevant since it signifies efficiency and speed of the developed method.
Environmental assessment tools | GAPI and AGREE methods | Relevant as these evaluate the eco-friendliness and sustainability of analytical techniques used.
Analytes quantified | APAP, NAPQI, and NAC | Relevant since the study focuses on determining these substances simultaneously.
HPLC advantage | Higher environmental friendliness, reduced solvent use | Relevant because it offers benefits in terms of sustainability and analytical efficiency compared to HPTLC.
BAGI ratings | HPTLC: 80; HPLC: 82.5 | Relevant as it measures the applicability and practicality of the methods developed.
Human plasma aliquots used | 1 mL plasma samples | Relevant as the study involves plasma-based pharmacokinetic analysis.
Patient population | Pediatric patients primarily considered, no specific number mentioned | Relevant since APAP is frequently prescribed for pediatric patients and their plasma levels were studied.
Sampling preparation | Plasma was spiked with working solutions of APAP, NAPQI, and NAC, centrifuged for 30 min., and evaporated under nitrogen prior to analysis | Relevant because it shows the preparation process for plasma samples used in pharmacokinetic studies.
Sampling time points | Not explicitly mentioned | Relevant as sampling time points are crucial but missing from the provided paper details.
Centrifugation speed | 4000 rpm | Relevant because it describes an essential part of the plasma sample preparation process in the study.
Analyte analysis methods | HPTLC and HPLC-DAD methods | Relevant because the methods were used to measure drug analytes in plasma for pharmacokinetics.
Preconcentration technique | Supernatant evaporated under nitrogen | Relevant because preconcentration steps are integral to analyzing plasma samples in PK studies.
Injection for HPLC | 20 μL injected via auto-sampler | Relevant because injection volumes are part of the analytical method in PK studies.
Model | Not explicitly stated but techniques adapted for APAP, NAPQI, and NAC quantification could align with a linear pharmacokinetic model perspective | The model or approach aligns with understanding analyte concentration dynamics via developed analytical techniques, but a formal compartmental model (like 1-compartment or 2-compartment) wasn't stated specifically in the paper.
Detection Wavelength (HPTLC) | 254 nm | The detection wavelength ensures sensitive quantification for APAP, NAPQI, and NAC in plasma samples.
Detection Wavelength (HPLC) | 254 nm | Utilized for accurate measurement of analytes using a photodiode array detector in the HPLC procedure.
Mobile Phase Composition (HPTLC) | Methanol: Ethyl Acetate: Glacial Acetic Acid (8:2:0.2, v/v/v) | Optimized for effective separation and resolution of analytes in the HPTLC method.
Mobile Phase Composition (HPLC) | Water: Methanol: Formic Acid (70:30:0.15, v/v/v) | Achieves optimal analyte separation in HPLC with minimal tailing.
Flow Rate (HPLC) | 1 mL min−1 | Allows efficient resolution and reduces solvent consumption during HPLC analysis.
Rf Values (HPTLC) | APAP (0.25), NAPQI (0.40), NAC (0.55) | Retardation factors confirm the effective separation and chromatographic performance in HPTLC.
Stock Solution Concentration | 1 mg mL−1 | Used as the basis for preparing various concentrations for both HPTLC and HPLC analysis and calibration.
Concentration Range (HPTLC) | APAP (0.5–6 μg band−1), NAPQI (1–7 μg band−1), NAC (0.3–18 μg band−1) | Important for establishing analyte detection accuracy and linearity in the HPTLC method.
Concentration Range (HPLC) | APAP (1–30 μg mL−1), NAPQI (10–200 μg mL−1), NAC (3–100 μg mL−1) | Linear range for quantifying analytes in HPLC for plasma analysis.
Run Time (HPLC) | 5 min | Highlights the method's efficiency and suitability for routine analysis.
Regression Equations (Plasma, HPTLC) | APAP: Y = 0.7543C + 1.545; NAPQI: Y = 0.432C + 0.243; NAC: Y = 0.254C + 0.466 | Derived regression lines in spiked plasma samples indicate high precision.
LOD and LOQ (HPLC) | Low values achieved | Demonstrate high sensitivity, especially critical for plasma-based analyte quantification.
System Suitability Parameters | Tailing Factor, Capacity Factor, Selectivity, Resolution | Ensures validated performance and reproducibility of chromatographic systems.
Saturation Time (HPTLC) | 20 minutes | Ensures optimal saturation for analyte development in thin-layer chromatography.
Calibration Curve Accuracy | Mean Percent Recoveries Adequate | Verifies established procedures for HPTLC and HPLC are accurate and precise.
Degassing Time (HPLC) | 15 min | Ensures a stable mobile phase for maintaining the performance of the HPLC system.
Approximately 5–10% of APAP undergoes hepatic metabolism by cytochrome CYP 2E1 | Metabolism of APAP is mediated by CYP 2E1 enzyme | This is relevant as the CYP 2E1-mediated metabolism influences the production of NAPQI, a toxic metabolite, thereby affecting APAP pharmacokinetics.
NAPQI detoxification by glutathione (GSH) forms non-toxic cysteine and mercapturic acid | Detoxification pathway through GSH conjugation | This relates to pharmacokinetics by explaining how the body metabolizes and eliminates the toxic intermediate NAPQI.
GSH depletion causes the accumulation of NAPQI in cases of overdose | Overdose affects GSH depletion | This is relevant as GSH depletion alters the pharmacokinetics by limiting NAPQI detoxification and increasing toxic effects.
Restoration of GSH levels using NAC | Utilization of NAC to replenish GSH | Relevant to pharmacokinetics, as NAC affects the metabolic process of detoxifying harmful metabolites like NAPQI.
NAPQI's covalent binding to macromolecules due to GSH depletion | Effect of GSH depletion on cellular proteins | Relevant to pharmacokinetics, since this binding causes cellular damage, altering APAP's effect and disposition.
Reactive oxygen species (ROS) and mitochondrial dysfunction caused by NAPQI | Mitochondrial failure and ROS generation | Shows how pharmacokinetics can be influenced by metabolic byproducts like ROS, affecting drug toxicity and clearance.