## Table 1
In vitro retrodialysis | Probe 1 (0.01% DMSO with 0.5% BSA Ringer) | Probe 2 (0.1% DMSO with 1% BSA Ringer)
1 h | 2.2 | < 1
2 h | 0.9 | < 1
3 h | < 1 | < 1
Stock | 102 | 70
Recovery rate (%) | 1.5% | 1% a

## Table 2
Unnamed: 0 | Measured selinexor concentration (ng/mL) in Ringer's solution | Percent difference to polypropylene tubes
Polypropylene tubes | 2536 ± 175 | 0
Dialysate tubes | 2554 ± 102 | −0.007
Glass | 2109 ± 23 | −20.25

## Table 3
Calibration method | Advantage | Disadvantage | Calculation
In vitro dialysis | No mice needed Easy to control | Differences between in vitro and in vivo | RR = C dialysate/C external (Epi)
II In vivo/vitro retrodialysis | Mass transfer resistance considered | Ensurance of drug‐free brain necessary | RR = (C perfusate—C dialysate)/C perfusate CECF = Cdialysate/recovery
III Internal standard technique | Continuous calibration during experiments | Interaction between drug and standard
IV Endogenous reference method | Continuous calibration during experiments | Different properties or interaction between drug and standard
V Zero/ultra slow flow rate method | Increased recovery rate | Small sample volume | RR = (1 − exp. − PeAQ) × 100
VI (Dynamic) no‐net‐flux method | Well investigated Less time‐consuming (dynamic) | Steady state necessary Large number of animals (dynamic)

## Freetext Data
Data | Value | Relevance
Cerebral microdialysis application | Successfully applied to various drugs but challenging for hydrophobic compounds due to non‐specific binding | Highlights the challenges associated with non‐specific binding of hydrophobic drugs, setting the context for addressing experimental limitations.
Challenges with hydrophobic drugs | Non‐specific binding to microdialysis components and low recovery rates | Demonstrates the need for specific strategies and adaptations for hydrophobic drugs, which are central to the study’s focus.
Study focus | Addressing non‐specific binding challenges for hydrophobic drugs and providing practical recommendations | Directly addresses the study's objective of minimizing issues for hydrophobic compounds in cerebral microdialysis.
Addition of supplements to perfusate | Bovine serum albumin (BSA) effectively reduces non‐specific binding | Key solution presented in the paper for improving drug recovery in hydrophobic drug experiments.
Key hydrophobic drug examples tested | Actinomycin D, Selinexor, Ulixertinib | Demonstrates application of the findings and experimental setup to specific hydrophobic drugs for broader relevance in drug development.
Probe material and performance | PES membrane (CMA7) provided highest recovery for actinomycin D (81.6%) | Practical example of selecting probe materials based on the hydrophobic compound's properties, central to experimental optimization.
Calibration necessity | Retrodialysis and other complementary methods recommended for probe recovery rates | Calibration is crucial for reliable data and underscores the study’s practical recommendations.
Systematic evaluation necessity | Drug adherence to vials, tubing, and other system components must be evaluated | Critical in ensuring robust insights into the challenges of non‐specific binding and consistent microdialysis outcomes.
Flow rate importance | Low flow rates are critical for stable equilibrium, but higher rates may be needed for short time-sampling | Highlights the technical adjustments needed based on experimental objectives and challenges.
Recommended probe membrane MWCO | ≤55 kDa sufficient for small molecular weight compounds | Strategy for adapting microdialysis components for enhanced drug recovery and accuracy.
Calibration methods comparison | Multiple methods advised (e.g., in vivo retrodialysis, no-net-flux) for reliability | Direct response to achieving accurate results in challenging experimental setups.
Drug transport factors influencing outcomes | BBB permeability and physicochemical properties like molecular weight and hydrophobicity | Key underlying factors influencing drug-specific microdialysis challenges discussed in the study.
Limitations in microdialysis studies | Technical limitations like non‐specific binding and sensitivity issues can exclude some drugs | Acknowledges scope of challenges, focusing on the study’s goal to devise solutions for better adaptability.
Model‐related factors | BBB permeability varies with age, recovery periods needed after probe placement | Important for guiding experimental models to align with preclinical and clinical translational aims.
Objective of Kp,uu | To estimate brain drug concentrations and differentiate between transport mechanisms | Central to the study's implications for pharmacokinetics and CNS drug development.
Importance of quantification method | UPLC-MS/MS required for sensitivity in microdialysis studies | Supports the need for advanced analytical techniques for precise data, especially critical for low-concentration drugs.
Decision tree contribution | Systematic steps for experimental setup including membrane, perfusate, calibration, and drug testing | Practical tool developed to improve experimental design for hydrophobic drugs, advancing translational science.
Clinical relevance | Framework translatable to human clinical contexts where microdialysis is already applied | Supports study's translational impact in clinical pharmacology and CNS drug research.
PK study design | "Blood samples collected from each animal to assess plasma PK and protein binding." | "Relevant to understanding how plasma pharmacokinetics (PK) were established to determine drug distribution."
Patient population | "Rodent models used for testing hydrophobic drugs actinomycin D, selinexor, ulixertinib." | "Relevant as rodents serve as the animal model for PK evaluations preclinically."
PK parameters | "Unbound drug concentrations in plasma and brain calculated to assess the unbound plasma‐to‐brain partition coefficient (Kp,uu)." | "Relevant as Kp,uu is essential for elucidating drug distribution across the blood-brain barrier in PK analysis."
Sample time points | "Dialysate samples were collected at 1-hour intervals over the equilibration period using fraction collectors." | "Relevant to temporal assessment of drug distribution at specific intervals during PK studies."
Analytical methods | "UPLC‐MS/MS used to quantify drug concentrations in plasma and dialysate." | "Relevant to ensure sensitive quantification of unbound drug samples in PK analysis."
Calibration | "In vitro and in vivo calibration methods validated recovery rates before in vivo experiments." | "Relevant for accurately determining recovery rates critical for PK measurement reliability."
Recovery period after surgery | "24 hours post probe implantation to minimize tissue damage effects before experiments." | "Relevant as it influences the integrity of PK data by accounting for confounding factors like BBB disruption."
Preconditions for drug study | "Addition of bovine serum albumin (BSA) to minimize non-specific binding in plasma and dialysate samples." | "Relevant for improving the accuracy and reliability of plasma and tissue PK data for hydrophobic drugs."
The study utilized cerebral microdialysis for in vivo experiments to determine unbound drug concentrations in brain tissue. | Physiologically based setup using cerebral microdialysis system components, i.e., probes with semipermeable membranes, Ringer’s solution as perfusate, and in vivo retrodialysis calibration methods. | This physiologically based method aligns with understanding drug distribution and unbound fraction in brain tissue, complementing pharmacokinetic models without directly relying on predefined compartments.
The plasma-to-brain partition coefficient (Kp,uu) serves as the metric calculated from the in vivo recovery and unbound drug concentrations in plasma and brain. | No explicit compartmental model specified, focused on plasma-brain relationships using microdialysis equipment and calibration techniques. | The study's model-free approach focuses on empirical parameter derivation (e.g., Kp,uu), critical for bridging observed data with mechanistic transport across the BBB.
Challenges with hydrophobic drugs were addressed through material calibrations and optimized microdialysis probe recovery rather than theoretical compartmental modeling. | Experimental physiologically based optimization without adherence to standard pharmacokinetic compartment models. | The approach ensures relevant in vivo data collection without dependency on abstract modeling frameworks like one or two compartments.
Drug calibration incorporated recovery methods (e.g., in vitro, retrodialysis) tailored to probe conditions and drug-specific physiochemical traits. | Focused on probe-determined recoveries instead of theoretical projection from compartmental assumptions or PK models. | The calibration methods directly validate physiological drug behavior, bypassing assumptions typical of compartmental or IVIVE setups.
Compound name | Actinomycin D | Relevant because actinomycin D is one of the hydrophobic drugs studied for microdialysis testing.
Compound name | Selinexor | Relevant because selinexor is one of the hydrophobic drugs tested, requiring specific adaptations in the microdialysis system.
Compound name | Ulixertinib | Relevant because ulixertinib is another hydrophobic drug studied, providing insights into experimental challenges and solutions.
Molecular weight | 1255.4 Da (actinomycin D) | Relevant for parameterizing probe membrane choice, influencing drug recovery rates in cerebral microdialysis.
Molecular weight | 443.3 Da (selinexor) | Relevant as it influences membrane efficiency and non-specific binding properties.
Molecular weight | 433.3 Da (ulixertinib) | Relevant for choosing compatible probe materials to optimize recovery rates.
MWCO for actinomycin D | 55 kDa (CMA7) | Relevant as it achieved high retrodialysis recovery rates of ~81% for actinomycin D, critical for probe setup.
MWCO for selinexor | 55 kDa (CMA7) | Relevant because it provided similar recovery rates (~80%), showing membrane compatibility with drug size and hydrophobic properties.
MWCO tested for actinomycin D | 20 kDa, 35 kDa, and 55 kDa | Relevant for determining the optimal MWCO for each compound to achieve reliable recovery.
MWCO tested for ulixertinib | 55 kDa (low recovery observed: 1.5%) | Relevant for showing that standard materials were unsuitable for this compound.
Probe material for actinomycin D | CMA7 (PES membrane) | Relevant as PES membranes reduced non-specific binding for actinomycin D.
Incubation temperature | 37°C | Relevant because drug stability and perfusate stability were tested under this condition to simulate in vivo setups.
Flow rate | 0.5 μL/min | Relevant as the flow rate was tested and optimized to assess drug recovery and avoid tubing and probe clogging.
Recovery rate (actinomycin D) | 81.6% ± 9.9% (CMA7 probe) | Relevant for validating the combination of probe type and perfusate in microdialysis experiments.
Recovery rate (selinexor) | 80.0% ± 7.25% (CMA7 probe) | Relevant for selecting probe membranes providing optimal recovery for drug quantification.
Recovery rate (ulixertinib) | 1.5% (CMA7 probe) | Relevant for identifying challenges in using standard microdialysis probe materials.
Perfusate solution | 1% BSA Ringer's solution | Relevant as it minimized non-specific binding to system components for hydrophobic drugs.
Additional solution (ulixertinib) | 0.01%-0.1% DMSO with 0.5%-1% BSA Ringer's solution | Relevant for improving recovery and reducing non-specific binding with ulixertinib.
Non-specific binding mitigating agent | 1% BSA (actinomycin D, selinexor) | Relevant as BSA addition directly addressed the challenge of non-specific binding to vials and tubing.
Alternative tubing materials tested | FEP and PEEK | Relevant as these tubing materials were evaluated for reduced binding with hydrophobic drugs.
Optimal tubing material | FEP (ulixertinib, selinexor) | Relevant as it demonstrated lower non-specific binding compared to PEEK.
Calibration methods | Retrodialysis, in vitro dialysis, no-net-flux method | Relevant for determining recovery rates critical for accurate microdialysis outcomes.
Temperature stability testing | RT, 37°C, −20°C, 4°C (actinomycin D) | Relevant for confirming stability under experimental and environmental conditions.
Quantification method | UPLC-MS/MS | Relevant for detecting low-concentration drug recovery and ensuring accuracy in pharmacokinetics measurements.
Photostability testing | Light and darkness conditions (actinomycin D) | Relevant to confirm experimental suitability for photosensitive compounds.
Setup material for photosensitivity | Aluminum foil, opaque tubing | Relevant to mitigate degradation during drug testing for photosensitive compounds.
Transport tubing swept volume | 1.2 μL per 10 cm (for FEP and PEEK) | Relevant as it determined compatibility with flow rates for optimal recovery.
Recovery rate reduction factor | Due to flow rate variability | Relevant for experimental setup design to ensure consistent calibration results.
Analytical sensitivity | Lower limit of quantification: 0.1 ng/mL | Relevant because sensitivity impacted the successful recovery rate evaluation of hydrophobic drugs.
Stability of drugs under experimental conditions | Actinomycin D was photosensitive but stable at room temperature, 37°C, 4°C, and −20°C under dark and light conditions | Stability influences pharmacokinetics by ensuring accurate measurements of drug concentrations under varying experimental conditions.
Drug solubility in aCSF/microdialysis buffer | Selinexor required 1% BSA in Ringer's solution to reduce non-specific binding | Solubility impacts the ability of drugs to maintain consistent pharmacokinetic profiles during experiments.
Molecular weight of drugs | Actinomycin D: 1255.4 Da, Selinexor: 443.3 Da, Ulixertinib: 433.3 Da | Drug size affects diffusion through microdialysis membranes and influences drug recovery rates, a key pharmacokinetic aspect.
Non-specific binding in setup materials | Selinexor adhered to glass and FEP/PEEK tubing; actinomycin D adhered to polypropylene vials | Non-specific binding alters the apparent drug concentration, thus influencing pharmacokinetic measurements.
Addition of stabilizing agents to reduce binding | 1% BSA added for Actinomycin D and Selinexor to improve recovery rates | Binding modifications directly affect the accurate determination of pharmacokinetics by stabilizing drug concentrations.
Membrane material and MWCO | A 55-kDa MWCO (CMA7) suited actinomycin D and selinexor; unsuitable for ulixertinib due to low recovery rate | Membrane properties interact with drug molecular characteristics to influence diffusion and recovery, critical for pharmacokinetics.
Flow rate during microdialysis | Low flow rates optimized recovery rates for hydrophobic drugs | Flow rate determines the equilibrium diffusion of drugs, thus influencing pharmacokinetic parameters in microdialysis.
Drug-specific adhesion evaluation | Ulixertinib showed no evidence of binding to tested vials | Adhesion properties of specific drugs define the stability and quantifiability within systems, crucial for pharmacokinetics.
Calibration models for probe recovery rates | Retrodialysis and zero flow rate for drugs with slow diffusion | Calibration ensures precise probe recovery and valid pharmacokinetics assessment.
Age-related BBB differences | Age-related permeability changes in rodents noted as considerations | Variations in BBB integrity with age impact drug pharmacokinetics and must be accounted for in experimental designs.
System component material | FEP tubing best recovered high hydrophobic drugs like selinexor and ulixertinib | Material choices affect pharmacokinetics by minimizing interaction losses.