## Table 1
Time (min) | Eluent C (%) | Eluent D (%) | Flow rate (μL)
0 | 40 | 60 | 400
0.5 | 40 | 60 | 400
2 | 5 | 95 | 400
3 | 5 | 95 | 400
3.1 | 40 | 60 | 400
5.5 | 40 | 60 | 400

## Table 2
Analyte | Monoisotopic mass | Collision energy | Used transitions | Retention time (minutes) | Signal intensity (E)
Dolutegravir | 419.13 | 25 | 277 | 3.05 | ≥3
Dolutegravir-d4 | 423.13 | 27 | 280.1 | 3.03 | ≥3
Efavirenz | 315.03 | 2 | 316 | 4.25 | ≥3
Nevirapine | 266.16 | 2 | 267.1 | 2.57 | ≥3
Rifampicin | 822.41 | 35 | 790.2 | 3.36 | ≥3
Rifampicin-d4 | 826.41 | 18 | 795.2 | 3.26 | ≥3
Rifapentine | 876.45 | 17 | 845.5 | 3.73 | ≥3

## Table 3
STD and QC | Plasma (μL) | Working Solution (μl) | DTG, EFV, NVP, RMP, RPT (μg/ml)
BLANK | 1000 | 0 | 0
STD1 | 995 | 5WS2 | 0.25
STD2 | 990 | 10WS2 | 0.5
STD3 | 950 | 50WS2 | 1
STD4 | 995 | 5WS1 | 2.5
STD5 | 975 | 25WS1 | 5
STD6 | 950 | 50WS1 | 10
QCL | 985 | 15WS2 | 0.75
QCM | 920 | 80WS2 | 4
QCH | 960 | 40WS1 | 8

## Table 4
Analyte | Recovery (%) (n=6) | PASS/FAIL*
Dolutegravir | 86.12 | Pass
Efavirenz | 88.88 | Pass
Nevirapine | 109.89 | Pass
Rifampicin | 103.65 | Pass
Rifapentine | 100.48 | Pass

## Table 5
Unnamed: 0_level_0_Analytes | Unnamed: 1_level_0_QC levels | Unnamed: 2_level_0_Referred concentration (μg/ml) | Precision (%RSD)_Intraday (n=6) | Precision (%RSD)_Interday (n=6) x3days | Accuracy_Intraday (n=6) | Accuracy_Interday (n=6) x3days
Dolutegravir | LLOQ | 0.25 | 7.73 | 14.41 | 101.87 | 99.56
Dolutegravir | Low | 0.75 | 9.86 | 11.57 | 98.11 | 97.54
Dolutegravir | Medium | 4 | 11.61 | 10.04 | 100.59 | 102.72
Dolutegravir | High | 8 | 9.92 | 7.55 | 99.64 | 99.34
Efavirenz | LLOQ | 0.25 | 7.69 | 9.12 | 103.4 | 100.4
Efavirenz | Low | 0.75 | 5.44 | 7.18 | 99.87 | 100.49
Efavirenz | Medium | 4 | 8.52 | 10.74 | 103.32 | 101.4
Efavirenz | High | 8 | 9.22 | 99.58 | 102.25 | 99.58
Nevirapine | LLOQ | 0.25 | 5.5 | 9.81 | 109.67 | 100.84
Nevirapine | Low | 0.75 | 9.8 | 9.94 | 106.6 | 102.36
Nevirapine | Medium | 4 | 6.79 | 5.42 | 104.36 | 99.89
Nevirapine | High | 8 | 3.79 | 6.23 | 101.19 | 96.92
Rifampicin | LLOQ | 0.25 | 12.39 | 14.72 | 106.27 | 98.13
Rifampicin | Low | 0.75 | 9.13 | 10.23 | 103.76 | 101.55
Rifampicin | Medium | 4 | 2.47 | 5.34 | 109.67 | 100.31
Rifampicin | High | 8 | 6.07 | 5.68 | 99.02 | 93.38
Rifapentine | LLOQ | 0.25 | 4.47 | 16.83 | 88.73 | 103.87
Rifapentine | Low | 0.75 | 9.31 | 7.94 | 107.73 | 104.3
Rifapentine | Medium | 4 | 4.01 | 6.58 | 95.69 | 101.11
Rifapentine | High | 8 | 7.24 | 10.44 | 96.25 | 97.43

## Table 6
Unnamed: 0_level_0_Analytes | Unnamed: 1_level_0_QC level | Unnamed: 2_level_0_Referred value (μg/ml) | Stability (%RSD)_Short-term (28 days) | Stability (%RSD)_Freeze-thaw (n=6) x3days | Stability (%RSD)_Autosampler (n=6)
Dolutegravir | LLOQ | 0.25 | 7.73 | 14.41 | 8.18
Dolutegravir | Low | 0.75 | 9.86 | 11.57 | 3.78
Dolutegravir | Medium | 4 | 11.61 | 10.04 | 8.65
Dolutegravir | High | 8 | 9.92 | 7.55 | 6.54
Efavirenz | LLOQ | 0.25 | 7.69 | 9.12 | 8.46
Efavirenz | Low | 0.75 | 5.44 | 7.18 | 7.6
Efavirenz | Medium | 4 | 8.52 | 10.74 | 11.84
Efavirenz | High | 8 | 9.22 | 8.87 | 10.93
Nevirapine | LLOQ | 0.25 | 5.5 | 9.81 | 12.05
Nevirapine | Low | 0.75 | 9.8 | 9.94 | 8.94
Nevirapine | Medium | 4 | 6.79 | 5.42 | 5.99
Nevirapine | High | 8 | 3.79 | 6.23 | 5.19
Rifampicin | LLOQ | 0.25 | 12.39 | 14.72 | 6.97
Rifampicin | Low | 0.75 | 9.13 | 10.23 | 4.54
Rifampicin | Medium | 4 | 2.47 | 5.34 | 5.88
Rifampicin | High | 8 | 6.07 | 5.68 | 10.15
Rifapentine | LLOQ | 0.25 | 4.47 | 16.83 | 5.12
Rifapentine | Low | 0.75 | 9.31 | 7.94 | 8.58
Rifapentine | Medium | 4 | 4.01 | 6.58 | 13.3
Rifapentine | High | 8 | 7.24 | 10.44 | 10.25

## Freetext Data
Data | Value | Relevance
Therapeutic Drug Monitoring (TDM) | Limited in low-and-middle income countries due to lack of affordable analytical methods | Essential for optimizing treatment outcomes among HIV-TB co-infected patients; underscores the need for accessible tests in these regions.
Bioanalysis Technique | LC-MS method for simultaneous determination of dolutegravir, efavirenz, nevirapine, rifampicin, isoniazid and rifapentine | Conventional chromatographic platforms less effective; LC-MS provides specificity and sensitivity for appropriate pharmacokinetics in TDM.
Challenges in HIV-TB Treatment | Drug-drug interactions, maladherence, stigma, inter-individual variability in drug metabolism | Polypharmacy hinders the attainment of therapeutic targets, making TDM crucial in these co-infected populations.
Validation Parameters | Accuracy ranged from 88.73%-109.67%, intra-day precision 2.47-12.39 RSD | Demonstrates robust method reliability required to gauge clinically relevant concentrations for TDM.
Antiretroviral Drugs | Dolutegravir, Efavirenz, Nevirapine (First-line drugs for PLWHIV) | Their plasma concentration measurement is necessary for monitoring adverse effects and therapeutic efficacy in co-treatment.
Anti-TB Drugs | Rifampicin, Isoniazid, Rifapentine (First-line TB therapy options) | Accurate plasma-level measurement aids in optimizing dosing to improve treatment success and minimize failure outcomes.
Method Sensitivity | LLOQ lower limit sensitivity at 0.25 μg/ml (for all analytes) | Sufficient sensitivity to detect therapeutic plasma levels and ensure patients are not exposed to subtherapeutic concentrations.
Analyte Extraction | Protein precipitation with acetonitrile at varying ratios | Ensures effective separation and recovery of antiretrovirals and anti-TB drugs from plasma samples in the method validation process.
Gradient Composition | Mobile phase with 0.1% formic acid in water and acetonitrile under varying ratios | Designed to optimize the elution of drugs while maintaining stability, key for simultaneous determination.
Proposed Use | Uganda Ministry of Health for patient TDM | Offers cost-effective, validated analytical framework suitable for public healthcare facilities combating HIV-TB co-infection.
Drug Recovery Efficiency | Range: 86.12%-109.89% (all analytes) | Confirms the accurate retrieval of drugs from plasma, reflecting the reliability of the sample preparation method.
Stability | Long-term and freeze-thaw stability validated for all analytes | Affirms suitability of this method for routine laboratory practice, including extended storage periods of patient samples.
Accuracy and Precision | Inter-day accuracy: 93.38%-104.30% | Validates the method's reliability for ensuring consistent measurement of drug concentrations over multiple runs.
Method Limitations | Exclusion of Isoniazid from validated analytes due to inconsistent peaks | Highlights scope for future method improvements, particularly for including polar molecules like isoniazid.
Cost Effectiveness | Saves time and reagents via simultaneous drug analysis | Reduces operational costs and enhances feasibility for resource-constrained health systems.
Sample source | De-identified fresh frozen plasma from Mulago hospital blood bank | The study leveraged plasma samples from human donors to prepare standards and quality control samples, indicating the foundational matrix for pharmacokinetic analysis.
Sampling batches | Plasma obtained in three batches on 14/03/2024, 08/04/2025 and 06/05/2024 | Relevant because these plasma sampling dates define batch acquisition timelines for consistency in sample handling.
Patient population | Focused on HIV-TB co-infected patients | Relevant for assessing HIV-TB pharmacokinetics because this population faces challenges like polypharmacy and drug-drug interactions.
Matrix | Human plasma | Plasma was chosen over other matrices like saliva or urine due to reduced influence from secondary organs and demographics like age, gender, which is critical for accurate pharmacokinetic evaluation.
Sampling methodology | Protein precipitation achieved using 100% acetonitrile | Relevant because the extraction methodology directly impacts the accuracy and reproducibility of drug concentration measurements.
Sampling process | Blank plasma used for calibration, standards, and quality controls | Relevant for ensuring the reliability and accuracy of pharmacokinetic measurements.
Injection volume | 10μL for rifampicin and rifapentine and 30μL for dolutegravir, efavirenz, nevirapine | Relevant because injection volumes signify method sensitivity and operational design critical for pharmacokinetic studies.
Stability testing | Conducted freeze-thaw stability cycles and long-term storage stability of plasma | Relevant as stability testing determines reliability in tracking drug concentrations over timepoints.
Analyte separation | Achieved by gradient elution on Waters Atlantis dC18 column | Relevant because the chromatographic separation impacts the detection and quantification of drug concentrations critical for pharmacokinetics.
Study approval | Makerere University College of Health Sciences ethics board (Approval number: SBS-2022–164) | Relevant for ensuring ethical compliance in sourcing plasma samples from human donors.
Model type | Not stated | The paper does not explicitly mention the use of a pharmacokinetic model such as 1-compartment, 2-compartment, physiologically based model, or IVIVE, thus lacks relevance to modeling-specific approaches.
Injection Volume | 10μL for rifampicin and rifapentine; 30μL for dolutegravir, efavirenz, and nevirapine | Relevant as it defines the sample size delivered to the LC-MS system during analysis, which is critical for achieving consistent measurements.
Column Temperature | 18°C to 26°C | Relevant to ensure optimal analyte separation based on thermal stability, as low temperatures prevent analyte degradation.
Mobile Phase Composition | Eluent C: 0.1% formic acid in water; Eluent D: 0.1% formic acid in 100% acetonitrile | Relevant because this chemistry establishes the environment for effective separation by the column during analysis.
Gradient Flow Rate | 400μL/min | Relevant since the determined flow rate directly impacts the analyte separation and run time efficiency.
Spray Voltage | 5.00 V | Relevant to the ionization of target analytes in the mass spectrometer, determining accurate mass-to-charge ratios.
Heater Temperature | 170°C | Relevant because it ensures the desolvation of analytes in ion spray, aiding source stability during ionization.
Capillary Temperature | 275°C | Relevant in thermally stabilizing ions during transportation from the spray source to the detector.
Capillary Voltage | 34.00V | Relevant for optimizing ion transfer to the mass detector, contributing to precise analyte identification.
Tube Lens Voltage | 90.00V | Relevant to facilitating focused ion transport through the mass spectrometer aperture system, aiding detection sensitivity.
Sheath Gas Flow Rate | 29 arb | Relevant as it assists in reliable desolvation and droplet dispersion during ionization.
Auxiliary Gas Flow Rate | 5 arb | Relevant since it prevents significant analyte condensation in spray, ensuring ionization consistency.
Precursor Ion Transition | Dolutegravir: 419.13→277.00; Efavirenz: 315.03→316.00; Nevirapine: 266.16→267.10; Rifampicin: 822.41→790.20; Rifapentine: 876.45→845.50 | Relevant because these transitions are used in MRM mode to specifically measure analytes of interest by their mass and fragment properties.
Stock Solution Concentrations | Dolutegravir: 10,000 μg/ml; Nevirapine, Efavirenz, Rifampicin, Rifapentine: 10,000 μg/ml | Relevant for calibration and validation steps as stock solutions define the concentration range for constructing calibration curves.
Working Solution Concentrations | Working Solution 1: 1000 μg/ml; Working Solution 2: 50 μg/ml | Relevant since working solutions prepare quality control samples, impacting method precision and accuracy.
Standard Curve Range | 0.25μg/ml to 10.00μg/ml | Relevant for assessing linearity and sensitivity during calibration and validation phases.
Method Recovery | 86.12% to 109.89% | Relevant in determining the efficiency of analyte extraction methods, which affects quantitative accuracy.
Accuracy | Intra-day: 88.73% to 109.67%; Inter-day: 93.38% to 104.30% | Relevant for assessing method reliability in consistently measuring concentrations within validation ranges.
Precision | Intra-day: 2.47 RSD to 12.39 RSD; Inter-day: 5.34 RSD to 16.83 RSD | Relevant as it identifies the consistency of results across time, critical for robust validation.
Lower Limit of Quantification (LLOQ) | 0.25μg/ml | Relevant as this parameter establishes sensitivity and the minimum analyte level at which accurate quantification is feasible.
Collision-Induced Dissociation Energy | Optimized for individual target analytes during MRM scan | Relevant since CID energy affects the ion fragmentation patterns, ensuring precise analyte detection.
Retention Time | Dolutegravir, Efavirenz, Nevirapine, Rifampicin, and Rifapentine peaks verified against blank plasma samples | Relevant for validating method selectivity and specific detection of analytes without interference.
Analyte Stability | Freeze-thaw (5.34 to 16.83), Long-term (-80C, 28 days), Autosampler (3.78 to 13.30) | Relevant for confirming analyte preservation across different conditions and timeframes, ensuring reliable measurements.
Inter-individual genetic variabilities | Presence among individuals | Genetic variability influences how drugs are metabolized, affecting pharmacokinetics outcomes.
Poor medication practices | Mentioned as a factor | Poor practices alter drug absorption, distribution, and clearance, directly affecting pharmacokinetics.
Presence of other morbidities | Reported in study | Comorbidities can modify drug metabolism and excretion, impacting pharmacokinetics in patients.
Drug-drug interactions | Highlighted in polypharmacy | Interactions can result in altered drug levels due to changes in metabolism or binding.
Drug-related toxicities | Cited as a limitation | Toxicities may necessitate dose adjustments, impacting pharmacokinetics parameters.
Polypharmacy | Observed in HIV-TB co-treatment | Concurrent drug usage can interfere with drug metabolism, influencing pharmacokinetics.
Lower drug concentrations | Empirical findings in HIV-TB patients | Reduced absorption or increased metabolism results in subtherapeutic levels, relevant to pharmacokinetics.
HIV-TB co-infection | Core analysis population | Co-infection creates complex pharmacokinetics due to interaction of disease-mediated effects on drug metabolism.
Stigma and maladherence | Related to HIV-TB treatment | These social factors lead to erratic drug administration, altering pharmacokinetics profiles.
Therapeutic drug monitoring | Proposed as solution | Helps understand variability in pharmacokinetics by measuring plasma drug concentrations.
Drug concentration measurement | Objective of study | Concentrations are fundamental in evaluating pharmacokinetics efficiency.
Plasma preference for TDM | Analytical framework | Plasma is less influenced by demographic factors, ideal for pharmacokinetics studies.
Drug recovery rates | 86.12% to 109.89% | Recovery efficiency reflects effectiveness of extraction and indicates pharmacokinetic stability.