## Table 1
Sequence | Mean T1 over 20 mm region at isocentre (ms) | Standard deviation (ms) | % difference from baseline
2D | 816 | 17 | 0.5
3D | 739 | 13 | 9

## Table 2
Unnamed: 0_level_0_Constant flow velocity (cm s−1) | 3D_% difference from baseline | 3D_Standard deviation (ms) | 2D_% difference from baseline | Unnamed: 4_level_0_Standard deviation (ms)
0 | 0 | 17 | 0 | 43
3.2 | 3.7 | 17 | 4.9 | 39
6.3 | 5.6 | 19 | 7.7 | 37
9.5 | 5.9 | 20 | 6.5 | 41
12.6 | 1.6 | 19 | 0.1 | 35
15.8 | −5.7 | 18
31.6 | −37.3 | 12
47.4 | −56.7 | 7
78.9 | −72.8 | 4
94.7 | −77.6 | 3

## Freetext Data
Data | Value | Relevance
The relationship between flow velocity and T1 measurement | Non-linear | This is critical for understanding how flow velocities affect T1 quantifications in measuring AIF in dynamic conditions.
Impact of pulsatile flow on T1 measurement | Increased measurement variance | Pulsatile flow mimics in vivo arterial conditions, which are vital for assessing its impact on T1 and subsequent AIF measurement.
Errors introduced by flow in T1 measures | Large and non-linear | Highlights fundamental challenges in using T1-based AIF measurements in flowing blood under realistic physiological conditions.
Pulsatile flow waveforms | Representative of carotid (31% mean-to-peak) and femoral arteries (12% mean-to-peak) | Pertinent to simulating arterial conditions for validating measurement techniques or errors.
Baseline T1 of blood-mimicking fluid (BMF) | 812 ms | Provides a reference point for analyzing deviations caused by flow effects in both the physical and virtual phantom experiments.
Techniques to mitigate flow effect | ROIs far from flow entry points; pulse sequence adjustments (higher flip angles); post-AIF acquisition | Methods to improve accuracy and reliability in AIF quantification under dynamic flow conditions.
Limitations of current AIF methods | Inaccuracy; reliance on population AIF; potential for relative vs. absolute measures | Explores why current methods do not address patient-specific arterial input function variations effectively.
Validation tools used | Physical and virtual phantoms | Demonstrates accurate control and replication of flow scenarios to validate errors and test hypotheses about measurement conditions.
T1 measurement sensitivity | High sensitivity to flow, flip angles, and pulsatile features | Addresses critical concerns in design adjustments for dynamic flow-sensitive sequences in DCE-MRI settings.
Recommendations for sequence design | Use well-mixed blood post-main DCE-MRI; utilize slower imaging sequences; higher flip angles | These are suggestions to minimize inflow effects and improve T1 measurement accuracy for real-world applications.
Errors in using pre-bolus experiments | Complicated by flow-related biases during bolus passage | Identifies practical challenges in using pre-bolus experiments to improve AIF derivation under realistic flow dynamics.
Clinical applications and relevance | Questioning absolute quantitative measures of vascular function using AIF; promoting focus on better tissue-based AIF approaches | Directly impacts DCE-MRI workflows aiming for standardization across clinical and trial applications.
Repeatability with population AIF | Better than individually measured AIF in some trials | Justifies continued reliance on population-based models despite potential limitations in individual accuracy.
Virtual phantom capabilities | Simulating a range of T1 values and flow velocities; validating errors | Essential for exploring the broader impact of in-flow effects under different conditions in controlled simulations.
Baseline T1 measurement method | Inversion recovery sequence using 19 inversion times between 0 and 4000 ms | Provides methodology used to establish baseline T1 of fluid mimicking blood; essential to simulate PK dynamics.
Physical phantom setup | Tube of 6.35 mm internal diameter mimicking major artery; attached to MR-compatible programmable syringe pump | Designed to simulate physiological blood flow for evaluating errors in T1 measurement affecting PK studies.
Blood mimicking fluid composition | 4:5 mix of glycerol and water with CuSO4 solution; T1 of 812 ms | Mimics human blood properties, vital for realistic simulation of AIF measurement errors in PK modeling.
Flow velocities tested | Continuous velocities (0-95 cm/s) and pulsatile waveforms mimicking femoral (~10% peak velocity) and carotid (~31% peak velocity) arteries | Reflects physiologically realistic flow conditions encountered during AIF sampling in PK studies.
Imaging sequences used | 3D and 2D spoiled gradient echo sequences with baseline estimates obtained at two flip angles | Critical for dynamic T1 measurement over time, a key factor in PK modeling derived from contrast agent kinetics.
Flip angles tested | 3° and 16° (reduced error) vs. 3.5° and 7° sequences (higher error) | Evaluates sequence performance under study conditions to refine patient-specific AIF measurement methodology in PK studies.
Dynamic T1 estimation method | VOIs used for averaging signals; Fram equation applied for T1 estimation | Provides critical step for accurate AIF data integration into PK modeling.
Pulsatile flow testing | Simulated with physiological waveforms from femoral and carotid arteries | Directly explores flow conditions impacting blood sampling timing and accuracy in PK studies.
Errors due to flow velocity and pulsatility | Non-linear increase in T1 measurement error with flow velocity and pulsatile effects | Highlights constraints in accurately measuring AIF under varied blood flow conditions in PK studies.
Virtual phantom validation | Simulated dynamic flow with Parker population AIF at 1/10th dose | Enables extrapolation of findings to in vivo PK studies, particularly pre-bolus approaches.
Virtual Phantom | Physiological Flow Model | The virtual phantom was designed to simulate blood flow through a major artery under physiologically realistic pulsatile waveforms, which directly impacts the imaging and AIF measurement process.
Pulsatile Flow Simulation | Physiologically Relevant Waveforms for Carotid and Femoral Arteries | Pulsatile flow modeling was included to reflect the actual dynamics of blood flow in major arteries, crucial for analyzing errors in T1 measurement affecting the AIF.
Pre-bolus AIF Simulation | Parker Population AIF | A scaled Parker population AIF governed the simulation of dynamic T1 variation during the passage of a contrast agent, establishing its suitability for analyzing errors during pre-bolus scenarios.
Signal Dynamics | Bloch Equations-Based Signal Simulation | Used to generate signals under different flow velocities, effectively modeling the non-linear error propagation during dynamic T1 measurements for the AIF evaluation process.
Baseline T1 of BMF | 812 ms | Relevant as this value was obtained using an inversion recovery sequence and is considered the true T1 for subsequent analysis.
Baseline T1 measurement method | Inversion recovery sequence at 1.5 T | Relevant because it establishes the method used to calculate the baseline T1 value of the blood-mimicking fluid, which serves as a standard reference.
Physical phantom details | 6.35 mm internal diameter tubing | Relevant for replicating the diameter of major arteries and ensuring physiological accuracy in the experiments.
Flow velocities tested (constant) | 0, 3, 6, 9, 13, 16, 32, 47, 63, 79, 95 cm.s−1 | Relevant because these velocities span physiological ranges and form the basis of the study's analysis of flow effects.
Pulsatile flow parameters | Waveforms with mean velocities of 31% (carotid artery) and 12% (femoral artery) of peak flow | Relevant to simulate physiological pulsatile conditions in major arteries, which are a focus of the study.
Baseline flip angles for 3D sequence | 3.5° and 7° | Relevant because these values are used in dynamic imaging to measure T1 of blood-mimicking fluid under flowing conditions and assess signal consistency.
Baseline flip angles for 2D sequence | 5° and 10° | Relevant since these parameters directly influence the imaging outcomes for T1 estimation in dynamic studies.
TR for 3D sequence | 3.0 ms | Relevant as it determines the time between successive radiofrequency pulses, influencing the steady state and overall signal quality in the imaging sequence.
TR for 2D sequence | 5.5 ms | Relevant because it impacts the imaging signal and time-dependent effects during the acquisition of 2D images, affecting T1 measurement accuracy.
Correction factor estimation method | Based on measured BMF signal data | Relevant as the correction accounts for B1 field inhomogeneities, critical for isolating flow-related measurement errors.
Programming of physical phantom | Using Compuflow 1000MR for pulsatile and constant flow | Relevant as it controls the flow patterns, critical for replicating physiological arterial conditions.
Virtual phantom parameters | Range of flow velocities, underlying blood T1 values, and flip angles | Relevant since the virtual phantom simulations are validated against the physical phantom data, allowing broader parameter exploration.
Flip angle correction factor profile | Derived from phantom images corrected for B1 inhomogeneities | Relevant as it ensures accurate interpretation of T1 values under flow conditions, highlighting measurement biases caused by inflow effects.
Gadolinium concentration simulation | Parker population AIF scaled by 1/10th | Relevant for simulating pre-bolus experiments to understand flow-induced errors under realistic contrast dynamics.
Contrast agent's effect on T1 | Reduced T1 makes measurement less prone to flow errors | Relevant because this informs sequence design and the accuracy of T1 measurements in vivo when contrast dynamics are involved.
T1 measurement error assessment | Validated by physical phantom and simulated with virtual phantom | Relevant as it establishes the degree of error introduced under different physiological conditions and flow velocities.
Blood flow velocity | Non-linear relationship between flow velocity and measured T1 | Flow velocity influences the accuracy of measured arterial input function (AIF), which is necessary for pharmacokinetic modeling of contrast uptake.
Pulsatile flow | Increased standard deviation in T1 measurements with pulsatile flow | Variations in pulsatile flow introduce measurement variability, complicating accurate pharmacokinetic modeling.
BMF viscosity | 4:5 mix of glycerol and water | BMF viscosity was designed to mimic blood, potentially influencing how the phantom measurements relate to in vivo blood pharmacokinetics.
Contrast agent concentration | Errors in measuring high concentrations of contrast agent in flowing blood | Accurate concentration measurements are critical for pharmacokinetics, and flow introduces non-linear measurement errors.
Flip angles | Influence of flip angles in mitigating signal error due to flow | Sequence design with higher flip angles reduces inflow effects, impacting the reliability of AIF for pharmacokinetic models.
Presence of magnetic field (B0) history | Blood entering the imaging volume has been in a B0 field for >4 x T1 | Spin polarization history affects T1 measurements, which are foundational for deriving pharmacokinetic parameters.
Patient variability | Natural variations in AIF observed between patients and sessions | Variability affects the reliability of pharmacokinetics modeling when using population-averaged AIFs instead of patient-specific data.
Temporal resolution of imaging sequences | Faster imaging sequences lead to higher variability in T1 due to pulsatile flow | Temporal resolution influences the error margin in pharmacokinetic measures, especially during bolus passage.
Pre-bolus approach | Use of a scaled pre-bolus to approximate main AIF | Affects pharmacokinetic modeling accuracy as the pre-bolus measurement strategy is influenced by flow dynamics.
Reference tissue approaches | Investigated for alternative AIF estimation | Contraindications between contrast dynamics may affect consistency in pharmacokinetics using reference tissue AIF methods.