## Table 1
Unnamed: 0 | Gentamicin | Gentamicin.1 | Gentamicin.2 | Gentamicin.3 | Gentamicin.4 | Gentamicin.5 | Gentamicin.6 | Gentamicin.7 | Vancomycin | Vancomycin.1 | Vancomycin.2 | Vancomycin.3 | Vancomycin.4 | Teicoplanin | Teicoplanin.1 | Doripenem | Doripenem.1
Membrane | SF21E | SF17UX | SF21UX | F80s | F80s | Fx80 | Fx80 | CAHP-170 | F60s | F80s | F80s | F80s | Fx80 | F60s | Fx80 | SF21UX | Fx80
BFR (ml/min) | 400 | 400 | 400 | 300 | 400 | 300 | 400 | 400 | 250 | 250 | 300 | 400 | 300 | 250 | 250 | 300 | 300
DFR (ml/min) | 800 | 800 | 800 | 500 | 800 | 500 | 800 | 800 | 500 | 500 | 500 | 800 | 500 | 500 | 500 | 500 | 500
Cl (ml/min) | 131.52 | 160.75 | 178.25 | 122.37 | 137.62 | 142.13 | 187.90 | 90.12 | 71.20 | 83.35 | 93.61 | 103.25 | 91.11 | 25.50 | 43.49 | 155.76 | 141.25
SD | 16.04 | 26.09 | 46.13 | 9.95 | 12.15 | 6.03 | 13.24 | 31.55 | 4.91 | 19.81 | 31.11 | 18.34 | 7.14 | 17.7 | 23.06 | 18.65 | 22.92
N | 12 | 12 | 17 | 12 | 12 | 12 | 12 | 16 | 16 | 8 | 12 | 12 | 12 | 16 | 12 | 12 | 12

## Table 2
Unnamed: 0 | Gentamicin | Gentamicin.1 | Vancomycin | Vancomycin.1 | Teicoplanin a | Doripenem
Membrane | CAHP-210 | F80s | F60s | F80s | F60s | SF21UX
Cl in vitro [ml/min] | 108.34 b | 137.62 | 71.20 | 93.61 | 38.95 | 155.76
SD [ml/min] | 24.32 b | 12.15 | 4.91 | 31.11 | 26.73 | 18.65
Cl in vivo [ml/min] | 104.03 Sowinski et al. (2008) | 116 Amin et al. (1999) | 73.0 Lanese et al. (1989) | 85.2 Lanese et al. (1989) | 39.70 Thalhammer et al. (1997) | 148.61 Vossen et al. (2018)
SD [ml/min] | 12.44 | 9 | 5.0 | 7.0 | 24.50 | 8.41

## Table 3
Unnamed: 0 | Vancomycin | Doripenem
Human | 41.5% Butterfield et al. (2011) | 8.0% HORI et al. (2006)
Bovine | 43.6% | 9.7%

## Freetext Data
Data | Value | Relevance
Clearance values for gentamicin | Range: 90.12 ml/min (Baxter CAHP-170 diacetate) to 187.90 ml/min (Fresenius Fx80 polysulfone, blood flow 400 ml/min, dialysate flow 800 ml/min) | Demonstrates variability in drug elimination based on membrane material, critical for optimizing drug dosing during hemodialysis.
Clearance values of Vancomycin vs Gentamicin | Fx80 polysulfone: Gentamicin 137.62 ml/min vs Vancomycin 103.25 ml/min | Highlights how different drugs interact uniquely with membrane material during dialysis, essential for tailoring pharmacokinetics to patient needs.
Drug concentration decline in in vitro setup | Faster than in vivo; time multiplication factor of 11.36 aligns curves | Clarifies how in vitro experiments simulate but differ from human conditions, aiding interpretation of clearance data.
Measured drugs | Gentamicin, vancomycin, teicoplanin, doripenem | Indicates drugs studied for clearance profiling, relevant for targeting recommendations for hemodialysis patients.
Dialysis parameters | Blood flow rate: 300 ml/min; dialysate flow rate: 500 ml/min | Provides context for standardized experimental conditions and their influence on clearance rates.
Impact of membrane evolution on clearance | Membranes progressed from diacetate (CAHP-170) to triacetate (Sureflux UX family) to polysulfone (Fx80); newer membranes show increased clearance | Essential for understanding technological advancements in membrane designs and their effect on drug elimination.
Plasma protein binding comparison | Vancomycin: ~51.8%; Doripenem matches human plasma data well | Demonstrates predictability and relevance of bovine blood as a proxy for human physiology, aiding dose determination.
Elimination influence factors | Flow rates, membrane material, pore size, molecular size, protein binding | Addresses the complexities in predicting drug clearance, emphasizing the necessity of tailored dose recommendations.
Drug-specific membrane interactions | Teicoplanin exhibits strong binding, low clearance despite high flux settings | Shows limitations of relying purely on molecular weight or binding affinity to predict elimination behavior during dialysis.
Correlation between in vitro and historical in vivo clearance | High correlation across all tested drugs | Validates the in vitro model as a tool for predicting pharmacokinetics for clinical applications.
Limitations of the model | No post-hemodialytic rebound effects; rapid plasma drug concentration decline compared to in vivo | Highlights drawbacks important for interpreting and contextualizing in vitro clearance data for real-world situations.
Funding sources | Austrian national bank science fund (ÖNB 15187); Paul-Ehrlich-Gesellschaft’s research funding | Supports the project's validity by showing non-commercial funding sources, ensuring unbiased methodology development.
Patient Sample Size | Not applicable; model used bovine blood not human patients | Study design relies on in vitro simulation rather than patient population, avoiding patient exposure to experimental treatments.
Blood/Tissue Sampling Time Points | Samples were collected at 1, 5, 10, 15, 20, 30, 40, and 60 minutes after system equilibration | Sampling time points were selected to capture rapid drug concentration decline due to reduced blood volume in the in vitro model, enabling accurate clearance measurement.
Patient Population | Not applicable; bovine blood used for experimentation | Bovine blood model mimics human plasma characteristics to simulate drug clearance during intermittent renal replacement therapy without involving human patients.
Reservoir Blood Volume | 2 liters of heparinized bovine blood at 37°C | Critical to simulate in vitro conditions similar to human blood volume encountered during hemodialysis.
Blood Flow Rates | Typically set at 300 ml/min; modified for specific drugs or validation comparisons | Reflects clinical parameters used in hemodialysis to standardize experimental conditions for varied flow rates and drug interactions.
Dialysate Flow Rates | Standardized at 500 ml/min; adjusted for specific tests where required | Important to simulate clinical dialysis scenarios and ensure comparability between in vitro and clinical trial data.
Target Drug Concentrations | Gentamicin 10 mg/L, Vancomycin 20 mg/L, Teicoplanin 25 mg/L, Doripenem 40 mg/L | Mimics plasma levels previously described in clinical trials for accurate simulation and external validation of the model.
Animals Used | Bovine blood obtained from veterinary clinic after ethical approvals | Reflects ethical constraints in model design; bovine blood chosen to mimic hematocrit values similar to human patients undergoing hemodialysis.
Temperature and Anticoagulation | Blood maintained at 37°C; 2,000 IU low molecular weight heparin added during blood draw and experiment | Ensures compatibility with physiological and pharmacokinetic conditions, avoiding clot formation during extended dialysis simulation.
In vitro pharmacokinetic model | Bovine blood-based circulation system | This model simulated intermittent renal replacement therapy specifically, giving insight into drug-membrane interactions and clearance that are pivotal for dosing recommendations.
Compartmental approach | Drug concentrations in plasma measured pre-filter and post-filter | The compartmental calculations for drug clearance can be indirectly inferred due to the segregation of evaluated blood flow parameters into hematocrit corrected/non-corrected readings.
Dialysis flow dynamics model | Utilized Fresenius 4008H dialysis machine | This system simulates realistic drug clearance during hemodialysis, allowing validation and prediction relative to clinical data.
IVIVE-related methods | Validation performed comparing bovine plasma protein binding to human plasma binding | Enables bridging between the in vitro model's predictions and in vivo applicability.
Empirical scaling factor | Vd ratio-based alignment to in vivo PK curves | Demonstrates a method akin to IVIVE (in vitro to in vivo extrapolation) using drug distribution volume correlation for temporal alignment.
Fixed conditions model | Blood flow rate and dialysate flow rate standardized at 300 ml/min and 500 ml/min | Establishes consistent experimental conditions to both control variability and replicate dialysis settings aligned to clinical operations.
--- | --- | ---
Targeted plasma levels for drugs | Gentamicin 10 mg/L, Vancomycin 20 mg/L, Teicoplanin 25 mg/L, Doripenem 40 mg/L | These values define the concentrations of drugs added to the bovine blood reservoir to mimic in vivo plasma levels.
Blood flow rate (BFR) | 300 mL/min | Standard flow rate for hemodialysis experiments ensures comparability.
Dialysate flow rate (DFR) | 500 mL/min | A standardized dialysate flow rate supports consistent clearance analysis across different drugs.
Temperature during experiment | 37°C | Relevant for maintaining physiological conditions of the bovine blood, ensuring accuracy in clearance simulation.
Hemofilter membrane surface area (examples) | CAHP 170: 1.7 m2, SF17UX: 1.7 m2, SF21UX: 2.1 m2, F80s: 1.8 m2 | Membrane surface area affects drug clearance, and these parameters are estimated from prior clinical and technical knowledge of membrane function.
Standard corrected clearance formula | (1-Hc)×BFR×Cpre−(Cpost(BFR−UFR))/(BFR×Cpost) | This formula accounts for hematocrit and other parameters to estimate drug clearance rates.
Hematocrit value assumption | Patient-similar values (~0.35-0.45) | Used for clearance correction to replicate typical patient conditions during hemodialysis.
Historical data comparison | Clearance of gentamicin, vancomycin, teicoplanin, doripenem validated with prior clinical trials (e.g., Amin et al. 1999, Lanese et al. 1989) | Clearance values estimated in vitro were cross-checked with previously published clinical pharmacokinetic trials.
Protein binding estimation | Vancomycin <54.8%, Doripenem <10%, derived from ultrafiltration measures and prior human plasma binding studies | Protein binding values affect drug clearance and were validated against historical data in human plasma.
Time multiplication factor | 11.36 | Estimated as the ratio between in vitro and in vivo volume of distribution (Vd) to align PK curves.
Hemodialysis machine used | Fresenius 4008H | Provides controlled in vitro simulation conditions for drug clearance analysis.
Validated in vitro membranes | Baxter CAHP 170, Nipro Sureflux 17UX, Fresenius F80s, FX80 | Ensures correspondence with clinical membrane types commonly used in dialysis for external validation.
Protein binding experimental method | Centrifugation and HPLC for doripenem; Millipore ultrafiltration for vancomycin | Provides accurate measurements for plasma protein interaction, a key parameter for clearance prediction.
Volume of bovine blood used | 2 L | Selected to mimic patient-level volumes while balancing practical constraints of in vitro models.
Elimination influencing factors | Molecular weight, membrane adhesion, flow rates (BFR, DFR), protein binding | Parameters critical to simulate realistic drug-membrane interactions for clearance evaluation.
Flow rate variability | Blood flow: 400 mL/min; Dialysate flow: 800 mL/min (gentamicin clearance testing with Fresenius Fx80) | Shows how variations in flow rates impact drug clearance outcomes.
Drug-specific elimination trends | Gentamicin clearance: Minimum 90.12 mL/min (CAHP-170) to 187.90 mL/min (Fx80); Doripenem: 141.25 mL/min (Fx80) | Membrane material and flow rate variability underline clearance estimations.
Simulation adjustment for validation | Matching the dialysis conditions from historic trials (membrane and flow rate replication) | Ensures external validation of model with historical datasets.
Membrane material and surface | Different hemofilter models showed variable clearance rates for drugs. For example, gentamicin clearance ranged from 90.12 ml/min with Baxter CAHP-170 diacetate to 187.90 ml/min with Fresenius FX80 polysulfone. | Demonstrates the influence of membrane composition and design on drug elimination during hemodialysis, which directly impacts pharmacokinetics.
Blood flow rate | Increased blood flow rates resulted in higher drug clearance rates. For example, higher blood flow rate correlates with increased clearance for gentamicin across different membranes. | Indicates that blood flow rate is a controllable parameter that greatly affects pharmacokinetics during hemodialysis.
Dialysate flow rate | Dialysate flow rate significantly impacts clearance values; gentamicin showed varied clearance at dialysate flow rates of 500 ml/min and 800 ml/min depending on the membrane type. | Highlights that dialysate flow rate is a critical factor in determining drug clearance during hemodialysis.
Drug-membrane interaction | Drugs such as teicoplanin exhibit specific behaviors based on material-binding affinities, which is linked to molecular size and protein binding. | Illustrates the role of membrane material adherence and molecular size in altering pharmacokinetics outcomes.
Plasma protein binding | Vancomycin (47.1%) and doripenem (7.3%) plasma protein binding showed consistency between bovine and human studies. Protein binding influences drug interaction with the membrane and overall clearance. | Plasma protein binding affects drug availability for filtration and hence pharmacokinetic profiles during renal replacement therapy.
Molecular size of drug | Large molecular weight drugs like teicoplanin showed lower clearance rates compared to smaller molecules like doripenem. | Larger molecules are less likely to diffuse through dialysis membranes, showcasing the inverse proportionality of molecular size to elimination rates in pharmacokinetics.
Hemodialysis conditions | Standard flow rates (BFR 300 ml/min, DFR 500 ml/min) and absence of ultrafiltration were standardized for drug clearance experiments. | Controlled experimental conditions directly impact pharmacokinetic data and aim for repeatable results influencing real-world dosing guidelines.
Animal blood model similarity | Bovine blood demonstrated comparable hematocrit and plasma protein binding values to human blood, enabling reliable predictions of human drug pharmacokinetics. | Validates the model for studying pharmacokinetic variables without human trials, relevant for ethical concerns and cost-effectiveness in drug testing.
Membrane evolution | Gentamicin clearance was higher with modern membranes compared to older designs (e.g., F80s vs FX80 polysulfone). | Evolution in membrane technology directly affects pharmacokinetic parameters by enhancing drug clearance efficiency.
Drug concentration in hematocrit | Hematocrit corrections were required for accurate clearance calculations, especially for drugs with rapid removal rates. | Supports precision in pharmacokinetic modeling by adjusting for red blood cell volume in calculations.