## Freetext Data
Data | Value | Relevance
B1+ field mapping method | Mapped by hand using a network analyzer with an S-parameter test set | This is relevant because understanding how the B1+ field is measured is key to evaluating the accuracy and reproducibility of hyperpolarized 13C MRI studies.
Homogeneous region dimensions | Approximately 12.9 × 11.5 × 13.4 cm (X, Y, Z directions) | Relevant to assessing patient positioning and ensuring imaging biomarkers are consistent within specific volumes of the 13C clamshell coil.
Impact of excitation angles | Higher excitation angles yielded higher signal and SNR but reduced the uniform measurement region size | Relevant to optimizing signal strength while balancing spatial homogeneity for metabolic imaging applications.
Optimal excitation angles | θP and θL set to 10° for largest homogenous region | Relevant for selecting excitation angles that ensure reproducibility and calibrate imaging biomarkers effectively across the target anatomy.
Error estimation approach | Percent difference calculated relative to values at isocenter | Relevant because it quantifies imaging biomarker variability caused by B1+ field inhomogeneity, aiding in gauging consistency across the imaging volume.
Imaging biomarkers assessed | Normalized lactate ratio (nLac) and apparent rate constant kPL | Relevant because these biomarkers are critical metrics for evaluating tumor metabolism and their sensitivity to coil inhomogeneities is central to the study's goals.
N/A | N/A | The paper does not provide specific details about the patient sample size, blood/tissue sampling time points, or patient population involved in the pharmacokinetic (PK) study. Instead, it focuses on simulations based on pharmacokinetic models to study the impact of B1+ field inhomogeneity on metabolic imaging biomarkers.
Pharmacokinetic Model | Simulated signal evolution as a function of position-dependent local excitation angles | This describes the model used in the study for assessing the impact of B1+ field inhomogeneities on metabolic imaging biomarkers, making it directly relevant to the paper's methodology.
B1+ field range | ±10% of the value at isocenter | Defines the inhomogeneity threshold used to characterize the B1+ field and its effects on measured biomarkers.
B1+ field mapping tool | Network analyzer with an S-parameter test set | Used to derive the B1+ field maps for the 13C clamshell coil.
Region of uniform B1+ field | 12.9 × 11.5 × 13.4 cm (X, Y, Z) centered in the coil | Defines the spatial zone within the coil where B1+ field deviations are minimal and biomarkers are reliably measured.
θP and θL value for largest uniform region | 10° | Represents the excitation angles that maximize the homogeneity of the B1+ field to maintain biomarker fidelity.
Higher excitation angle tradeoff | Higher signal and SNR but reduced uniformity | Demonstrates the relationship between signal intensity and the homogeneous region's size in imaging applications.
Pyruvate excitation angle (θP) | Nominal excitation angles assumed calibrated at isocenter | Indicates the parameter governing signal excitation and its dependence on position relative to the isocenter.
Lactate excitation angle (θL) | Nominal excitation angles assumed calibrated at isocenter | Reflects its role in position-dependent signal quantification and imaging biomarker variability.
B1+ Field Inhomogeneity | Variations in ±10% of value at isocenter | B1+ field variations directly influence excitation angles, impacting signal evolution and subsequent pharmacokinetic modeling.
Excitation Angle Calibration | Nominal excitation angles assumed to be accurately calibrated at isocenter | Accurate calibration of excitation angles is critical for reliable pharmacokinetic modeling and minimization of error in metabolic biomarker quantification.
Lactate Ratio (nLac) Sensitivity | Errors estimated via normalized lactate ratio (nLac) | Lactate ratios are a key semi-quantitative biomarker influenced by B1+ field; errors in this metric affect pharmacokinetic assessment.
Apparent Rate Constant (kPL) Sensitivity | Errors calculated using kPL conversion rate | The conversion constant reflects the pyruvate-lactate metabolic pathway; B1+ field uniformity ensures reliable computation of this pharmacokinetic parameter.
Excitation Angle Tradeoff | Larger region uniform at lower angles (θP and θL = 10°) | Lower excitation angles favor larger uniform regions, reducing pharmacokinetic variability due to spatial inhomogeneity.
Signal-to-Noise Ratio (SNR) | Higher excitation angles yield increased SNR | Although improving SNR enhances signal interpretation, it narrows the region for accurate pharmacokinetic modeling, influenced by B1+ profile.
Position Dependency | Errors vary with position in the coil | Spatial position impacts local excitation angles and, consequently, pharmacokinetic signal analysis, requiring precise patient positioning.