## Table 1
0 | 1
Hardware name | HeartView Data Acquisition System
Subject area | • Biomedical Engineering • Medical sensor
Hardware type | • Clinical research tool • Flow sensor and pressure data acquisition system • Clinical hemodynamic monitoring research applicable • Edwards Lifesciences TruWave pressure sensor compatible
Open source license | Creative Commons Attribution-ShareAlike 4.0
Cost of hardware | ~US$90
Source file repository | https://doi.org/10.17605/osf.io/24zuh

## Table 2
Design filename | File type | Open source license | File location
Production Files
DAQ_Housing_Base.stl | 3D-printable | CC BY-SA 4.0 | DAQ/
DAQ_Housing_Lid.stl | 3D-printable | CC BY-SA 4.0 | DAQ/
DAQ_PCB.zip | Contains Gerber and drill files | CC BY-SA 4.0 | DAQ/
Arduino_read.ino | Arduino project | CC BY-SA 4.0 | DAQ/
Loadcell_Housing_Base.stl | 3D-printable | CC BY-SA 4.0 | Hanging Scales/
Loadcell_Housing_Lid.stl | 3D-printable | CC BY-SA 4.0 | Hanging Scales/
GUI.zip | Contains executable software application | CC BY-SA 4.0 | User Interface/
Cable_Cover_A.stl | 3D-printable | CC BY-SA 4.0 | Connectors/
Cable_Cover_B.stl | 3D-printable | CC BY-SA 4.0 | Connectors/
Source Files
DAQ_Housing_Base.FCStd | FreeCAD project | CC BY-SA 4.0 | DAQ/
DAQ_Housing_Lid.FCStd | FreeCAD project | CC BY-SA 4.0 | DAQ/
DAQ_PCB.pro | KiCAD Project | CC BY-SA 4.0 | DAQ/
Loadcell_Housing_Base.FCStd | FreeCAD project | CC BY-SA 4.0 | Hanging Scales/
Loadcell_Housing_Lid.FCStd | FreeCAD project | CC BY-SA 4.0 | Hanging Scales/
GUI_main.py | Python File | CC BY-SA 4.0 | User Interface/
GUI_setup_exe.py | Python File | CC BY-SA 4.0 | User Interface/
Cable_Cover_A.FCStd | FreeCAD project | CC BY-SA 4.0 | Connectors/
Cable_Cover_B.FCStd | FreeCAD project | CC BY-SA 4.0 | Connectors/

## Table 3
Component | Qty | Total Cost (USD)
Key Components
Edwards Connector IBP Cable | 1 | $29.99
Arduino Nano ATMega328 | 1 | $20.70
INA828 | 2 | $11.16
Load Cell 3 kg | 1 | $1.40
System totals
DAQ Unit | 1 | $42.12
Hanging Scales | 1 | $6.77
Connectors | 1 | $38.71
Total Device |  | $87.61

## Freetext Data
Data | Value | Relevance
Set-up Arduino | Arduino code "Arduino_read.ino" is loaded onto the Arduino Nano | Relevant as it describes the preparation procedure needed to operate the data acquisition system.
Set-up system components | Hang fluid bag from hanging scales; connect scales to DAQ via RJ-45 connector; secure connection with cable cover; attach Edwards TruWave pressure transducer to DAQ via hybrid pressure transducer-RJ-45 connector; connect DAQ to computer via USB | Relevant as it outlines the necessary setup process for utilizing the device with its components.
Calibration for load cell | Load cell calibration is done using known weights hung from the sensor | Relevant as it ensures the accuracy of fluid infusion rate measurements by accounting for varying attachment weights.
Calibration for pressure sensor | Pressure transducer calibration is performed using a water column | Relevant since accurate pressure measurements are critical to the device’s performance in clinical research applications.
Pressure recording range | 0–200 mmHg | Relevant to describe the operational parameters for hemodynamic monitoring compatibility.
Fluid load recording range | 0–2 kg | Relevant for understanding the maximum fluid infusion that can be monitored.
Resolution for pressure | 0.4 mmHg | Relevant because it indicates the precision of pressure readings, sufficient for capturing clinically meaningful changes.
Resolution for load cell | 6 g | Relevant as it describes the precision of fluid measurements, acceptable within typical clinical fluid delivery volumes.
Data sampling rate | 500 Hz | Relevant as it relates to the capability of the device to record high-resolution physiological waveforms.
Validation of pressure signal | Bland-Altman bias and limits of agreement were −1 [-10, 8] mmHg | Relevant as it evaluates the device's performance by comparing its pressure signal data with a clinical monitor.
Validation of fluid rate | Bland–Altman bias and limits of agreement were −0.08 [-0.89, 0.73] ml/s | Relevant to assess the fluid infusion rate measurement accuracy of the device.
Arduino Controller | Arduino Nano (ATmega328p) | Relevant as it is the central processing unit of the device, enabling data acquisition and output.
Software interface | Python application, “HeartView_application.exe,” used for viewing real-time outputs | Relevant as it provides the tool for controlling and analyzing data acquired by the system.
Compatibility | Edwards Lifesciences TruWave pressure transducers | Relevant as the device is specifically designed for compatibility with widely used clinical pressure transducers.
Medical relevance | Hemodynamic monitoring in ICU, especially for research in cardiovascular response to fluids and inotropes | Relevant as it indicates the clinical applications the device was designed to support.
Electrical hardware | PCB with Arduino Nano, two INA828 amplifiers; powered by USB | Relevant to understanding the electronic design and functionality of the system.
Cost | ~US$90 for each unit | Relevant to highlight the affordability of the system compared to commercial alternatives.
Data output | .csv file format | Relevant as it demonstrates compatibility with standard data analysis tools used in research.
Patient sample size | Not specified | The paper does not discuss the sample size for a pharmacokinetic (PK) study; instead, it describes validation using simulations and mechanical testing, which may imply no human patients were involved.
Blood sampling time points | Not applicable | Sampling time points specific to PK studies are not mentioned as the focus of the paper is on developing and validating a data acquisition system.
Tissue sampling time points | Not applicable | The study does not involve tissue sampling or PK study design but focuses on acquiring pressure waveforms and fluid infusion data.
Patient population | Not conducted; validation in a controlled hospital setting | The study mentions validating the device in a hospital using clinical sensors, but no specific patient population or direct patient data collection is described, making it not relevant for PK study design.
Device validation in hospital | Controlled environment with sensor coupling simulation | Relevant for understanding the context of data acquisition and its potential application, though unrelated to direct PK study involving patients or sampling schedules.
Model Type | Physiological | The described system builds upon computational models designed to utilize ICU monitoring data for improving patient outcomes, aligning with physiologically based modeling approaches aimed at characterizing hemodynamic responses.
System Classification | Model-based System | The study emphasizes using computational models for clinical care improvements related to the cardiovascular system, indicating reliance on physiological modeling techniques suitable for research contexts.
Operational range for pressure acquisition | 0–200 mmHg | Indicates the device's capability to measure arterial pressure within a clinically relevant range.
Operational range for load cell | 0–2 kg (corresponding to 0–2 L saline) | Relevant as it defines the maximum measurable weight for fluid infusion monitoring, important for practical use in ICU settings.
Resolution for pressure | 0.4 mmHg | Relevant for capturing clinically meaningful pressure changes of ≥5 mmHg, ensuring precision in research applications.
Resolution for load cell | 6 g | Relevant for monitoring fluid delivery with a typical resolution of 100–500 ml, critical for fluid management in hemodynamic research.
Sampling rate | 500 Hz | Relevant because it determines the temporal resolution of waveform data, enabling detailed real-time analysis.
Coefficient of determination for pressure calibration | R² > 0.99 | Indicates high linearity and low noise of the pressure sensor, validating its reliability for research.
Coefficient of determination for load cell calibration | R² > 0.99 | Validates the load cell's precision and linearity for accurate infusion rate measurements.
Arduino ADC-input voltage increment | 0.0049 V | Determines the resolution of the signal, which is a limiting factor defined by the hardware.
Calibration method for pressure | Water column | Specifies an accurate and convenient procedure to establish pressure measurement baselines.
Calibration method for load cell | Known weights hung from the sensor | Provides a reliable method to calibrate fluid weight measurements, essential for accurate fluid infusion rate calculations.
Pressure sensor validation | Compared to Philips Intellivue MX800 Monitor | Provides assurance of the accuracy of arterial pressure measurements in clinical settings.
Bland–Altman mean bias for pressure comparison | −1 mmHg | Validates that the device closely agrees with recognized clinical monitors, ensuring relevance in clinical research.
Limits of agreement for pressure comparison | [-10, 8] mmHg | Relevant as it quantifies the measurement accuracy and error range for physiological monitoring.
Fluid infusion rate validation | Measured via weighed infusion bags draining at different rates | Ensures accuracy for tracking fluid delivery rates, critical for intervention studies.
Bland–Altman mean bias for fluid rate comparison | −0.08 ml/s | Establishes that the device accurately records infusion rates, matching manual methods without significant error.
Limits of agreement for infusion rates | [-0.89, 0.73] ml/s | Provides confidence in the precision of fluid infusion measurements during experiments.
Resolution limiting factor | 10-bit ADC of Arduino Nano | Indicates the hardware constraint affecting overall measurement precision.
Calibration method | Calibration of the fluid infusion measurement involves testing the load cell with known weights, and pressure transducers using a water column | Calibration impacts the accuracy of monitoring fluid infusion, potentially influencing pharmacokinetics by affecting the measurement of drug infusion rates.
Sampling rate | Signals were recorded with a sampling rate of 500 Hz | Accurate sampling impacts the monitoring of physiological responses, which could relate to drug pharmacokinetics.
Resolution | The resolution is 0.4 mmHg for pressure and 6 g for the load cell | Precision in physiological measurement may significantly influence pharmacokinetic modeling by accurately reflecting changes induced by drug administration.
Validation method | The device was validated using fluid infusion bags and pressure transducers in hospital settings | Validation ensures reliable data for studying pharmacokinetics, particularly in fluid and drug delivery systems.
Fluid infusion measurement | Fluid infusion rate measurement was validated using drainage tests and a mathematical filtering process | Precise fluid measurement influences drug dispersion and, consequently, pharmacokinetics.
Pressure signal comparison | Pressure acquisition was compared with the Philips Intellivue MX800; the range of error was minimal | Accurate pressure measures may impact pharmacokinetics by ensuring appropriate monitoring of hemodynamic conditions affected by drugs.