## Table 1
Unnamed: 0_level_0_Concentration(µg/mL) | Intra-Day (n = 5)_Mean ± SD(µg/mL) | Intra-Day (n = 5)_CV (%) | Intra-Day (n = 5)_Bias (%) | Inter-Day (n = 5)_Mean ± SD(µg/mL) | Inter-Day (n = 5)_CV (%) | Inter-Day (n = 5)_Bias (%)
0.1 | 0.11 ± 0.01 | 1.23 | 12.50 | 0.10 ± 0.01 | 5.82 | 2.96
0.5 | 0.49 ± 0.02 | 4.81 | −2.10 | 0.56 ± 0.01 | 0.88 | 12.89
1 | 0.89 ± 0.04 | 4.02 | −10.99 | 1.09 ± 0.03 | 3.04 | −9.05
2 | 1.84 ± 0.11 | 6.04 | −7.96 | 2.09 ± 0.06 | 2.99 | 4.83

## Table 2
Unnamed: 0_level_0_Concentration (µg/mL) | Recovery (%)_Mean ± SD | Recovery (%)_CV (%)
0.5 | 89.66 ± 6.93 | 7.73
1 | 95.02 ± 6.77 | 7.12
2 | 96.92 ± 4.65 | 4.8

## Table 3
Condition_Unnamed: 0_level_1 | Condition_Unnamed: 1_level_1 | QC Low 0.5 (µg/mL)_Mean ± SD (µg/mL) | QC Low 0.5 (µg/mL)_CV (%) | QC Low 0.5 (µg/mL)_Bias (%) | QC High 2.0 (µg/mL)_Mean ± SD (µg/mL) | QC High 2.0 (µg/mL)_CV (%) | QC High 2.0 (µg/mL)_Bias (%)
Short-term stability | 0 h | 0.47 ± 0.03 | 5.86 | −4.91 | 2.05 ± 0.05 | 2.52 | 2.83
Short-term stability | 24 h (RT) | 0.44 ± 0.02 | 3.9 | −11.00 | 1.83 ± 0.10 | 5.53 | −8.37
Short-term stability | 24 h (4 °C) | 0.44 ± 0.01 | 2.99 | −10.54 | 1.86 ± 0.04 | 2.44 | −7.17
Long-term stability (−40 °C) | 15 days (−40 °C) | 0.45 ± 0.03 | 6.26 | −8.73 | 1.83 ± 0.02 | 0.94 | −8.51
Long-term stability (−40 °C) | 30 days (−40 °C) | 0.48 ± 0.03 | 6.61 | −2.58 | 1.80 ± 0.05 | 2.83 | −9.77
5 Freeze–thaw cycles (−40 °C) | 5 days of freeze–thaw, (−40 °C) | 0.48 ± 0.05 | 10.79 | −3.31 | 1.78 ± 0.07 | 3.87 | −10.76

## Table 4
Condition_Unnamed: 0_level_1 | Condition_Unnamed: 1_level_1 | Stock Solution Low 2.5 (µg/mL)_Mean ± SD (µg/mL) | Stock Solution Low 2.5 (µg/mL)_CV (%) | Stock Solution Low 2.5 (µg/mL)_Bias (%) | Stock Solution High 25 (µg/mL)_Mean ± SD (µg/mL) | Stock Solution High 25 (µg/mL)_CV (%) | Stock Solution High 25 (µg/mL)_Bias (%)
Short-term stability | 0 h | 2.45 ± 0.16 | 6.59 | −2.13 | 25.20 ± 1.20 | 4.76 | 0.81
Short-term stability | 7 days (−40 °C) | 2.59 ± 0.01 | 0.19 | 3.89 | 24.18 ± 0.16 | 0.63 | −3.28
Long-term stability (−40 °C) | 20 days (−40 °C) | 2.53 ± 0.14 | 5.56 | 1.22 | 24.73 ± 1.44 | 4.63 | −1.07
Long-term stability (−40 °C) | 60 days (−40 °C) | 2.41 ± 0.23 | 9.41 | −3.67 | 24.54 ± 0.32 | 1.29 | −1.82

## Freetext Data
Data | Value | Relevance
Tylvalosin (TV) | Third-generation macrolide antibiotic | Tylvalosin's classification as a macrolide antibiotic relevant for veterinary treatment of bacterial infections in swine and poultry.
Usage of TV | Treating respiratory and enteric bacterial infections | Relevant to explaining why the drug is used in specific veterinary species.
FDA guidelines | Used for validation | Relevant since the study adhered to regulatory standards ensuring method reliability.
HPLC methodology | Quantifies TV in pig plasma | Relevant as the paper focuses on developing an analytical approach using HPLC for drug measurement.
UV detection | 254 nm chosen as final wavelength | Relevant for the technical setup of the quantification method.
Mobile phase | 0.3% formic acid in water: acetonitrile (78:22) | Relevant for defining the conditions under which TV is quantified.
Sample preparation | Protein precipitation (PP) used | Relevant as this method simplifies the analytical process and aligns with eco-friendly principles.
LOQ | 0.1 µg/mL | Relevant to the method's sensitivity in detecting small concentrations of TV.
LOD | 0.05 µg/mL | Relevant for determining the minimum detectable concentration of TV.
Calibration curve range | 0.1–5 µg/mL | Relevant as this range supports the method’s applicability for pharmacokinetic studies.
Precision | Intra-day ≤ 6.1%, Inter-day ≤ 6.1% | Relevant for assessing the reproducibility of the HPLC method.
Accuracy | Intra-day (-7.96% to 12.50%), Inter-day (-9.05% to 12.89%) | Relevant for determining the reliability of measurements in this validated method.
Recovery values | 89.66–96.92% | Relevant indicating method efficiency in isolating TV from pig plasma.
Matrix effects | Low; no interfering peaks | Relevant as it supports the method’s specificity for TV in pig plasma.
Stability | Stable for ≥30 days at -40°C | Relevant to ensure plasma samples maintain integrity over time for analysis.
Test species | Swine | Relevant to the species-specific focus of the method for pharmacokinetic applications.
Dilution accuracy | ≤ ±15% of nominal concentration | Relevant to the reliability of method under dilution conditions.
Post-administration TV detection | Detected at 2, 4, and 24 h post-dose | Relevant as it demonstrates the method's application in real pharmacokinetic studies.
System | HPLC with UV detection | Relevant as the main technological framework used in the research.
Principles | Follows green chemistry | Relevant due to the method's cost and environmental sustainability emphasis.
Plasma Samples | 2 pigs used for the study | Relevant to describe the sample size in the pharmacokinetics (PK) study.
Animal Weight | Pigs weighed approximately 20 kg | Relevant to define the subject characteristics for the PK study.
TV Administration | Single oral dose of 5 mg/kg | Relevant for the PK study in describing administered drug dosing.
Sampling Time Points | Blood collected at 2, 4, and 24 hours | Relevant to identify the sampling time points in the PK study.
Administration Method | TV dissolved in water and delivered as a single bolus by oral gavage | Relevant for understanding how the antibiotic was administered during the PK study.
Blood Sampling Location | Blood collected from the jugular veins | Relevant for understanding the methodology of blood sampling in the PK study.
Sample Preservation | Plasma stored at −40 °C until analysis | Relevant to assure stability and integrity of the biological samples for PK analysis.
Ethical Approval | Approved by the Bioethical Committee of the University of Murcia (CEEA 558/2019) | Relevant to ensure ethical considerations in PK study design.
N/A | The manuscript does not explicitly mention the use of a specific pharmacokinetic model such as 1-compartment, 2-compartment, physiologically based model, or IVIVE. | The relevance lies in understanding the methodology; since the study focused on the analytical method development and not on modeling, no explicit pharmacokinetic model is addressed.
Calibration Curve Range | 0.1 to 5 μg/mL | Relevant to the validation of the analytical method and its applicability to accurately determine tylvalosin plasma concentrations.
LOD (Limit of Detection) | 0.05 μg/mL | Relevant for highlighting the sensitivity of the method in detecting tylvalosin accurately.
LOQ (Limit of Quantification) | 0.1 μg/mL | Relevant to indicate the minimum concentration that can be accurately quantified using the method.
Mobile Phase Composition | 78:22 of 0.3% formic acid in water:acetonitrile | Relevant for replicating conditions under the validated chromatographic method.
Flow Rate | 1.0 mL/min | Relevant to reproducing the specific conditions of the chromatographic method for tylvalosin quantification.
Column Temperature | 22°C | Relevant for maintaining consistent chromatographic conditions for method reproducibility.
Injection Volume | 50 µL | Relevant as a parameter in the method for tylvalosin quantification at validated concentrations.
Wavelength of UV Detection | 254 nm | Relevant to ensure the detection settings are appropriate for tylvalosin's optimum absorption spectrum.
Retention Time for TV | 14.1 min | Relevant for identifying tylvalosin's elution in chromatograms and ensuring accuracy in detection.
Retention Time for IS (Enrofloxacin) | 5.9 min | Relevant for validating the internal standard's role in the quantification process.
Recovery Range | 89.66% to 96.92% | Relevant for confirming the method's efficiency in extracting tylvalosin from pig plasma.
Short-term Stability (QC at 4°C/Room Temperature) | CV < 5.9%; Bias ≤ 11.0% | Relevant to validate the stability parameters for tylvalosin under controlled conditions during short-term handling.
Long-term Stability (QC at −40°C for 15 Days) | Bias at Low QC: −8.73%, CV 6.26%; Bias at High QC: −8.51%, CV 0.94% | Relevant for understanding tylvalosin's integrity under extended storage conditions.
Long-term Stability (QC at −40°C for 30 Days) | Bias at Low QC: −2.58%, CV 6.61%; Bias at High QC: −9.77%, CV 2.83% | Relevant to further validate the method for long-term storage of tylvalosin.
Freeze–Thaw Stability (5 Cycles, −40°C) | Low QC: Decrease 3.31%, CV 10.79%; High QC: Decrease 10.76%, CV 3.87% | Relevant to ensure the reliability of tylvalosin's quantification after multiple freeze–thaw cycles.
Linear Regression Equation | y = 0.9993x + 0.0021 | Relevant to defining the linear relationship between concentration and response in the calibration curve for tylvalosin measurement.
Correlation Coefficient (r²) | ≥ 0.9998 | Relevant to measure the reliability and linearity of the calibration curve.
Chromatogram Run Time | 15 min | Relevant for understanding method efficiency and analysis speed.
Recovery from Solvent | Acetonitrile with 0.1% Formic Acid | Relevant for ensuring optimal extraction solvent for tylvalosin from pig plasma samples.
Intra-day Precision (CV) | < 6.1% | Relevant for assessing the consistency of measurements within the same day.
Inter-day Precision (CV) | < 6.1% | Relevant for assessing the reproducibility of measurements across multiple days.
Intra-day Accuracy | −7.96% to 12.50% | Relevant to evaluating tylvalosin measurement biases within the same day.
Inter-day Accuracy | −9.05% to 12.89% | Relevant to evaluating tylvalosin measurement biases across different days.
Differences in blood composition and functional properties | Higher levels of hemoglobin and plasma protein content in porcine blood compared to avian blood; porcine blood had a less nutritive amino acid composition and differences in heme iron content as well as gelation properties | Relevant because differences in blood composition can influence the pharmacokinetics by affecting the extraction ratio and the biological matrix specificity.
Plasma protein content | Higher in porcine blood compared to avian species | Relevant for pharmacokinetics as plasma protein levels affect drug binding and distribution in the body.
Extraction method for plasma samples | Protein precipitation (PP) preferred for pig plasma due to its simplicity and eco-friendliness compared to liquid-liquid extraction (LLE) | Relevant because sample preparation methods can affect the reliability and reproducibility of pharmacokinetic analysis.
Absorption spectrum wavelength | Different absorption areas identified for tylvalosin at 254 nm using pig plasma vs. 236 nm previously found | Relevant as wavelength selection directly affects the accuracy and precision in determining tylvalosin concentration, which in turn impacts pharmacokinetic evaluations.
Impact of storage conditions | Stability tests revealed tylvalosin concentration in plasma decreases with prolonged storage at −40°C and repeat freeze–thaw cycles | Relevant because stability directly influences the reliability of pharmacokinetic studies if samples degrade over time or during handling.