## Table 1
Component | Retention time (min) | Positive ion mode MRM transition trace (m/z) | Cone voltage (V) | Collision energy (eV)
Pemetrexed | 2.71 | 428.08 > 281.06 | 14 | 18
Pemetrexed Q | 2.71 | 428.08 > 163.00 | 14 | 36
[13C5]‐Pemetrexed | 2.71 | 433.10 > 281.01 | 14 | 20

## Table 2
Validation sample | Nominal concentration (μg/L) | Within‐day accuracy (n = 5) (%) | Within‐day imprecision (RSD%) | Between‐day accuracy (n = 15) (%) | Between‐day imprecision (RSD%)
LLOQ | 0.0252 | 74.2 | 6.04 | 77.4 | 5.39
QC low | 0.0602 | 88.4 | 4.03 | 90.2 | 0.78
Alternative LLOQ | 0.075 | 107 | 6.05 | 103 | 3.38
QC alternative low | 0.12 | 90.6 | 4.57 | 98. 6 | 6.77
QC medium | 1.2 | 108 | 8.79 | 104 | 3.99
QC high | 12 | 102 | 2.86 | 101 | 0
HLOQ | 25.2 | 105 | 5.93 | 102 | 3.23
Average |  | 96.5 |  | 96.5

## Table 3
Matrix | Condition | Time interval | Mean concentration compared to nominal concentration (%)
Stock solutions | –40°C | 24 months | 102 ± 0.3 (n = 3)
Spiked plasma | Room temperature 4°C –40°C Three freeze–thaw cycles | 7 days 7 days 14 months | 99.3 ± 2.4 (n = 6) 94.9 ± 6.3 (n = 6) 108 ± 4.8 (n = 6) 107 ± 5.5 (n = 6)
Spiked whole blood | Room temperature 4°C | 3 days 3 days | 97.9 ± 12.2 a (n = 6) 102 ± 6.7 (n = 6)

## Freetext Data
Data | Value | Relevance
Study Title | Highly sensitive quantification of pemetrexed in human plasma using UPLC‐MS/MS to support microdosing studies | Provides context and relevance of the study focus—assay development for pharmacokinetics of a microdose drug.
Drug Studied | Pemetrexed | Central focus of the study; directly relevant to drug assay development and application in clinical contexts.
Drug Type | Multitargeted antifolate | Relevant to understanding the mechanism of action and therapeutic category of the drug being studied.
Target Conditions | Non‐small‐cell lung cancer, Mesothelioma | Highlights the clinical application and therapeutic relevance of pemetrexed.
Clinical Usage | Monotherapy or combination treatment for lung cancer | Explains real-world use in oncology and the importance of accurate dosing methods.
Microdose Amount | 100 μg | Indicates the novel approach to microdosing for pharmacokinetic profiling with reduced toxicity.
Assay Methodology | Reversed‐phase UPLC‐MS/MS | Core approach developed for highly sensitive drug quantification.
Internal Standard Used | [13C5]-Pemetrexed | Key component ensuring assay precision and accuracy.
Sample Preparation | Solid-phase extraction from plasma | Relevant to technical procedures improving sensitivity and recovery rates in the assay.
Detection Technique | UPLC-MS/MS in positive electrospray ionization (MRM mode) | Explains the analytical process used to detect low concentrations of the drug.
Validation Range | 0.025–25.0 μg/L | Identifies the dynamic range of the assay, demonstrating its suitability for microdose studies.
Accuracy Percentage | 96.5% | Assesses reliability of the assay in measuring pemetrexed concentrations.
Precision Metrics | Within-day and between-day coefficient of variations <8.8% | Demonstrates assay reproducibility, essential for pharmacokinetics.
Extraction Recovery | 59 ± 1% for pemetrexed, 55 ± 5% for IS | Indicates efficiency and consistency in drug recovery during sample preparation.
Stability of Processed Samples | 2 days at 10°C | Relevant for practical considerations and ensuring the assay produces valid data over time.
Study Application | Successful quantification of pharmacokinetic curve in clinical microdosing study | Validates the effectiveness and clinical utility of the developed assay for low-dose pharmacokinetics.
Type of Plasma Used | EDTA-treated plasma | Relevant to assay validation and ensuring consistency across clinical sampling.
Carryover Specification | <20% for pemetrexed, <5% for IS | Assures methodological reliability and reduces potential for cross-sample contamination.
Validated Calibration Model | Quadratic regression, weighting factor 1/x2 | Details specific statistical methods for calibration and assay reliability.
Autosampler Temperature | 10°C | Critical factor for maintaining the integrity of processed samples during analysis.
Mobile Phase Composition | 0.1% formic acid in water and acetonitrile | Elucidates chemical optimization for chromatographic separation in UPLC-MS/MS.
Column Specification | BEH C18 column (100 × 2.1 mm i.d.) | Relevant for chromatographic performance and drug sensitivity during separation.
Analyte Stability at Room Temperature | 97.9 ± 12.2% | Indicative of reliable analysis conditions for study replication.
Funding Source | ZonMW, Grant Number 848016010 | Identifies the funding organization and supports transparency in potential biases.
Clinical Study Identifier | NCT03655834 | Provides direct access to the study for further investigation and validation.
Endogenous Interference | Plasma free from interference | Ensures assay specificity and minimizes confounding factors for quantification.
Conflict of Interest Statement | Authors declare no conflicts | Vital for maintaining the integrity and reliability of scientific contributions.
Patient Population | Plasma of six oncological patients was evaluated for potential interferences by endogenous substances. Clinical samples included EDTA-anticoagulated whole blood from oncological patients. | Relevant as it shows the study included plasma samples from cancer patients for PK and assay validation.
Blood Sampling Time Points | Plasma concentration–time curve measured at 0.5–1, 1–2, 3–4, and 6–8 hours after administration of a 100 μg dose | Relevant as it establishes the sampling time points for tracking pemetrexed pharmacokinetics post microdose.
Dose Administered | Single dose of 100 μg pemetrexed administered | Relevant for understanding the dosing used in the study to assess microdosing PK.
Clinical Identifier | Pharmacokinetic curve obtained as part of ClinicalTrials.gov study NCT03655834 | Relevant as it links the PK study design to an ongoing clinical trial for credibility and context.
Patient Sample Preparation | EDTA-anticoagulated whole blood was used, centrifuged for plasma preparation, stored at -40°C | Relevant as it explains how patient samples were processed and stored for PK analysis.
Monte Carlo simulation population pharmacokinetic model | Previously developed population pharmacokinetic model by Latz et al., 2006 | Used to predict expected pemetrexed concentrations associated with a 100-μg dose, supporting validation of assay range and microdosing study feasibility.
Accuracy | 96.5% | Relevant for assay validation to ensure the method's consistency in measuring pemetrexed concentrations.
Within-day precision (maximum coefficient of variation) | <8.8% | Demonstrates the reliability of the assay across multiple replicates within a single day's measurements.
Between-day precision (maximum coefficient of variation) | <8.8% | Indicates the reproducibility of the assay on different days.
Extraction recovery for pemetrexed | 59 ± 1% | Indicates assay efficiency in extracting pemetrexed from plasma, critical to method reliability.
Extraction recovery for internal standard | 55 ± 5% | Ensures accurate quantification of pemetrexed given its reliance on internal standard comparison during analysis.
Tested concentration range | 0.0250–25.0 μg/L | Estimated from Monte Carlo simulation of pharmacokinetics from previous population studies for expected plasma concentrations post-microdosing.
Optimal cone voltage | 2.7 kV | Mass spectrometric parameter essential for sensitivity during quantification.
Optimal collision energy (pemetrexed) | Not specified in a numeric detail | Critical for fragmentation and detection sensitivity in MS/MS.
Optimal collision energy ([13C5]-Pemetrexed IS) | Not specified in numeric detail | Ensures internal standard fragment intensity for calibration consistency.
Optimal source temperature | 150°C | Parameter tuned for stability and sensitivity enhancement during ionization in mass spectrometry.
Optimal desolvation temperature | 350°C | Enhances evaporation and ionization efficiency during the MS analysis.
Desolvation gas flow | 950 L/h | Parameter set for detector stability and maximal sensitivity in MS ionization.
Optimal mobile phase components | 0.1% formic acid in water (A) and 0.1% formic acid in acetonitrile (B) | Optimized for chromatography baseline, retention time, and peak clarity.
Mobile phase starting composition | 85% A and 15% B | Ensures initial conditions for retention of pemetrexed.
Flow rate | 0.3 mL/min | Controls the eluent speed for consistent peak resolution during UPLC analysis.
Validation guideline references | EMA 2011 & FDA 2018 | Used for systematic validation of assay methods.
Column type | BEH C18 column | Chosen for its ruggedness and resolve sensitivity characteristics suitable for pemetrexed analysis.
Retention time | 2.8 min | Critical for chromatographic resolution of pemetrexed and avoiding baseline disturbances.
Internal standard concentration | 50 μg/L (stock solution) | Calibrates and standardizes quantification depending on the nominal levels and dynamic ranges of both analyte and plasma interferences.
Carryover limit for pemetrexed | <20% of LLOQ | Maintains the assay's specificity and separation efficiency during subsequent extractions.
Carryover limit for internal standard | <5% | Ensures the robustness of sensitive minimal signal under subsequent instrument injections.
Processed sample stability parameters | Stable for 2 days at 10°C | Ensures practical clinical application during pharmacokinetic studies, despite processed samples held >72 hours.
Calibration curve regression coefficient | 0.9978 ± 0.0004 | Determined using quadratic least-square fits for dependably patient micro-dose processing context visibility consistency.assertj
Carryover effect during SPE | Substantial carryover effect was observed for pemetrexed during SPE leading to blank samples showing higher concentrations after extracting high-concentration calibrators. | Relevant as it may impact the accuracy of pharmacokinetic measurements if not addressed properly.
Hemolysis effect on plasma samples | Mean concentration deviations of 9.24% for QC low and 4.32% for QC high, and relative standard deviations of 5.25% and 2.00%, respectively. | Relevant because hemolysis impacts plasma sample integrity, which may alter observed pemetrexed levels, affecting pharmacokinetic analysis.
Lipemic plasma effect | Deviations of 5.71% and 4.22% from nominal concentration for 150 and 300 mg/dL intralipid, with relative standard deviations of 5.04% and 3.26%. | Relevant since lipemic plasma alters matrix composition, potentially influencing drug concentration measurements in pharmacokinetic studies.
Autosampler Stability | Processed samples were stable for 2 days in autosampler at 10°C with mean concentration deviations of 102 ± 7.41%. | Relevant as processed sample stability ensures integrity in pharmacokinetic modeling over extended analytical run times.
Pemetrexed chemical degradation | Degradation occurs via oxidation and hydrolysis; stability improved by adding ascorbic acid. | Relevant since degradation influences the reliability of observed drug concentrations in pharmacokinetic assessments.
Matrix factor (MF) variation | Relative standard deviation of the IS-normalized MF was less than 15% across six plasma lots at QC high and QC low concentrations. | Relevant as matrix variability impacts observered drug response and accurate interpretation of pharmacokinetic data.
Batch-to-batch SPE cartridge difference | The difference in mean concentrations for pemetrexed QC samples measured with two cartridge batches was less than 15%. | Relevant to ensure consistent recovery and reproducibility in pharmacokinetic quantifications across cartridge lots.
Plasma type contamination | Defrosted Li-heparin and EDTA–plasma led to obstruction of SPE cartridges, decreasing flow and resulting in lower response for pemetrexed and IS. | Relevant as plasma preparation methods affect extraction efficiency and overall drug quantification in pharmacokinetics.