## Table 1
OXY spiked conc. (ng/ml)_OXY spiked conc. (ng/ml) | Found C (ng/ml)_Found C (ng/ml) | Intra‐day (n = 5)_RSD (%) | Intra‐day (n = 5)_RE (%) | Inter‐day (n = 5)_RSD (%) | Inter‐day (n = 5)_RE (%)
80 | 79.5 ± 3.2 | 7 | −2.7 | 4.1 | −0.7
20 | 20.1 ± 0.6 | 4.2 | 1.7 | 2.8 | 0.4
4 | 4.0 ± 0.3 | 2.9 | 9.7 | 7.6 | 0.04
2 | 1.9 ± 0.08 | 4.3 | −6.8 | 3.9 | −3.9

## Table 2
Unnamed: 0_level_0_Unnamed: 0_level_1 | OXY (ng/ml)_80 | OXY (ng/ml)_20 | OXY (ng/ml)_4
Recovery (%, n = 5) | 102.6 ± 3.6 | 111.1 ± 2.6 | 102.0 ± 4.3
Matrix effects (%, n = 3) | 104.0 ± 0.5 | 107.6 ± 3.1 | 101.9 ± 3.6

## Table 3
Storage condition_Storage condition_Storage condition | OXY (%RSD)_HQC_80 ng/ml | OXY (%RSD)_MQC_20 ng/ml | OXY (%RSD)_LQC_4 ng/ml
Post‐preparative stability 8 h in autosampler at 8°C | 0.3 | −0.2 | 0.3
Bench‐top stability 18 h at room temperature | −7.8 | −4.6 | −8.1
Freeze–thaw stability a | −1.3 | 1.0 | −1.0
Short‐term stability 2 weeks at −30°C | −0.6 | 2.0 | 2.8
Long‐term stability 6 months at −30°C | −14.9 | −11.2 | −0.2

## Freetext Data
Data | Value | Relevance
Calibration curve range for OXY | 2–100 ng/ml | Relevant to the question of measurement feasibility; demonstrates the range within which plasma concentrations of OXY can be quantified for clinical application.
Measurement time for OXY in plasma | 4 minutes | Relevant to the question of establishing a quick method for plasma analysis; addresses the need for expedient clinical measurement.
Amount of plasma used in method | 100 μl | Relevant to the question of efficient resource usage; demonstrates minimal biological sample requirements for the procedure.
Preparation time for sample | Approximately 40 minutes | Relevant to the question of method establishment; contributes to the practicality of the LC–MS/MS system for clinical adjustments in OXY dosing.
Stability testing of OXY in plasma at -30°C | Stable for 6 months | Relevant to the question of method validation; confirms reliable storage conditions for long-term analysis.
Maximum plasma concentration reported in pharmacokinetics (10 mg dose) | 9.81 ± 2.74 ng/ml | Relevant to establishing clinically relevant ranges; aligns with therapeutic range addressed by the paper’s system.
Intra-day validation precision (RSD) | <8% | Relevant to method reliability; ensures reproducibility of results for clinical applications.
Inter-day validation accuracy (RE) | ±15% | Relevant to accuracy requirements; confirms the method's reliability across different days for clinical use.
Internal standard used | Papaverine (PAV) | Relevant to system design; provides a cost-effective and consistent alternative to deuterated standards in OXY quantification.
Time for drying eluents during SPE | Reduced compared to previous reports | Relevant to method optimization; demonstrates improvement over existing methods in sample preparation duration.
Retention time for OXY and IS | Approximately 3 minutes | Relevant to LC–MS/MS measurement efficiency; allows for rapid co-detection of compounds within a short timeframe.
Post-preparative stability | Stable for 8 hours in autosampler at 8°C | Relevant to operational conditions; assures integrity of prepared samples during measurement.
Matrix effect evaluation range | Found not significant | Relevant to method robustness; eliminates concerns about interference from plasma components during analysis.
Carryover assessment | No significant carryover detected | Relevant to measurement reliability; confirms prevention of contamination between sample injections.
Plasma OXY measurement reproducibility | RSD <8% and accuracy within acceptable range | Relevant to method consistency; supports reliable integration in clinical dosing adjustments.
Blood Volume Used | 100 μl | The limited use of plasma indicates the study's compatibility with minimal sample volume requirements for PK analysis.
Patient Population | Human pooled plasma | The study was conducted using human plasma, ensuring relevance to human PK simulations.
Stock Solution Concentration | 10 mg/ml OXY injection | Concentration of the stock solution addresses preparation details for PK testing.
Sample Preparation Time | Approximately 40 minutes | Rapid sample preparation facilitates efficient PK studies.
Sampling Method | Solid-phase extraction (SPE) | SPE ensures reliable analyte extraction for accurate PK evaluation.
Analytical Run Time | 4 minutes | Short analytical time supports a high-throughput PK study design.
Standard Curve Range | 2–100 ng/ml | The calibration range reflects expected therapeutic and toxic plasma concentration levels relevant to PK assessments.
Measurement Method | LC–MS/MS | Use of LC–MS/MS is appropriate for precise quantification needed in PK studies.
Stability Studies | Conducted under bench-top, freeze-thaw, and long-term conditions | Stability assessment is essential for ensuring reliable PK data over time.
Time Frame for Stability Checks | Up to 6 months | Long-term viability of the data supports extended PK evaluation scenarios.
Injection Volume | 2 μl | Small injection volume highlights the method's sensitivity, a critical factor in PK sampling efficiency.
Ionization Method | ESI positive mode | Electrospray ionization ensures accurate detection necessary in PK studies.
N/A | N/A | The paper focuses on developing a method to measure oxycodone in plasma using LC-MS/MS, but it does not report or utilize a pharmacokinetic model such as 1-compartment, 2-compartment, physiologically based model, or IVIVE.
Nebulizer gas | 3.0 L/min | These parameters optimize the LC–MS/MS system's instrumental conditions to accurately measure OXY in plasma samples.
Desolvation line temperature | 250°C | Selected to aid in optimal ionization of OXY and IS under electrospray ionization (ESI) conditions.
Heat-block temperature | 400°C | Helps control the thermal environment during ionization for efficient LC–MS/MS detection of OXY and IS.
Drying gas | 10.0 L/min | Supports solvent evaporation and ensures stable ionization for LC–MS/MS analysis.
Q1 Pre Rod Bias for OXY | 16 V | These voltages were specifically optimized for OXY detection under positive mode to establish a reliable LC–MS/MS assay.
Collision Energy for OXY | 19 V | Ensures generation of fragment ions for reliable multiple reaction monitoring (MRM) transitions.
Q3 Pre Rod Bias for OXY | 20 V | Fine-tuned values for detecting OXY precursor and product ions accurately.
Transitions for OXY | m/z 316.0 → 298.2 | Defines the MRM transition selected for quantifying OXY to avoid interference and ensure signal specificity.
Q1 Pre Rod Bias for IS | 17 V | Optimized for stable ionization of IS (Papaverine) to serve as a reference in quantifications.
Collision Energy for IS | 30 V | Produces specific fragment ions (m/z 340.4 → 324.1) to enable consistent internal standard monitoring.
Q3 Pre Rod Bias for IS | 21 V | hese optimized voltage parameters contribute to accuracy in OXY measurement through LC–MS/MS system calibration.
Mobile phase | 1 mM ammonium formate and acetonitrile (50:50) with 0.1% formic acid | Ensures proper separation and peak shape of OXY and IS on the Unison UK silica column during the chromatographic process.
Plasma OXY range | 2–100 ng/ml | Linear calibration curve range addresses therapeutic and toxic plasma concentration levels.
Sample injection volume | 2 μl | Small volume ensures sample economy while achieving effective quantitative detection in LC–MS/MS.
Column temperature | 40°C | Maintains reproducible chromatographic retention times and peak shapes during OXY analysis.
Autosampler storage temperature | 8°C | Ensures sample stability during batch runs for LC–MS/MS analysis.
Plasma processing conditions | Solid-phase extraction using hydrophilic–lipophilic balance cartridges | These parameters were designed to recover OXY and IS efficiently from plasma.
Analyte retention time | ~3.0 min | Short runtime facilitates rapid measurement of OXY, which is important in clinical settings for dose adjustment.
Reproducibility criteria | <15% relative standard deviation (RSD) | Ensures the reliability and precision of the developed LC–MS/MS method.
Recovery rates | OXY extraction reproducibility <5% RSD | Confirms efficiency and consistency of the solid-phase extraction step critical in plasma measurements.
Freeze-thaw cycles stability | Up to 3 cycles | Stability assessment ensures confidence in sample handling for clinical and research purposes.
Long-term storage | Stable at −30°C for 6 months | Ensures practicality of sample storage for ongoing OXY concentration monitoring.
Relevance | Lab Solutions LCMS software | Used for data acquisition and integration to calculate plasma OXY concentrations from raw signals.
Age | Various backgrounds including age influence the range of plasma OXY concentrations for pain relief. | Age is associated with individual variability in plasma OXY pharmacokinetics due to differences in metabolism and drug sensitivity.
Cancer progression | Plasma OXY concentration variability is influenced by cancer progression. | Cancer progression might alter drug metabolism, affecting OXY pharmacokinetics.
History of opioid therapy | Range of plasma OXY concentration varies with a history of opioid therapy. | Prior exposure to opioids may result in tolerance, influencing plasma levels required for pain relief.
Pain sensitivity | Pain sensitivity affects individual plasma OXY concentration for therapeutic efficacy. | Different sensitivities to pain may result in varying effective OXY plasma concentrations among patients.
Concomitant medications | Drug interactions with concomitant medications influence plasma OXY concentration. | Other medications could alter CYP enzyme activities, significantly affecting OXY metabolism and pharmacokinetics.
CYP3A4/5 enzyme activity | OXY is primarily metabolized by CYP3A4/5 enzymes to noroxycodone. | Variability in CYP3A4/5 activity can influence the rate and extent of OXY metabolism.
CYP2D6 enzyme activity | OXY metabolism to oxymorphone involves CYP2D6 enzymes. | Polymorphisms or variability in CYP2D6 activity can affect OXY pharmacokinetics.
Glucuronosylation by UGT | Noroxymorphone metabolism includes glucuronosylation by UDP‐glucuronosyltransferase. | Differences in UGT activity contribute to variability in OXY pharmacokinetics.