## Table 1
S/N | Recovery | RSD injection replication | Reference
1 | Procedural and spike blank |  | [83],
2 | Spiked sample | 5 injections | [84]
3 | Method and sample blank |  | [77]
4 | Procedural and spikes blank, duplicate samples |  | [85]
5 | Blank and real sample | 7 replicates | [26]
6 | Method and blank sample |  | [86]
7 | Field and equipment blank |  | [87]
8 | Matrix spike duplicate, and matrix spike triplicate |  | [52]
9 | Surrogate recovery |  | [88,89]
10 | Spiked blank sample | 7 replicates | [90,91]
11 | Solvent blank, a matrix blank and a matrix spike and its duplicate |  | [92]
12 | Spiked blank samples |  | [93]
13 | Field and method blank |  | [94]
14 | Field and lab blank | 10 replicates | [95]
15 | Solvent, field, and procedural blank |  | [66]
16 | Procedural, matrix and sample blank |  | [96]
17 | Spiked sample | 5 replicates | [19]
18 | Field blank |  | [51]
19 | Field and procedural blank | 3 replicates | [97]
20 | Procedural blank |  | [31,98]
21 | Method blank, parallel samples, blank solvent recovery surrogate, internal standard, QC standard sample |  | [99]
22 | Blank sample |  | [100]
23 |  | 6 replicates | [101]
24 |  | 7 replicates | [102]
25 | Blank matrix | 7 replicates | [103]
26 | Field blank |  | [104]
27 | Duplicate samples, blank and spike samples, CRM |  | [105]

## Table 2
MATRIX | No. of COMPOUNDS | LOD (liquid: ng/mL; Solid: ng/g; Air: ng/m3) | LOQ (liquid: ng/mL; Solid: ng/g; Air: ng/m3) | RECOVERY(%) | RSD (%) | INSTRUMENT/INSTRUMENTAL CONDITIONS/TOTAL RUN TIME | REFERENCE
River | 7 | 2 × 10-−6–5.1 × 10−5 | 1 × 10−5–1.69x10-4 | 73.3–115.5 | 4.5–14.6 | GC-MS. “splitless mode with initial column flow rate of 1.9 mL min− 1. Initial temperature of oven at 60 °C for 3 min, 110 °C (2 min) at the rate of 5 °C min− 1, reaching 200 °C at 8 °C min− 1 and finally attaining the temperature of 315 °C at 5 °C min− 1 with final holding time of 10 min” | [106]
River water | 14 |  |  | 66–80 |  | GC-MS/MS: “The initial temperature of oven was set at 150 °C for 3 min, and then raised to 290 °C at a rate of 4 °C/min and held for 10 min” | [66]
River water | 8 | 0.001–0.005 |  | 81.6–109 |  | GC-ECD: Split. “the column temperature starts at 160 °C with a hold time of 2 min and increased to 270 °C at a rate of 4 °C·min−1 with a 2-min hold. Tail gas flow was 60 mL/min” | [34]
Snow | 13 | 1.1x10-6-2.6x10-5 | 3.2x10-6-7.8x10-5 | CRM: (27–137Internal standard: 70–98 | 10–25 | GC-MS: “Initial temperature of oven at 40 °C for 2 min, 150 °C (0 min) at the rate of 25 °C min− 1, reaching 200 °C at 0 °C min− 1, 280 °C at 10 °C min− 1 and finally attaining the temperature of 300 °C at 5 °C min− 1 with final holding time of 5 min” RT 12.90–13.90 | [80]
Glacier-ice core | 25 | 4 × 10−6 -1.53 × 10−4 |  | 9.67–112.4 |  | GC-MS: splitless, “temperature program: 80 °C (2.00 min) to 110 °C at 7 °C min-1 then 3 °C min-1 to 250 °C and a final ramp of 10 °C min-1 to 285 °C with a hold of 5 min” ECDs: “temperature program: 80 °C for 2.00 min, then 10 °C min-1 to 150 °C then 2 °C min-1 to 280 °C hold for 5 min. Hydrogen was used as a carrier at 1.0 mL min-1” | [81]
Water wells | 20 | 0.02–0.74 | 0.05–2.46 |  |  | GC-MS: “the initial temperature of 90 °C (hold 1.3 min), increased to 125 °C at 15 °C/min, 165 °C at 5 °C/min, 195 °C at 2.5 °C/min, and finally, it was increased to 280 °C at 20 °C/min (hold for 4 min)”. RT 18.38–30.59 | [107]
farmland, river, and fishpond water | 4 | 0.46–0.97 nMa |  | 90.20–109.4 | 4.23 | GC-MS | [36]
Seawater | 10 | 0.0001–0.0004 |  | Standard spike:20–115; surrogate: 66-84 |  | GC-MS/GC-ECD | [35]
Seawater and sediment | 50 | 1 × 10−5–5.6 × 10−4 (seawater) 0.01–1.04 (sediment) |  | 69.5–118.2; 75.2–120.4 | 9.5 (water); 8 (sediment) | GC-MS | [108]
Surface water and Sediment | 18 | Water: 9 × 10−6-2.5 × 10 −4; sediment: 0.018–0.500 |  | Water: 690–122; Sediment: 77-108 |  | GC-MS/MS: Splitless. “the GC temperature programs were as follows: 40 °C for 1 min; 40 °C–120 °C (40 °C min−1), held for 0 min; 120 °C–240 °C (5 °C min−1), held for 0 min; 240 °C–300 °C (12 °C min−1)” | [109]
Water, sediment | 22 | IDL:2.9–112.2 pga; MDL: 5-227 |  | 22.42 % | 0.75–57.2 % | “Gradient temp: 80 °C for 0.75 min, 40.0 °C/min to 170 °C, 2.5 °C/min to 236 °C, then 40.0 °C/min to 275 °C, hold 9.62 min. Run time: 40.0 min. 1 μL injection” | [82]
Water and sediment | 11 | Water: 0.2–0.4; sediment: 1-3 | Water: 0.7–1.3; sediment: 4-15 | Water: 94–98; sediment:94-103 |  | GC-ECD: “The column temperature was programmed from 90 °C (3 min) at a rate of 25 °C min to 200 °C, fixed for 15 min and then at a charge per unit of 30 °C min to 265 °C held for 5 min to 275 °C (15 min) at 3 ° C min−1”. | [110]
Sediment, river water | 23 |  | MDL:RW: 2.6 × 10−7 –9.2 × 10−6; sed: 1.8 × 10−4–7.6 × 10−4; 1.5 × 10−4–4.1 × 10−4 | RW 62.8–86.3; sed: 66.4–84.3; muscle: 73.1–105 | <15 | GC-HRMS: “The oven program was 100 °C (1 min), 20 °C/min to 200 °C, and 2.5 °C/min to 270 °C (5 min)” | [111]
Water, sediment, benthos, phytoplankton | 20 | MDLs: Water: 1 × 10−5–2.57 × 10−3; sediment: 0.01–2.81 |  | 80–112 | <5 | GC-μECD | [112]
Surface water | 20 | 1 × 10−6-7.2 × 10−5 |  | Surrogate recovery: 88–117 and 85–110 |  | GC-ECD: Splitless, “the initial temperature of 120 °C was ramped to a final temperature of 290 °C with a hold time of 3.5 min” | [113]
Water | 10 | MDL: 0.00010 to 0.00040 |  | 20 to 115; surrogate: 66 to 84 |  |  | [114]
water | 19 | 250–3400 |  | 89–105 | 3.9–13.2 | GC-MS: “an initial temperature of 50 °C, which was then to 180 °C at 20 °C/min and held for 1.5 min. This was then increased to 200 °C 3 °C/min and held for 3 min. Finally, it was increased to 230 °C at 5 °C/min and held for 5 min. The total run time was 28.667 min” | [115]
water | 19 | 0.00001–0.002 |  | 87–105 |  | GC-ECD, GC-MS: Splitless, “the column temperature was programmed from 60 °C to 170 °C at 10 °C/min, a hold for 2 min, followed by an increase to 280 °C at 5 °C/min, and a hold for 3 min, and finally at 15 °C/min to 300 °C″
Water | 14 | IDL: 0.001–0.116 | IQL: 3.6–388 | 67.3–129 | 9.2–18.1 | HPLC-MS/MS: “The gradient elution program was as follows: 0–0.5 min, 10 % B; 3.0 min, 90 % B; 11.0 min, 90 % B; 11.1 min, 10 % B; 11.1–15.0 min, 10 % B″ | [116]
River and pond water | 6 | 0.000011–0.000058 | 0.00009–0.0.00193 | River: 85.9–122; pond: 82.9–96.2 | 2.7–14.5 | GC-MS: Splitless, “the initial oven temperature was at 80 °C and held for 0.5 min, increased to 200 °C at a rate of 30 °C/min, then increased to 250 °C at a rate of 20 °C/min and held for 2 min, and finally increased to 300 °C a rate of 25 °C/min and held for 2 min”. RT 13 min | [117]
Agricultural effluents, WWTP | 20 | 0.0001–0.005 | 0.0005–0.01 |  |  | GC-MS: “the initial temperature of 80 °C (hold 1 min) increased to 170 °C at 10 °C/min (hold 10 min) then increased to 230 °C at 4 °C/min and finally increased to 280 °C at 3 °C/min (hold for 2 min), with a total acquisition time of 41 min” | [118]
Water, DP, SPM | 25 | MDLs: 0.00021 to 0.00152 | MQLs: 0.00070 to 0.00507 | surrogate spiked: 61.3–108.6 |  | GC-MS: RT. 13.47–21.72 min | [119]
Surface and ground water | 7 | 0.00008–0.00018 |  | 78–91 | 2–6 | GC-MS: “The oven temperature began at 80 °C–200 °C at a rate of 10 °C per minute (5 min hold time) and increased to 270 °C at a rate of 5 °C per min (5 min hold time)” | [120]
Surface and groundwater | 8 | 0.00003–0.00018 |  | 85–92.7 | 3.9–11.2 | GC-ECD: “The oven temperature began at 60C for 1 min and increased to 290C (10 min hold time) at a rate of 4 C/min” | [121]
Virtual Organism, sediment | 30 |  | VO: 0.0043–0.0435; sediment: 0.00039–0.0276 | VO: 22–72; sediment: 19-90 | VO: 13–39; sediment: 12-40 | HRGC/HRMS | [27]
Water and sediment | 18 | Water: 8.8 to 16.4; sediment: 0.0219 to 0.0298 | water: 0.0279 to 0.0497 sediment: 0.0698 to 0.0902 | Water: 74–110; sediment: 91–125 | 0.9–1.8 | GC-MS: Splitless, “the oven temperature was set at 150 °C and held for 2 min, raised to 270 °C at 14 °C/min and held for a further 2 min” | [38]
Water and sediment | 30 | Water: 0.00003–0.0001; sediments: 0.00050–0.001 |  | Water: 88.2 to 103.1; sediment: 72.7 to 95.5 |  | GC-MS: “The oven temperature of the GC was initially held at 70 °C for 5 min and raised to 280 °C at 4 °C/min and kept for 5 min”” | [39]
WATER AND SOIL | 26 | WATER: MDLs: 0.00001–0.00089; SOIL: 10-980 |  | Water: 41–113 soil: 57-121 |  | GC−MS/MS: “The oven temperature was programmed as follows: the initial temperature (50 °C) was held for 1 min and increased by 25 °C/min to 125 °C and then increased by 10 °C/min to 300 °C and held for 2 min” | [122]
Water sample | 5 | 240–330 | 740–980 |  | 1.4–3.5 | GC-ECD: “The oven temperature was set as follows: initial oven temperature of 170 °C for 2 min, increasing to 300 °C at a rate of 10 °C/min”. RT 7.45–13.26 | [123]
Water | 8 |  |  | Mixed reference sample: 70–140; surrogate: 70-130 | <5 | GC-MS: Splitless,. “the initial oven temperature was 80 °C, which was maintained for 1 min, then raised to 160 °C at a rate of 30 °C/min, maintained for 1 min, and finally raised to 265 °C at a rate of 3 °C/min and maintained for 1 min” | [124]
Water | 24 | MDL: 0.00001 to 0.00002 |  | Spike mean: 66; surrogate mean: 92 | <20 | GC-MS; GC-ECD | [99]
Water | 8 | 0.0024–0.0462 |  | 77.7–106.3 |  | GC-MS: “oven temperature program began at 100 °C (held for 0.5 min) and was raised to 175 °C at 40 °C·min−1 (held for 2 min), before ramping to 190 °C at 5 °C·min−1. It was then further increased to 205 °C at 20 °C·min−1” | [125]
Surface water | 13 | MDLs: 0.000022–0.000069 |  | Surrogate: 80–110 | 1.50–5.54 | GC-ECD: “initially at 80 °C (equilibrium time 1 min), increased to 150 °C at the rate of 20 °C/min and held for 2 min, before reaching at 300 °C at the rate of 5 °C/min, and then held for 5 min” | [126]
Surface water | 9 | 0.000002–0.00003 | 0.000003–0.000061 | Int std: 72-93 | <1.8 | GC-MSD: “The column temperature was initially held at 50 °C for 1 min, then increased to 160 °C at 7 °C/min and finally increased to 240 °C at 10 °C/min, and was held for 18 min” | [127]
Surface water | 14 |  | Water: 100; SM: 0.1 | 92–120 |  | GC-MS | [128]
Water lake | 20 | MDL: 2.81 e−6-1.58e-4 | MQL: 2.81e-06-0.000158 |  | 1.4–11.9 | EC-ECD | [129]
Surface water | 7 |  | 0.0006–0.003 | 80–115 | 8–15 | GC-ECD: Split, “temperature program of hold for 1 min at 100 °C, drop the temperature at 10 °C/min to 240 °C and hold for 1 min, drop at 1 °C/min to 260 °C and hold 1 min and finally drop at 10 °C/min to 300 °C and hold 10 min” | [130]
Surface water | 13 | 0.000006–0.00002 |  | 82–124 | <20 | GC-ECD: “Oven temperature (Initial) 100 °C (1 min) → 180 °C at 25 °C/min (2 min), ramp 5 °C/min → 240 °C (1 min), and at 4.5 °C/min to 260 °C (2 min) → at 10 °C/min to 280 °C (5 min). RT: 8.839–23.195 min” | [131]
Water | 8 |  | 0.001–0.01 | 94.5–106.8 |  | GC-MS: “initial temperature of 80 °C which was maintained for 1 min, followed by an increase in temperature to 230 °C at a rate of 10 °C min −1, which was then maintained for another 4 min” | [132]
water |  | 0.00028–0.01655 | 0.00092–0.05516 | 63.6–125 | 1.03–17 | GC-MS | [133]
Water | 4 | 40–97 | 12–29 | 91.8–103.5 | 2.7–4.5 | MSPE-HPLC-UV: “The column temperature was set at 30 °C and the detection wavelength was set at 238 nm” | [134]
Water | 51 | MDLS:0.0001–0.0031 |  | 75.2–105.8 | 1.8–11.30 | GC-MS | [135]
water | 10 | 0.00006–0.003 | 0.0002–0.01 | 91–109 | 3–10 | GC-MS: “The oven was programmed at 80 °C and increased to 290 °C at 20 °C/min and held at 290 °C for 4.75 min” | [45]
Lagoon water | 7 |  |  | 51–120 |  | GC-ECD: “The initial oven temperature is 60 °C (for 2-min hold), raised to 160 °C at 20 °C/min then 200 °C (for 10-min hold), and was maintained at 250 °C for 2 min” | [136]
Drinking eater | 5 | 305–2340 |  | 87–92 | 4.16–6.06 % | “The oven temperature was programmed with an initial isothermal at 150 °C held for 2 min, which was then increased to 210 °C at 10 °C/min, and was held for 1 min, to increase the temperature to the final value of 300 °C at 5 °C/min and held for 5 min” | [137]
Drinking water | 6 | 10–20 | 40–200 |  |  | GC-MS:“The oven temperature was set at 80C and kept at this temperature for 1 min. Then, the temperature was increased to 175C with a ramp of 30C/min and then kept for 4 min. Next, the oven temperature was increased to 225C with a ramp of 3C/min and kept at this final temperature for 10 min” | [138]
Tap water | 26 | 7e-07-0.00009 |  | Spiked: 72–122; internal std: 84-108 | <4 | GC-MS: “80 °C for 1 min; 80 °C-150 °C (20 °C/min); 150 °C-287.5 °C (5 °C/min); with temperatures of 150 °C, 230 °C, and 280 °C for the quadrupole, ion source, and interface, respectively” | [139]
Surface water, SPM | 14 |  | MQLs: 2.5e-05–5e-05 | Water: 63–127;SPM: 72-126 | <5 | GC-ECD: Splitless, “oven heating started at 90 °C and 1min hold, ramped to 210 °C at the rate of 10 °C/min with 1min hold, and then to 230 °C at 1 °C/min with 10min hold, and finally, it increased to 250 °C at 1 °C/min” | [140]
Water, SPM | 26 | 0.00024–0.00183 | 0.0008–0.00614 | SF: 62.5–126; SPM: 60.8–114.8 |  | GC-MS: “The oven temperature was programmed as follows: initially at 70 °C (equilibrium time 1.5 min), increased to 200 °C at the rate of 10 °C/min, continually increased before reaching at 320 °C at the rate of 7 °C/min, and then held for 3 min” | [7]WWEE
Water, WTP | 6 | 25–512 | 800–1630 |  |  | GC-MS: “The oven temperature was started at 80 °C and kept at this temperature for 1 min. Then, the temperature was increased to 175 °C with a ramp of 30 °C/min and then kept for 4 min. Next, the oven temperature was increased to 225 °C with a ramp of 3 °C/min and kept at this final temperature for 10 min” | [141]
Water, sediment | 25 | Water: 6.3–21; sediment: 1.3–4.1 | Water: 21–68 ng/mL; sediment: 4.2–14 | CRM: 79.2 to 117; Water: 46.6 to 119; sediment: 86.7–117 | Water: 0.05–20; sediment: 0.7–17 | GC-MS: “initial temperature of 50 °C, held for 1 min, ramped at 5 °C/min to 180 °C, held for 3 min, and increased to 300 °C at 10 °C/min, maintained for 7 min, equating to a total run time of 49 min” | [71]
Water, Soil, sediment | 20 |  |  | Surrogate: 65 and 105 | <15 | GC-ECD: “The oven temperature program began at 100 °C, held for 1 min, raised to 200 °C at 4 °C·min−1, then to 230 °C at 2 °C·min−1, and ramped to 280 °C at 8 °C·min−1, and held for 15 min” | [142]
Water, planktons and corals | 8 | MDL: Seawater: 2 e−09-1.6e-08: marine biota; 0.0009–0.0013 |  | 74–110 |  | GC-MS/MS: “The oven temperature was programmed as follows: 80 °C for 5min, at 20 °C min-1 to 160 °C, at 4 °C min −1 to 240 °C, at 10 °C min −1 to 295 °C with a final hold for 2min”. RT: 11.4–26.3 | [89]
Water, biota | 16 | MDL: 3.3e -07-1.25 e−06 |  | 75–89 | 5.9 |  | [143]
lake | 12 | 0.01 |  | 40–100 |  | GC-MS/MS: “The GC programming was as: The primary temperature was 70 °C and hold for 2 min, then ramped up to 150 °C at 25 °C/min, then ramped up to 200 °C at 3 °C/min, finally ramped up to 280 °C at 8 °C/min and hold for 10 min” | [144]
Gas and particulate phase | 13 | 0.044–0.145 |  | 74.6, 89.6 |  | GC-MS/MS: “The oven temperature program was as follows: initial temperature of 100 °C (hold for 1min), heat to 220 °C at a rate of 5 °C/min (hold for 10min), then heat to 280 °C at a rate of 4 °C/min and hold for 5min” | [145]
SOIL | 10 | 0.012–0.061 |  | 78.80–110.48 | <10 | GC-ECD: “The column oven-heating program was as follows: initial temperature of 80 °C, ramped up to 260 °C at the rate of 10 °C⋅min− 1, then heated to 280 °C at the rate of 20 °C⋅min− 1, and held for 5 min” | [146]
Soil | 8 | 0.012–0.061 |  | 78.80–110.48 | <10 | GC-ECD: “The oven temperature of gas chromatography was set to start at 120 °C, ramped to 260 °C at the rate of 10 °C·min−1, and then to 280 °C at a rate of 20 °C·min−1, and finally held for 5 min” | [147]
Soil | 23 | MDL: 0.004–0.360 |  | 53.9–112 | 1.6–8.5 | GC-HRMS: Splitless, “the oven temperature program was as follows: 100 °C (1 min) → 20 °C/min → 180 °C (5 min) → 0.5 °C/min → 186.5 °C (20 min) → 10 °C/min → 250 °C (24 min) → 40 °C/min → 300 °C (30 min)” | [90]
Soil | 24 | 0.00011–0.00549 |  | Surrogate: 60-107 | <10 | GC-MS/MS: “The temperature program of the GC oven was initially set at 100 °C, then increase to 200 °C at the rate of 10 °C/min, keep heating to 230 °C at the rate of 1 °C/min with the holding time of 1 min, and finally reach 290 °C at the rate of 10 °C/min with the holding time of 10 min” | [94]
Soil | 19 |  |  | 86–135 |  | GC/HR-MS: “The GC temperature was programmed as follows: 60 °C was held for 1.5 min, then increased to 140 °C at a rate of 10 °C/min, then to 220 °C at a rate of 4 °C/min, 250 °C at a rate of 2 °C/min, and finally to 300 °C at a rate of 8 °C/min”. “Chromatographic analysis was performed at the following temperatures: 120 °C for 1 min and then increased to 180 °C at a rate of 20 °C/min, then to 260 °C at a rate of 2 °C/min, and finally to 300 °C at a rate of 5 °C/min for 4 min” | [95]
Soil | 3 | 0.022–0.024 |  | 97.8–111.3 | 7.1–13.1 | GC-MS: Splitless. “the column temperature was at 80 °C for 1 min, increased to 150 °C at a rate of 20 °C/min, then increased to 210 °C at a rate of 5 °C/min, and finally increased to 280 °C at a rate of 10 °C/min; thereafter, the temperature remained constant for 5 min” | [148]
Soil | 15 | 500–2000 | 1000–2x104 |  |  | GC-MS: Splitless, “the oven temperature was programmed starting at 45 °C and held 0.5 min, followed by an increase of 30 °C min─1 up to 90 °C, held 0.5 min, then 10 °C min─1 up to 310 °C for 3 min” | [149]
Soil | 28 | 0.036–0.183 |  |  |  | GC-ECD: “the oven temperature began at 100 °C (equilibrium time 1 min), rose to 200 °C at 4 °C/min, then to 230 °C at 2 °C/min, and at last reached 280 °C at a rate of 8 °C/min, held for 15 min” | [150]
Soil | 25 | 0.018–0.058 | 0.062–0.199 | Surrogate: 59.2–124 |  | GC-ECD: “The oven temperature was set to start from 100 °C (held for 1 min), and then increased by a rate of 4 °C/min to 200 °C, by a rate of 2 °C/min to 230 °C, and at last by a rate of 8 °C/min to 280 °C, with a final holding time of 15 min” | [151]
Soil | 25 | MDL: 0.001–0.02 |  | 74.8 and 86.9 | <20 %; | GC-ECD/GC-MS: “initially 100 °C for 1 min, 4 °C/min to 200 °C, 2 °C/min to 230 °C, and 8 °C/min to 280 °C for 15 min” | [152]
Soil | 26 | MDLs: 0.002–0.008 |  | 72.8–100.3 | 1.5–7.7 | HRGC-HRMS: Splitless, “120 °C (2 min hold)-20 °C/min to 200 °C (2 min hold)-5 °C/min to 240 °C (2 min hold)−2.5 °C/min to 270 °C (2 min hold)-5 °C/min to 300 °C (2 min hold).” Total 35 min” | [153]
Soil | 20 | 0.05–0.5 | 0.16–1.55 | 80.3–109 | <10 | GC-MS: “The initial oven temperature was maintained at 100 °C for 2 min, ramp to 180 °C at a rate of 15 °C/min, and raised to 300 °C at a flow rate of 3 °C/min and held for 9 min” | [154]
Soil | 18 | MISOLEX: 0.005–1.6; QuEChERS: 0.001–1.48 | MISOLEX: 0.02–3.85; QuEChERS: 0.04–4.93 | MISOLEX: 65.8–180.9; QuEChERS: 67.8–169.3 | MISOLEX: 1.3–22.4; QuEChERS: 1.2–40.9 | HS–SPME–GC–MS: SIM, “the initial oven temperature was set to 90 °C and held for 3 min. The temperature was ramped to 150 °C at a rate of 15 °C·min−1. Then, it was ramped to 280 °C at a rate of 5 °C·min−1 and held for 3 min. RT 18.99–33.44 min” | [75]
Soil | 213 Pesticides; 12 OCPs |  |  | OCPs: 69–140; Pesticides: 21–197.3 | OCPs: 8.2–41.6; pesticides: 0.8–63.6 | LC-MS: “The sheath gas was set at 12 L min−1 at 330 °C. The desolvation and nebulizing gas temperature was 190 °C and the flow rate was 11 L min−1 with a pressure of 26 psi. The capillary voltages were set at 3900 and 2600 V in positive and negative ionization mode, respectively. The cycle time was 700 ms and dwell time 3–83 min; GC-MS: Splitless, “The column temperature was maintained at 80 °C for 1.8 min, increased to 170 °C at a rate of 40 °C/min, then increased to 310 °C at a rate of 10 °C/min and held for 3 min” | [76]
Soil | 12 | 0.048 and 3.125 | 0.5 and 20 | 67.9–119.6 | 3.1–31.6 | GC-MS: “The temperature ramp was programmed as follows: (a) 80 ◦C—1.8 min; (b) 80 °C to 170 °C at a rate of 40 °C min−1; (c) 170 °C to 310 °C at a rate of 10 °C min−1; (d) 310 °C for 3 min. The total time for each analysis was 20.75 min” | [15]
Soil | 8 | DL: 0.002–0.05 | QL: 0.158–0.169 | 80 to 110 | 5–10 | GC-ECD: “The oven temperature of gas chromatography was set to start at 120 °C for 1 min, ramped to 240 °C at the rate of 7 °C/min and held for 5 min. RT: 9.07–16.912 min” | [155]
Soil | 10 |  |  | 50–120 |  | GC-ECD: “initial oven temperature was 100 °C (equilibrium time 1 min), raised to 200 °C at a rate of 4 °C·min−1, then increased to 230 °C at 2 °C·min−1, then raised to 280 °C at 8 °C·min−1 and held for 15 min” | [156]
Soil | 14 | 0.01 |  | 81–118 |  | GC-MS | [157]
Soil |  | 0.03–0.20 |  | 90.1–93.7 |  | GC-MS: Splitless, “the initial oven temperature was maintained at 100 °C for 2 min, ramp to 180 °C at a rate of 15 °C/min and then raises to 300 °C at a flow rate of 3 °C/min and held for 9 min” | [158]
Soil | OCPs (23), CUPs(24) | 0.001–7.83 | 0.003–26.1 | Procedural blank: 70-130Spiked blank:71.3–126 | <20 | GC-TQ-MS: Splitless, “held in 1 min at 60 °C, then ramped to 120 °C at the rate of 40 °C/min, and finally up to 310 °C with 5 °C/min and held for 3 min” | [83]
Soil | 8 | 4–10 | 2.4–32 | 75–10 | 0.1–4.2 | GC-MS: “The temperature of 80 °C lasted for 2 min, then raised to 210 °C (10 °C/min), and remained constant for 10 min, and further raised to 250 °C (10 °C/min) and remained constant for 5 min (PCBs and OCPs)” | [159]
Soil | 4 | 0.002–0.003 |  | 75–105 | <10 | GC-ECD: “The initial temperature was held at 100 °C (equilibration time of 1 min) and raised to 200 °C at 4 °C/min, then to 230 °C at 2 °C/min, and finally raised to 280 °C at a rate of 8 °C/min, and held for 15 min” | [160]
Soil | 20 | 0.1 |  | 72.5–124.1 | 1–6.7 | GC-ECD: “The oven temperature was set at 80 °C initially for 2 min, and increased to 170 °C at the rate of 25 °C/min, then ramped to 250 °C at 5 °C/min and held for 2 min, and then raised to 280 °C at 25 °C/min and eventually held for 2 min” | [161]
Soil | 23 | 0.04–1.21 |  | 63.9 and 112.1 | 0.4 and 26.2 | GC-MS/MS | [162]
Soil | 14 |  |  | 98–102 |  | GC-MS/MS: “temperature program 100 °C, 1 min; 10 °C/min to 160 °C, 4 min and 10 °C/min to 260 °C, 2 min, transfer line temperature 280 °C, total analysis time 19.2 min” | [163]
Soil | 8 | 0.0045–0.034 |  | 78–107 | 3.0–6.9 | GC-QTOF-MS: Splitless, “the GC oven temperature profile was started at 100 °C and maintained at this temperature for 3 min, then it was ramped to 170 °C at a rate of 5 °C/min, subsequently ramped to 300 °C at a rate of 10 °C/min, and then the back-flush was performed for 0.5 min at 310 °C″ | [164]
soil |  | 0.02–0.15 |  |  |  | GC-ECD: Split, “initial run temperature 80 °C, ramped at 50 °C/min to 210 °C, then ramped at 1 °C/min to 220 °C and held for 1 min, and finally ramped at 2 °C/min to 230 °C″ | [165]
soil | 14 | IDL: 0.001–0.013 |  | 63–127 |  | GC-ECD: “Increase the temperature to 80 °C during 1 min, then increase the temperature to 150 °C by the rate of 15 °C/min, up to 250 °C with 5 °C/min, and eventually up to 300 °C during 5 min” | [166]
Soil | 4 | 0.628–3.68 | 2.093–12.27 | 81.42–110.7 | 1.68–9.43 | GC-ECD: Split, column temperature: 250 °C″ | [167]
Soil | 6 | 0.10 |  | 86–90 | 3–10 | ECD: “0 °C for 1 min, 30 °C/min to 180 °C, 180 °C, for 1 min, 3 °C/min to 205 °C, 205 °C for 4 min, 20 °C/min to 290 °C, 290 °C for 7 min”; GC-MS/MS: “70 °C for 2 min 25 °C/min to 150 °C for 0 min; 3 °C/min to 200 °C for 0 min; 8 °C/min to 280 °C for 10 min hold time” | [168]
Topsoil | (8 OCPs and chlorpyrifos; 32 CUPs | UPLC-MS/MS: 0.002–0.23; GC–MS: 0.1–1.2 | UPLC-MS/MS: 0.02–0.75; GC–MS: 0.3–4.3 | 76.7–110.3 | 1.7–11.8 | GC-MS, UPLC-MS/MS | [169]
Soil, moss | 7 | 0.006–0.032 |  | 76–114 | <4 | GC-ECD: “The column temperature was from 90 °C (with 1 min hold) to up to 180 °C at 30 °C/min (1 min hold), and then up to 240 °C at 2 °C/min with a hold of 20 min and finally up to 260 °C at 5 °C/min with an 8 min hold” | [170]
Surface soil | 20 | 0.001–0.0028 |  | 85–110 |  | GC-MS: “Oven temperature for GC-MS analysis was set at 60 °C for 1 min and increased to 290 °C (10 min hold time) at the rate of 4 °C/min” | [90]
soil, spring water, river water, spring sediment, and river sediment | 24 | MDL: water: 1 × 10−5-2x10−5 sediment/soil: 1 × 10−5-2x10−5 |  | Surrogate: 70.8 and 86.6 | <20 | GC-ECD: “The GC oven temperature program was set as: initially 100 °C for 1 min, 4 °C/min to 200 °C, 2 °C/min to 230 °C, and 8 °C/min to 280 °C for 15 min” | [171]
Air, soil water, sediment | Endodulphan | MDL: Air: 2.5 × 10−4-5.8 × 10−4 soil & sediment: 2.95 × 10−3-5.59 × 10−3; water: 4.34 × 10−6 -3.02 × 10−5 |  | Air:51–112, soil: 57–102, water: 57–105 sediment: 53–107 |  | GC-HRMS: “Splitless, the initial oven temperature was set at 80 °C, raised to 150 °C at 20 °C/min, then to 235 °C at 5 °C/min, and finally to 300 °C at 15 °C/min, where it was held for 5 min” | [91]
Dust | OCPs |  |  | 77 and 95 |  | GC-MS: “The initial temperature of the column chamber was 60 °C and held for 1 min, then increased at a rate of 4 °C/min to 280 °C and held for 10 min, then increased at a rate of 10 °C/min to 300 °C and held for 10 min” | [172]
Dust | 20 | 0.06–0.38 | 0.06–0.38 | 71–112 | 2–17 | GC-MS/MS: “The GC oven temperature program was set as follows: held at 70 °C for 5 min, ramped at 10 °C/min to 160 °C and then at 5 °C/min to 280 °C, held for 5 min, ramped at 20 °C/min to 300 °C, and finally held for 5 min” | [173]
Indoor dust | 16 | MDL: 0.004–0.2 |  | Surrogate: 46.4–120.51 |  | GC-MS: “initial oven temperature was maintained at 100 °C at the rate of 20 °C/min and later increased to 240 °C and 296 °C at the rate of 4 °C/min and 8 °C/min” | [5]
Indoor dust | 26 | MDL: 0.00131–0.0073 |  | Surrogate: 88 to 110 and 99–121 |  | GC-MS: “The GC oven program used was: 60 °C for 1 min, 30 °C/min to 220 °C held for 0 min, and 5 °C/min to 300 °C held for 15 min” | [174]
Sediment | 11 | 0.002–042 | 0.007–0.141 | 78.46–98.47 |  | GC/MS-MS TSQ: “The column temperature started at 80 °C for 2 min at a rate of 30 °C, increased to 150 °C at a rate of 5 °C/min, then increased to 280 °C at a rate of 3 °C/min, and finally increased to 320 °C and held for 1 min” | [175]
Sediment |  | 0.001–0.004 | 0.004–0.079 |  |  | GC-MS/MS “Oven temperature: 80 °C, 2 min-1, 150 °C/min to 290, 0 min, 310 °C, final hold of 1 min. RT: 11.21–22.35 min” | [176]
Soil, Sediment | 20 | MDLs: Soil: 0.0020–0.058; sediment: 0.0013 to 0.50 |  | Surrogate spike: 86-100 | <10 | GC-MS: “The GC temperature program was initially maintained at 80 °C for 2 min, and then increased to 180 °C at a rate of 20 °C/min, and then increased to 220 °C at a rate of 2 °C/min, and then increased to 245 °C at a rate of 1 °C/min, and then increased to 310 °C at a rate of 10 °C/min, finally held for 10 min” | [96]
Sediment |  | 0.005–0.048 | 0.018–0.159 |  |  | GC Ultra TSQ Quantum XLS: “Oven temperature: 90 °C, 1 min,15 °C/min to 160 °C3 °C/min to 225 °C, 6 °C/min to 305, 8 min, 3 °C/min to 310 °C, final hold of 17 min” | [177]
Sediment | 15 | 0.0029–0.35 |  | 80-120; surrogate: 79 and 73 | 3.4–14 | GC-MS: EIM mode, “initial at 80 °C hold for 2 min, 20 °C/min to 180 °C, 2 °C/min to 220 °C, 1 °C min −1 to 245 °C, 10 °C min −1 to 300 °C, hold for 10 min. RT: 17.569–29.913 min” | [97]
Sediment | 13 |  |  |  |  | LC-HRMS: “The full MS scan ranged from 100 to 900 m/z, and the collision gas was N2. The capillary temperature and probe heater temperature were set at 325 °C and 350 °C, respectively, and the spray voltage was 3000 V, RT: 0.49–9.90” | [68]
Sediment | 19 | 0.01–0.59 |  | 71–106 | 3–19 | GC-MS: “The temperature program was: initial temperature: 60 °C, 6 °C/min to 290 °C (10 min). The carrier gas was nitrogen at a flow rate of 1.2 mL/min” | [178]
Sediment | 18 | 0.04–0.20 |  | Blank and sample spiked surrogate: 90.25–110.49 | <35 | GC-MS: “The oven temperature programming was set to start at 50 °C, holding for 3 min, increased to 300 °C at a heating rate of 12 °C/min, and holding to 45min” | [179]
Sediment | 18 | 0.01–0.20 |  | >87 |  | GC-ECD: “initial temperature: 120 °C, 15 °C/min to 210 °C (1 min), 2 °C/min to 245 °C (1 min), 30 °C/min to 290 °C (3 min)” | [37]
Sediment | HCHs and DDTs | 0.00188 to 0.013 for HCHs, 0.00017 to 0.0123 for DDTs |  | Surrogates: 51-85 |  | GC-MS-ECNI | [180]
Sediment | 7 | 0.02–0.16 |  | Surrogate: 52–118; individual recoveries: 71–107; spiked sediment and blank: 75-112 |  | GC-ECD: “The following oven heating ramp was used: (1) 100 °C for 1 min, (2) 100–140 °C at 5 °C/min for 1 min, (3) 140–250 °C at 1.5 °C/min for 1 min, and (4) 250–300 °C at 10 °C/min for 10 min (total run = 90 min)” | [102]
Sediment | 22 | MDL: 0.10–1.59 | 0.34–5.30 | 87–118 |  | GC-HRMS: “120 °C(0.5 min) to 180 °C at 10 °C min. −1 to 210 °C at 4 °C min. −1 (12 min) to 300 °C at 10 °C min −1” | [181]
Sediment | 24 | MDL: 0.004–0.13 |  | Surrogate: 69.1 and 90.4 |  | GC-ECD: “initial temperature was 100 °C for 1min, raised to 200 °C at a rate of 4 °C/min, then increased to 230 °C at 2 °C/min, finally raised to 280 °C at 8 °C/min and held for 15 min” | [182]
Sediment | Insecticides(34) | 0.013–0.33 |  | Spiked samples: 40–130; surrogate: 65.3–61.5 |  | GC-MS: “The oven had an initial temperature of 70 °C holding for 1 min, and being heated to (i) 180 °C at 20 °C/min, (ii) 260 °C at 4 °C/min, and (iii) 300 °C at 15 °C/min, which was held for 6 min” | [103]
Sediment | 19 | MDL: 0.01–0.16 |  | Surrogate: 93 ± 13 and 83 ± 7; spiked standard: 75–95 | 5–12 | “The oven temperature was initiated at 100 °C (held for 1 min) and increased to 200 °C at 4 °C/min, 230 °C at 2 °C/min, and finally 280 °C at 8 °C/min (held for 15 min)” | [183]
Sediment | 24 | 0.01–0.02 |  | 71.2–106 |  | GC-ECD | [184]
sediment | 17 | 0.001–0.005 |  | Matrix: 70–109; procedural: 77-114 |  | GC-ECD: “oven-heated started at 80 °C with a 2 min hold, ramped to 190 °C at a rate of 25 °C/min, then 5 °C/min to 280 °C, and finally 25 °C/min to 300 °C with a 5 min hold” | [185]
sediment | 22 | MDL:2.78 × 10−4-0.020579 |  | 62–124 | <20 | GC-MS/MS | [186]
sediment | 6 |  |  | Spiked: 74.5–106.2; sur: 68.95 |  | GC-MS/MS: “an initial temperature of 80 °C was maintained for 2 min and increased at a rate of 20 °C/min to 180 °C and held for 5 min. Then, it was increased at a rate of 10 °C/min to 290 °C and maintained for 15 min” | [187]
Sediment |  | 0.00015–0.0003 ng/mL |  | 72–110 |  | GC-MS: “The oven temperature was programmed as follows: maintain 80 °C for 1 min, then increase at 20 °C/min to 150 °C and then at 5 °C/min to 300 °C, and then hold at the final temperature for 5 min” | [188]
Sediment | 14 | MS/NCI: 0.03–2.24; ECD: 1.2–4.6 | MCI: 0.11–7.41L; ECD: 3.9–15.2 | NCI: 88.6–120 %; ECD: 81.8–117 % | NCI: 0.2–1.7; ECD: 0.4–2.7 | ECD: “Initial temperature 90 °C held for 1 min; ramped at 30 °C/min to 200 °C, not held, 2 °C/min to 235 °C, not held, 20 °C/min to 300 °C, held for 10 min, RT: 50–60 min”; GC-MS/NCI: “Initial temperature 40 °C held for 1min; ramped 50 °C/min to 110 °C, not held; 5 °C/min to303 °C, not held: 20 °C/min to 335 °C, held for 5.4 min. RT: 45–48 min” | [189]
Sediment | 18 | MDL: 0.056–0.918 |  | Sur: 60-13 | <20 | GC-ECD: “Initial column temperature started at 100 °C, was kept for 1 min and raised to 320 °C at a rate of 4 °C/min and held for 5 min” | [190]
sediment | 21 | 0.003–0.080 |  | Spiked recovery: 81.3–94.2; surrogate recovery: 73.6–112.7 % and; 68.7–101.8 | <15 | GC-μECD and GC–MS: “Splitless, oven temperatures were programmed from 60 to 170 °C (at the rate of 25 °C/min, hold 2 min), from 170 to 250 °C (at the rate of 3 °C/min, hold 4 min), from 250 to 280 °C (at the rate of 4 °C/min, hold 1 min), and finally from 280 to 300 °C (at the rate of 20 °C/min) | [86]
Sediment | 11 | 0.05 |  | 80.26–95.89 |  | GC-MS/MS | [191]
Sediment | 20 | 0.003–0.08 |  | 78–110 | 0.19–1.93 | GC-μECD: “Splitless, the oven temperature programs for OCPs and PCBs were set as follows: holding at 60 °C for 1 min, heating to 170 °C at a rate of 10 °C/min holding for 2 min, and then heating to 280 °C at a rate of 5 °C/min holding for 2 min, and finally heating to 300 °C at the rate of 20 °C/min” | [192]
Sediment | 25 |  | MQL: 0.0025 | 53–119 |  | GC-ECD: “The oven temperature was programmed to 50 °C for 1 min, increased at 20 °C/min to 200 °C and at 10 °C/min to 300 °C, and held for 5 min” | [193]
Sediment | HCH, DDT | MDL: 0.005–0.01 |  | 79 and 97 |  | GC-ECD: “The oven temperature began at 60 °C for 1 min and increased to 290 °C (10 min hold time) at a rate of 4 °C/min” | [194]
Sediment | 21 | MDL: 0.003–0.86 |  | 78–110 |  | GC-ECD: “The oven temperature was programmed from 60 to 170 °C (2 min hold) at the rate of 10 °C/min, to 280 °C (3 min hold) at the rate of 5 °C/min, and finally to 300 °C at the rate of 15 °C/min” | [195]
Sediment | 13 | 0.25 |  | 75–115 |  | GC-ECD | [196]
Sediment |  | MDL: 0.020–0.04 |  | Matrix-spiked: 75–110; surrogate: 70-105 |  | GC-MS | [197]
sediment | DDT |  | QL: 0,002 | Sur: 55–92; spiked bl: 63–110; sample std: 76-101 |  | GC-MS/MS: “Oven temperature ramps were programmed as follows: 50 °C for 1 min, increased at 20 °C/min to 200 °C, then at 10 °C/min to 300 °C, with a final hold of 5 min” | [198]
Sediment, worm | 10 |  | 0.01–0.5 | 84–94 |  | GC-MS: “The oven temperature programme was set at 90 °C where it was held for 2 min, it was ramped to 280 °C at a rate of 6 °C/min” | [199]
Sediment | 21 | 0.003–0.080 |  | Spiked recover: 81.3–94.2; surrogate recovery: 73.6–112.7 and 68.7–101.8 | <15 | GC-μECD: “Oven temperatures were programmed from 60 to 170 °C (at the rate of 25 °C/min, hold 2 min), from 170 to 250 °C (at the rate of 3 °C/min, hold 4 min), from 250 to 280 °C (at the rate of 4 °C/min, hold 1 min), and finally from 280 to 300 °C (at the rate of 20 °C/min)”” | [86]
Sediment | 22 | MDL: 0.002–15.21 |  | 61 ± 8 | 0.4–20 | HRGC/LRMS: “Split, oven temp: 75 °C (holding time of 1 min) to 100 °C at 25 °C min−1, then to 225 °C at 6 °C min−1 (holding time of 1 min), finally to 255 °C at 6 °C min−1 (holding time of 1 min); 30 min run time” | [200]
Sediment | 22 | 0.04–0.71 |  | 82–94 |  | GC-MS: “The GC oven temperature was programed as follows: initial temperature 70 °C (2 min), then raised to 175 °C (10 °C/min) and held isothermally for 4 min, then raised to 320 °C (5 °C/min) and finally held for 1 min. The injector was set from 50 °C (0.1 min isothermal) to 250 °C (200 °C/min) held for 10 min” | [201]
Sediment | 20 | 0.03–0.2 |  | Surrogate: 88.6 and 91.7 % |  | “Splitless, the initial temperature of the oven was 90 °C with a 1-min grip, and the temperature was increased to 210 °C at the rate of 10 °C/min with a 1-min hold; the temperature was enhanced to 230 °C with increments of 1 °C/min with a 10-min hold; ultimately, the temperature was increased to 250 °C at the rate of 1 °C/min” | [202]
Sediment, soil | 4) | 1.6 to 2.3; sediment: 0.2–4 |  | Surrogate: 70-92 |  | GC-MS: “Splitless, the oven temperature was programmed to 140 °C with a 0.5 min hold, ramping at 20 °C/min to 270 °C and a 5 min hold, then to 290 °C at 3 °C/min and a 2 min hold, and finally to 310 °C at 40 °C/min and a 3 min hold” | [203]
Sediment, soil | 4 | 2 × 10−5-5x10−5 |  | 82–106 |  |  | [204]
Sediment, soil | 18 | 0.05–0.25 |  | Spiked: 88-1-102.1 | <6 | GC-MS: “The column oven temperature was initially 50 °C for 1 min, ramped up at 30 °C/min to 150 °C for 1 min, 10 °C/min to 180 °C for 3 min, 3 °C/min to 210 °C for 1 min, 5 °C/min to 250 °C, and finally 15 °C/min to 290 °C for 3 min (RT: 36 min)” | [205]
Sediment | 17 | DL: 0.42–3.1 |  | CRM: 85.1–123.4 | spiked revovery: 8.9–22.9 | GC X GC-TOFMS: “The GC oven was set to 75 °C (held for 0.5 min), then raised to 280 °C at a rate of 8 °C/min and maintained at this temperature for 2 min” | [206]
Sediment | 13 |  |  | Sur: 75-110 |  | GC-MSD: “Oven temperature was set at 60 °C for 1 min and increased to 290 °C (10 min hold time) at a rate of 4 °C/min” | [207]
Sewage sludge | 17 | 0.04–0.49 | 0.22–2.17 | 85.2–99.6 | 3.5–9.6 | GC/μ-ECD: “The oven was held at 100 °C for 1 min, and ramped up at 20 °C per min to 180 °C, at 5 °C/min to 270 °C, and at 20 °C/min to 320 °C″ | [208]
Sediment, plant | 16 | MDL: soil: 0.007–0.028; plant: 0.030–0.137 |  | 75-125; surrogate: 84.1–110.6 |  | GC-ECD, GC-MS: “The temperature programming of oven is: 110 °C (hold 2 min), raising 10 °C/minutes to 180 °C, then 5 °C/minutes to 280 °C (hold 20 min)” | [209]
Sediment, shrimp | 6 | 0.1–0.5 |  | 64.4–94.6 | <10 | GC-MS/MS: “The oven temperature was programmed to an initial temperature of 50 °C, held for 1 min, increased at a rate of 25 °C min−1 to 125 °C, and finally increased at a rate of 10 °C min−1 to 310 °C and held for 2 min” | [210]
Water, Sediment | 20 | 100–1500 |  | 65–110 | <20 | GC-MS: “Splitless, the oven temperature was programmed from 35 °C, increased to 150 °C with a ramping rate of 15 °C/min and held for 5 min, and then increased to 290 °C with a ramping rate of 3 °C/min and held for 2 min” | [211]
Water, sediment | 8 | 0.07–0.31 | 0.21–0.95 | 92.4 |  | GC-MS: “Splitless, the oven temperature program was: 70 °C raised to 140 °C (25 °C/min), a second ramp at 15 °C/min to 210 °C and a final ramp at 10 °C/min to 300 °C″ | [212]
Water, sediment | 19 |  | Water: 0.8–2; sediment: 0.075–0.122 | water: 71–106; sed: 77-110 | <15 | GC-ECD: “80–196 °C (4 °C/min rate, 2 min), from 196 to 224 °C (4 °C/min rate, 2 min), from 224 to 240 °C (4 °C/min rate, 2 min), and from 240 to 275 °C (rate 4 °C/min, 2 min)” | [213]
Sediment, SPM | 7 |  |  | 70 and 106 % |  | “The GC temperature program started at 80 °C (1.5 min hold), ramped 40 °C/min to 200 °C (18 min hold), and finally ramped 5 °C/min to 305 °C″ | [214]
DP SPM SED | 16 | D & SPM: 78.8–102.7; SED = 0.0005–0.0050 |  | Procedural blank DP:80.5; SPM: 79.3; sediment 83.7. Spiked blank:78.8–102.7 |  |  | [215]
sediment, mud | 9 | 0.5 | <5 | 80–120 | <35 | GC-QqQ/MS: “the GC oven is temperature-programmed from 80 °C (isotherm 2 min) to 300 °C at 12 °C/min, then held for 5 min until run time 25.33 min” | [216]
Sediment, aquatic food | 19 | Sediment: 0.04–0.09; 0.03–0.21 |  | Sediment: 81.7–109.8; aquatic food: 86.4–110.4 |  | GC-MS: “column temperature was initially set at 80 °C for 1 min, ramped at 30 °C per minute to 170 °C, 5 °C per minute to 240 °C, 30 °C per minute to 320 °C, and then held at 320 °C for 3 min; RT: 10.55–17.96” | [217]
Soil air, and deposited sample | 24 | For PUF: 0.05, DP: 0.00005,Soil: 0.005 |  | PUF:82; Filter: 79, DP: 100; Soil |  | GC-HRMS: “Splitless, the oven was held for 2 min, then increased to 160 °C at 10 °C/min, to 225 °C at 4 °C/min, and finally to 320 °C at 10 °C/min and held for 2 min” | [218]
Air, soil | 10 | MDL: soil: 7 × 10−5–1.13 × 10−3; air: 1 × 10−6–4.6 × 10−5 |  | Surrogate: soil:70.2, 99.5; air: 75.4, 98.6 | <10 | GC-MS/MS: “The GC oven temperature was programmed as: initially 100 °C, 10 °C/min to 200 °C, 1 °C/min to 230 °C for 1 min, and 10 °C/min to 290 °C for 10 min” | [219]
Soil and water | 8 | 10.4 (water); 490 (soil) | 58.3 (water) 1.63(soil) | 79.4–129 | 4.52–18.8 | LC-QqQ-MS/MS; GC-MS/MS: “The column temperature and ion source temperature were controlled at 30 °C and 200 °C, respectively. N2 was used as nebulizer with 40 psi, 450 °C, and 11 L min−1. The capillary voltage was controlled at 4000 V″ | [220]
Air | 22 | MDL: 0.000037–0.000269 | 0.118–0.940 | 61.2–114.5: CRM: 58.1–116.1 |  | GC-HRMS: “The initial oven temperature was maintained at 100 °C for 1 min, which was then increased to 180 °C, 200 °C, 220 °C, and 300 °C at a rate of 20 °C·min−1, 1.4 °C·min−1, 5 °C·min−1 (held for 0.314 min), and 40 °C·min−1 (held for 4.4 min), respectively” | [221]
Air | 1 | 0.01–0.05 |  | 79%-105 |  | GC-MS-ECD: “The oven temperature program was as follows: the temperature started at 100 °C, held for 1 min, and then increased at 4 °C/min to 200 °C, at 2 °C/min to 230 °C, and at last at 8 °C/min to 280 °C, with a final holding time of 15 min” | [222]
Air | 13 | 0.000011–0.000073 | 0.006–0.030 | Surrogate:62-93 |  | GC-MS: “The temperature program for GC started at 80 °C (1.5 min hold), then continued with 40 °C/min to 200 °C (18 min hold) and lastly 5 °C/min to 305 °C (no hold)” | [30]
Air | 18 | DL: 0.05 |  | 101 | 22 | GC-HRMS: “The oven was held for 2 min, then increased to 160 °C at 10 °C/min, to 225 °C at 4 °C/min, and finally to 320 °C at 10 °C/min and held for 2 min” | [223]
Air | OCPs(27); CUPs(25) |  |  | OCPs: 38.2–148; CUPs: 59.3–170 |  | GC-APCI-MS/MS: splitless, “the oven temperature programme was 90 °C (1 min hold), then 40 °C.min−1 until 200 °C, followed by 2 °C.min−1 until 240 °C, and 40 °C.min−1 until 310 °C (5 min hold)”; GC-MS: “The temperature program for GC started at 80 °C (1.5 min hold), then 40 °C.min−1 to 200 °C and finally 5 °C.min−1 to 305 °C (no hold)”; CUPs: HPLC-MS/MS | [51]
Air | 4 | IDLs: 0.011–0.071 pga; MDL: 0.000005–0.00111 |  | Ave:108 | <15 | GC-MS/MS | [224]
Air | 10 | 0.24–4.83 pga |  | >70 |  | GC-ECD | [225]
Air | 25 | 0.000002–0.0007 |  | Surrogate: 33.9–155 | 0.64–16 | GC-HRMS: “The oven temperature was 110 (1 min), 20 °C/min up to 210 °C， 1.5 °C/min up to 218 °C (1 min), 2/min up to 260 °C (1 min)” | [226]
Air | 29 |  |  | Samples: 3–210; method blank: 11–135; fields: 12-141 | 0–17 | GC-MS: “The programmed oven temperature went from 55 °C (hold time 2 min) to 200 °C at 70 °C min−1 (hold time 1 min), and then to 280 °C (10 °C/min) (hold time 1 min)., and then to 310 °C at 10 °C/min and finally to 325 °C at 70 °C (hold time 10 min)” | [227]
Air | 33 | 0.02–0.3 |  | 82–106 | 2–9 | GC-MS/MS: “The programmed oven temperature went from 90 °C (hold time 5 min) to 180 °C for 5 min, and then to 180 °C (5 °C/min) (hold time 3 min)., and then to 280 °C at 10 °C/min and finally to 300 °C (hold time 3 min)” | [228]
Air | 17 | 0.0001–0.055 |  | Surrogate: 118-123 |  | GC/ECNI-MS | [229]
Air | OCPs(27), CUP(3) | 2 × 10−5-3x10−4 | 5 × 10−5-1.03 × 10−3 | Surrogate: 75 and 105 |  | GC-ECD | [230]
Air | 9 | 0.71–2.13; IDL: 0.04–0.15 |  | 69–117 |  | GC-μ-ECD: “The initial oven temperature was held at 50 °C for 1 min and raised to 200 °C at 25 1 C/min, 200–300 1C at 8 °C/min” | [231]
Air | 13 | MDL: 0.007–1.27 ng/samplea |  | Int stan: 44.1–211 |  | GC-MS: “The GC temperature program started at 80 °C (1.5 min hold), then continued with 40 °C min1 to 200 °C (18 min hold) and lastly 5 min1 to 305 °C (no hold)” | [232]
Air | 100 | IDL: 0.04–4.83 pga |  | >75 |  | GC-ECD: “the oven program for OCPs was 80 °C (1 min) and then 20 °C/min to 300 °C. The injector and detector inlet temperatures for both PCBs and OCPs were 250 °C and 320 °C, respectively” | [233]
Air | 15 |  | MQL: 0.00080–0.097 | 40–110 | <10 | GC-MS: “The oven temperature program was as follows: 80 °C for 1 min, increased at 30 °C/min to 190 °C, followed by 2.5 °C/min to 230 °C, then increased at 20 °C/min to 260 °C and hold for 12 min, finally followed by 20 °C/min to 320 °C and hold for 23 min” | [234]
Air |  | 0.21 |  | 78.2–93.1 | 7.48–17.1 | GC × GC-LRMS ECD: “The initial oven temperature was maintained at 100 °C for 1 min and then 30 °C/min to 160 °C maintained for 5 min and increased to 300 °C at 1.5 °C/min and maintained for 2 min” | [235]
Air | 20 |  |  | 81.4–115.31 |  | GC-ECD: “The column temperature was increased from 80 °C to 210 °C at 10 °C/min, then increased at 0.8 °C/min to 250 °C and held for 1 min, and finally increased to 290 °C at a rate of 10 °C/min and held for 12 min” | [236]
Air suspended particles | 16 | MDL: 0.0005–0.0027 |  | 66–108 |  | GC-MS: “The oven was operated with a temperature ramp: 40 °C for 1 min, 50 °C/min up to 110 °C, 5 °C/min up to 303 °C, 20 °C/min up to 335 °C for 20 min” | [237]
Air, water | 4 | DP: 0.1–5.94; GP: 1.8–5.52 | DP: 0.11–7.15; GP: 2.31–8.31 |  |  | GC-ECD: “The programmed oven temperature went from 90 °C (hold time 1 min) to 190 °C at 20 °C min−1, and then to 310 °C (3 °C/min) (hold time 18 min)” | [100]
Air, water | 12 | Air: 1 × 10−7-1.4 × 10−6; Water: 1 × 10−7-7x10−7 |  | Surrogate: 45–64 and 72−98 |  | GC-MS/MS: “Splitless, the GC temperature programme was 80 °C (1 min hold), then 40 °C min-1 to 200 °C, and finally 5 °C min-1 to 305 °C”. | [104]
Air, water | 9 | IDL: 7.7 × 10−5-9.6 × 10−5; MDL: Air: 0.01342–0.0175; water: 1.991 × 10−5-2.649 × 10−5 |  |  |  | GC-ECD, GC-MS: “Temperature program: 60C, 1.5 min, 25C min1, 140C, 8C min1, 300C, 20 min” | [238]
Air, seawater, coral tissues | 22 | MDL: Seawater: 2.8 × 10−9-5.64 × 10−8, Air: 9 × 10−7-1.85 × 10−5, Coral tissue: 0.001–0.0358 |  | Surrogate: 63.2–110 |  | GC-MS/MS | [239]
Air, soil | 22 | IDL: 2.5 × 10−5-5.31 × 10−4 |  | 64.5–95.0;CRM: 73.1–152 |  | GC-MS: “The GC oven temperature program was as follows: 50 °C hold for 1 min; 25 °C/min to 100 °C; 5 °C/min to 260 °C (hold 1 min); 10 °C/min to 300 °C (hold 5 min)” | [240]
Gas-phase, particle phase | 22 | 0.00112–0.0076 | 0.00232–0.0054 | 64–103 |  | GC-MS: “Splitless, temperature program was 2 min at 80 °C, 10 °C/min to 285 °C (wait for 5 min), 25 °C/min to 315 °C (wait for 5 min)” | [241]
Particle-bound, sea | 20 | MDL: 1.5 × 10−6-3.3 × 10−4 |  | 76–81 |  | GC-MS: “The oven temperature was programmed as follow: 80 °C at the beginning and held for 2 min, increased to 180 °C at a rate of 20 °C per min, increased to 220 °C at a rate of 2 °C per min, increased to 245 °C at a rate of 1 °C per min, increased to 310 °C at a rate of 10 °C per min and held for 10 min” | [242]
Gas, seawater | 221 | Gaseous: 2 × 10−6–0.038; aqueous: 3 × 10−9–7.6 × 10−5 |  | Gaseous: 44–128; seawater: 42–127; median recoveries in spiked samples: 52–110 in air and 43–136 seawater |  | GC-MS/MS: “Splitless, the GC column temperature was initiated at 60 °C (held for 1 min), increased to 120 °C at 40 °C/min (held for 0 min), then to 310 °C at 5 °C/min (held for 0 min). RT: 5.6–56.5 min” | [243]
Gas/particular matter phase | 31 | MDL: 1–345 fg/m3 a |  | surrogate 81–121 | 6.5–15.1 | GC-MSD | [244]
Gas phase | 17 | MDLs: 0.0007 |  | 87–101.8 | <5 | GC-ECD | [245]
Gas-particule | 23 | MDL: 1.1–344.4 fg/m3 a |  | 83–120 | 7.4–13.7 | GC-ECNI: “oven program was set to 50 °C initial temperature, hold for 1 min, 30 °C/min to 140 °C, then 2.20 °C/min to 285 °C, and finally 15 °C/min to 300 °C, hold for 15 min (85.91 min total run time)” | [246]
Aerosol |  | 0.14–0.44 | 0.63–1.5 | 90–14 | 0.19–3.4 | GC-MS | [247]
Silicone sheet | 7 | 0.036–0.173 | 0.0008–0.078 |  |  | GC-MS/MS | [248]
plastics | 18 | 30–500 |  | Surrogate: 74.4–80 |  | GC-MS: “The GC oven temperature was raised from 70 °C (held for 2 min) to 200 °C at 20 °C/min and then increased to 250 °C at 10 °C/min” | [249]
plastics | 8 |  |  | Surrogate: 82.5–114; spike: 95-105 |  | GC-ECD: “The column oven temperature was programmed as follows: 80 °C for 2 min, increased at 8 °C per minute to a final 280 °C, which was held for 20 min. Total run time was 49 min” | [250]
Water, microplastics | 8 | 0.2–10 |  | 69–85 |  | GC-MS, GC-HRMS: “The oven temperature was initially set at 75 °C for 2 min, increased at first to 150 °C at a rate of 20 °C/min and held for 2 min, then increased to 260 °C at a rate of 3 °C/min and held for 2 min, and finally increased to 300 °C at a rate of 20 °C/min and held for 1 min” | [251]

## Table 3
S/N | LOD | MDL | LOQ | Reference
1 | 3σ |  | 10σ | [17]
2 | 3SD |  |  | [255,122,[274], [275], [276]]. [98,256]
4 | concentration of native Component in the quantification standard divided by 3S/N |  |  | [26,87],
4 | 3SD |  | 10SD | [189]
5 | mean of solvent blank + 25 ppb of surrogate standard, + 3SD |  |  | [275]
6 |  |  | 10S/N | [85]
7 | 2SD |  | 10SD | [277]
8 | S + 3.3σ |  | 3*LOD | [93]
9 |  | Sum of average of blank data and 3SD | 10SD (i.e. MQL) | [177]
10 |  | Average of all blanks + 3SD |  | [73,[94], [95], [96], [97],278]
11 | Average lab blank +3SD |  |  | [51,66,145,229,279]
12 |  | SD of the replicate analysis * one-tailed t statistics for the 99 % confidence interval |  | [203]
12 | Mean of 10 blank + 3 fold of their SD |  | Mean of 10 blank + 10 fold of their SD | [280]
13 | Lowest calibration point of an amount producing a S/N of 3.3 |  | Lowest calibration point of an amount producing a S/N of 10 | [30]
14 | Lowest calibration standard with the S/N > 3 |  | Lowest calibration standard with the S/N > 10 | [74]
15 | 3 N/S |  | 10 N/S | [19,52,88,257,149,281]
16 |  | Student's t-value of 3.14 * SD of spiked blank |  | [91,282]
17 |  | 3SD | 10SD | [181]
19 |  | Mean blank+3SD or 1/2IDL |  | [283]
20 | IDL calculated from linear extrapolation based on the lowest calibration standard | Average of field and method blank concentration+3SD |  | [232,240,241,284],
21 | Average amount of each analyte in blank +5SD+3SD |  |  | [100]
22 | 3*baseline noise in the chromatogram |  | 5*baseline noise in the chromatogram | [29] Or [285]
23 | SD of the blank matrix samples with the lowest concentrations of target standards by 7 replicates |  |  | [103]
24 |  | 3SD blank conc |  | [104]
25 | 3σb/S |  |  | [36]
26 | 3.3 SD of the smallest amount of the analyte that gave a S/N ≥ 3 |  |  | [28]
27 | 3σb/b |  | 10σb/b | [286]
28 |  |  | Mean + student t-test (3.75) SD of field blank sample | [254]
|  | IDL = 3S/N; MDL = IDL/sample's average mass |  | [89,239]
29 |  |  | 9S/N | [27]
| 3S/N |  | 9S/N | [213]
|  |  | highest detected value | [287]
30 | 3.3(S/b) |  |  | [288]
31 | 3.3σ/S |  | 10σ/S | [267,289],
32 | Sy/S | (LOD*final volume)/(sample mass*injected volume) |  | [190]
33 | Student's t-value of 3.14 * SD of spiked blank |  |  | [26,90],

## Table 4
MATRIX | No. OF COMPOUNDs | LOD | LOQ | RECOVERY | RSD | INSTRUMENT/INSTRUMENTAL CONDITIONS/TOTAL RUN TIME | REFERENCE
Vegetables | 11 | 0.07–0.18 |  | 76.21–93.51 | 4.76–11.1 | GC-ECD: “Splitless, the initial temperature at 100 °C for 1 min, ramping to 190 °C at 12 °C/min, and held for 8 min, followed by continuous ramping to 250 °C at 3 °C/min, and held for 10 min” | [290]
Vegetable | 10 | 0.0011 to 0.021 | 0.0013–0.034 | 95–110 |  | HPLC: “40 % A at 70.1–75.0 min and 7 % A at 75.1–90.1 min. The column temperature was maintained at 30 °C” | [291]
Vegetables | 20 | 0.01–0.08 | 0.03–0.24 | 88.6–102 | <6 | GC-MSD: “Splitless, Initial oven temperature was 150 °C, increase to 280 °C at 6 °C/min and final temperature was 300 °C” | [292]
Leafy vegetable(spinach, lettuce, oilseed rape, cabbage | 8 | 0.15–0.32 | 0.45–0.96 | 78.6–107.7 | 1.1–7.5 | GC-MS: “Splitless, the oven temperature program was set as follows: 50 °C as initial temperature, maintained for 1 min, raised to 200 °C at 20 °C/min, raised to 230 °C at 5 °C/min, and maintained for 5 min, raised to 280 °C at 10 °C/min, and held for 1 min. The total run time was 25.5 min” | [293]
Vegetable | 15 | 20–4500 |  | 60–120 | 0.2–19.8 | GC-ECD: “Splitless, the oven temperature was kept at 90 °C for 1.0 min and then programmed at 3.5 °C min-1 to 170 °C followed by a final ramp to 280 °C at 5.0 °C min-1” | [294]
Vegetable oils | 20 | 0.10–1.84 |  | 44–159 |  | GC-MS/MS: “Splitless, oven temperature: 50 °C for 1 min, increase at a rate of 25 °C/min up to 125 °C, then increase at a rate of 10 °C/min up to 300 °C and held for 2 min” | [295]
Vegetables and fruits | 19 | 0.019 to 0.033 (vegetable) 0.017–0.038 (fruit) | 0.048–0.081 (fruit) 0.049–0.088 (vegetable) | 87.2–99.27 (vegetables); 89.54–100.3 (fruits) |  | GC-MS/MS: “The temperature profile for GC was ranged from 80 °C to 290 °C. RT 14.23 30.54” | [48]
Fruits and vegetables | 11 | 12–987 | 40–3290 | 79.72–104.83 | 0.48–11.58 | “The oven temperature was programmed initially at 60 °C for 1 min, then raised to 140 °C at 12 °C min−1, and finally raised to 280 °C at 8 °C min−1” | [270]
Vegetables and fruits | 19 |  |  | 91.31–99.29 |  | GC-PFPD: “Split, The oven temperature was maintained initially at 60 °C for 2 min, first Ramp at 10 °C/min to 200 °C, then second Ramp at 8 °C/min to 300 °C and held constant at 5 min” | [88]
Root of ginseng | 20 | 0.18–2.0 | 0.55–6 | 51–156 | 1.8–12.6 | GC-μ-ECD, GC-EI-MS/MS | [296]
Raw and dregs Ginseng | 5 | Liquid: 100–400; soild: 0.0003–0.0012 | Liquid:20–120;soild:0.001–0.004 | Liquid: 70.3–85.6; solid: 83.4–106.9 | Liqiud: 2.7–6; soild: 4.7–14.9 | GC-MS/MS: “The column temperature was 40 °C at the start, held for 1 min; increased to 120 °C at a rate of 30 °C/min; increased to 180 °C at 20 °C/min; and increased to 280 °C at 15 °C/min, maintained for 10 min; GC-ECD: The column temperature was 120 °C at the start, held for 1 min; increased to 150 °C at a rate of 8 °C/min, held for 2 min; and then increased to 270 °C at 4 °C/min, maintained for 7 min. RT: 4.83–7.34 min”” | [297]
Fruits and vegetables | 6 | 0.03–0.30 | 0.11.0 | 73.7–111.6 | 3.4–11.9 | GC-ECD: “Splitless, initial at 70 °C, then heating at 5 °C/min to 135 °C, at 2 °C/min to 140 °C (held for 1 min), and finally ramped to 280 °C at 30 °C/min. Ultrapure nitrogen (99.999 %) was used as the carrier gas at 1.2 Ml min−1” | [281]
Fruits, vegetables | 20 | 10–100 | 100–10 × 104 | 44–101 |  | GC-ECD: “oven temperature 80 °C for 2 min ramped to 180 °C @ 20 °C/min up to 230 °C @ 5 °C/min and finally up to 280 °C @ 20 °C/min and held for 5 min” | [69]
apple, peach, Chinese cabbage and cucumber samples | 11 | 0.01–0.20 | 0.03–0.67 | 83.7–124 | 2.1–10.9 | GC-ECD: “The initial oven temperature was set at 90 °C, followed by increasing it to 184 °C at 30 °C/min, to 198 °C at 1 °C/min and to 280 °C at 20 °C/min” | [84]
Foods, fruits and vegetable | 19 |  | 0.009–0.156 | 24.3–154.7 |  | GC–NCI–MS: “The temperature was immediately increased to 240 °C at a rate of 3 °C/min, maintained for 5 min, then increased to 300 °C at 10 °C/min and maintained for 5 min” | [298]
cereals, vegetables, and fruits, | 20 |  | 0.009–0.156 | 31.5–92.3 |  | GC–NCI–MS: “The temperature was immediately increased to 240 °C at a rate of 3 °C/min, maintained for 5 min, then increased to 300 °C at 10 °C/min and maintained for 5 min” | [299]
meat products, aquatic foods, dairy products, edible oils, chicken eggs, fruits, vegetables, and cereals | 18) | 0.0001–0.0004 | 0.0003–0.001 | 83–112 | 4–18 | GC- Μecd: “The oven temperature was 75 °C (0.5 min) initially and increased to 300 °C at 10 °C/min (2 min) for a total run time of 25 min” | [300]
fruits, vegetables, cheese, yogurt, egg, meat, and fish, samples | 13 |  |  | 7–95 | 0.6–4.4 | GC-MS/MS “The GC temperature programme was 80 °C (1 min hold), then 15 °C per minute to 180 °C, and finally 5 °C per minute to 300 °C (5 min hold). RT 11.05–32.99 min” | [49]
Yam, cassava, cocoyam, sweet potato | 14 | 0.45–3.51 | 3.94–10.69 | 86.56–98.74 |  | GC-TOFMS: “Splitless, oven temperature from 70 °C (2 min hold), then raised to 130 °C at the rate of 25 °C/min, afterwards raised to 220 °C at 2 °C/min, and then raised to 280 °C at 10 °C/min, and eventually 4.6 min hold. RT: 5.58–17.5 min” | [301]
Carrot, Onion, Cabbage, Garlic and Ginger | 17 | 0.017–0.405 | 0.17–4.05 | 86.64–93.86 |  | GC-ECD “initial oven temperature of 150 °C which was later increased to 280 °C at 6 °C/min and the total run time was 21.67 min. RT: 5.117–18.875 min” | [302]
Water, fruits and vegetable | 7 | 0.0005–0.004 | 0.0015–0.015 | Water: 69.61–121.95; veg: 59.83–132.67; fruit: 67.41–115.37 | Water: <9.76; veg: 10.37; fruit: <10.03 | GC-ECD: “Split, The oven temperature program was started at 160 °C for 0.5 min, raised to 180 °C at a rate of 20 °C min–1 and held for 0.5 min, raised to 190 °C at 20 °C min–1 and held for 0.5 min, raised to 250 °C at 40 °C min–1 and held for 4 min, and then to 300 °C at 25 °C min–1 and held for 2 min” | [303]
Tobacco leaves | 20) | 2–65.5 | 6.5 to 198 | 70–100 | 2.9–64.8 | GC-ECD: “Splitless, the GC oven temperature was programmed as follows: 100C (0.5 min), 35C min1 to 220C, 10C min1 to 240C, 2C min1 to 250C, and 10C min1 to 290C (2 min). RT: 6.542–15.628” | [304]
Needle leaves | 11 | 2.3 × 10−4-0.026 |  | 52–117 |  | HRGC-HRMS: “The oven temperature was 60 °C initially, held for 1.5 min and ramped at 10 °C/min to 140 °C, then increased to 220 °C at 4 °C/min, and ramped at 2 °C/min to 250 °C, and finally ramped to 300 °C at 8 °C/min” | [305]
Pine needle | 10 | 0–3.25 pga; IDL: 0.04 pga | 0–9.087 pga | 50–120 |  | GC-MS: “initial temperature 50 °C (1min), 25 °C/min up to 200 °C, 8 °C/min up to 300 °C, 5.5 min stayed, 5 °C/min rise to 310 °C for 3 min” | [306]
Meat | 19 | 0.27–1.51 | 1.10–5.20 | 81.6–116.3 | 1.78–18.2 | GC-MS: “Split, Column 50 °C (0.4 min hold) to 195 °C at 25 °C/min; hold to 265 °C for 1.5 min at 8 °C/min; maintained at 315 °C for 1.25 min at 20 °C/min” | [307]
Walnut, soil | 21 | 0.1–0.5 | 0.1–15.76 | 80–110 |  | GC-MS/MS: “Splitless, the initial temperature of 70 °C was maintained for 2 min, and then ramped up to 150 °C at 25 °C/min, then to 200 °C at 3 °C/min, and finally to 250 °C at 8 °C/min for 10 min. RT: 9.19–29.41” | [85]
Rice, soil | 22 | 0.02–0.76 | 0.06–2.6 | Surrogate: 76- 110 |  | GC-ECD: “Oven temperature program was initially set at 80 °C and held for 2 min, then increased to 195 °C at the rate of 10 °C/min. The increased temperature ramped to 230 °C at the rate of 3 °C/min, finally ramped to 310 °C at the rate of 10 °C/min, and held for 5 min” | [308]
Therapeutic tea | 17 | 0.4–5.1 | Spiked samples: 89.9–102.2 |  | <12 | GC-ECD: “Splitless, the initial oven temperature was 150 °C and was increased to 280 °C at 6 °C/min, RT 22 min” | [309]
Tea | 20 | 0.16–2.06 | 0.54–3.84 | 80–94 | 2.43–2.93 | GC-MS: “The column temperature programs s were as the following: 70 °C (2 min), 150 °C (25 °C/min), 200 °C (3.0 °C/min), 280 °C (9.0 °C/min), isotherm (10 min), the total run time for the OCPs was 40.8. 70 °C (1.2 min) and the ramp rate was 10 °C/min to 280 °C (18 min)” | [310]
Tea | 9 | 60–720 | 30–2400 | 75.87–111.56 | 0.8–9.0 | GC-ECD: “Oven temperature program was: initial temperature of 80 °C, ramped at 30 °C/min to 180 °C, ramped at 3 °C/min to 205 °C, held for 4 min, ramped at 20 °C/min to 290 °C, held for 8 min, ramped at 50 °C/min to 325 °C. The total GC run time was 27.92 min” | [269]
Tea | 13 | 1.4–7.2 |  | 86.1–100.3 | 3.1–11 | GC-MS: “The initial temperature was 50 °C, which was increased to 180 °C at 10 °C/min and maintained for 4 min. It was then increased to 210 °C at 2 °C/min and held for 4 min. Finally, it was raised to 231 °C at 3 °C/min and held for 5 min” | [311]
Corn, Corn floor, barn | 26 | MDL:C: 0.01–3.45; cf: 0.01–2.83; B: 0.01–1.98 |  | 48-129; 37–159; 35-135 |  | GC-MS/MS: “The oven temperature was programmed as follows: the initial temperature (50 °C) was held for 1 min and increased by 25 °C/min to 125 °C and then increased by 10 °C/min to 300 °C and held for 2 min” | [122]
Maize, flour | 19 | 14.23–32.31 | 63.68–100.98 | 66.74–100.65 |  | GC-MS: “The initial temperature was 90 °C for 2 min and later increased to 260 °C at 5 °C/min and held for 5 min” | [93]
maize | 17 | 1–4 | 3–12 | 89.6–100.91 | 0.69–12 | GC-ECD: “Oven temperature started at 150 °C and increased to 280 °C at 6 °C per minute. The injection was through a splitless injector, using helium as a carrier gas at a flow rate of 2 mL/min. The run time was 21.67 min” | [50]
Maize, sediment, water | 23 | Soil & maize: 0.00117–0.470; water: 2.4 × 10−7–7.83 × 10−5 |  | Surrogate: 65.7–101 |  | GC-MS/MS: “The GC oven temperature was set at 100 °C for 0.5 min, then 20 °C/min to 160 °C, 4 °C/min to 290 °C, and finally 10 °C/min to 300 °C, and hold for 10 min” | [266]
Water, plant and substrate | 9 | MDL: Water: 0.005 to 0.01; Plant: 0.01–0.05; substrate: 0.05–0.1 |  | Spiked: water: 72.4–116.5; plant: 69.9–109.2; substrate: 72.4–116.5 |  | GC-MS: “The column temperature was initiated at 80 °C (kept for 1 min), increased to 230 °C at 10 °C⋅min− 1 (kept for 4 min)” | [312]
Water, plant, soil | 35 | Water: 0.0025–0.35; plant: 0.05–7.0; soil: 0.05–7.0 |  |  |  | GC–MS/MS and UPLC-MS/MS: “Splitless, an initial temperature of 75 °C and a maximal temperature of 300 °C at the end of the injection's transfer phase (rate: 10 °C/s in 2.5 min)” | [313]
Plant tissue, soil | 6 | Plant: 25; soil: 100 | Plant: 0.1 |  |  | GC-ECD: “The temperature of the thermostat column was programmed from 40 °C (hold time 1 min) to 160 °C (hold time 3 min) at a heating rate of 20 °C/min, followed by heating to 250 °C (hold time 5 min) at 3 °C/min” | [314]
Plant | 11 |  | 3.2–9.9 | 82.3–91.9 | 6–12 | GC-ECD: “The initial column oven temperature was 150 °C; it was then ramped up at 2.5 °C per min to 270 °C and kept for 15min” | [315]
Plant | 18 | 0.001–0.004 | 0.004–0.011 | 81–96 | 0.1–1.5 | GC-qMS: “The oven temperature was programmed from 80 °C (initial time, 2 min) to 205 °C at a rate of 30 °C/min (hold time 5 min) and then heated to 290 °C at a rate of 10 °C/min with a final holding time of 3 min” | [316]
Dietary supplement extract | 8 | 0.00899–0.0931 |  |  | 4.48–12.9 | GC-MS: “The GC oven temperature was ramped at 10 °C/min from 40 °C to 290 °C, and then held at 290 °C” | [317]
Kaht plant | 15 | DDT: 0.1; HCH: 0.15–0.45 | HCH: 0.3; HCH: 0.5–1.5 | 86.72–114.75 (DDX),; 97.33–111.12 (HCH) | 2.84–11.7 | GC-MS: “The initial oven temperature was set at 90 °C and held for 3 min. The temperature was ramped to 150 °C at a rate of 15 °C/min. Then, it was ramped to 280 °C at a rate of 5 °C/min and held for 3 min” | [265]
Tree bark |  | 3.85 × 10−5–0.0131 |  | 51–98 |  |  | [318]
Seaweed | 20 | 0.001–0.004 | 0.005–0.017 | 72–120 | <12 | HPLC: “The column temperature was maintained at 35 °C″ | [19]
seaweeds | 17 | 1 |  | 80–108 |  | GC-MS: “initial oven temperature at 90 °C for 0.5 min, which was then increased to 280 °C at 8 °C min-1 and again to 300 °C at 15 °C min-1 for 2.5 min. RT: 13.65–22.74 min” | [319]
wheat | 15 | 0.02–0.03 | 0.010–0.05 | 80–92 | 2–8 | GC-MS: “initial 30 °C (2 min hold), ramped to 180 °C at 15 °C/min, ramped to 190 °C at 2 °C/min, ramped to 290 °C at 2.5 °C/min, and ramped to 320 °C at 5 °C/min (15 min hold)” | [320]
Olive oil | 10 | 0–3.25 Pga IDL: 0.04 Pga | 0–9.087 Pga | 83–117 |  | GC-MS: “The oven temperature analysis was as follows: the furnace was kept at 80 °C for 1 min and then, increased to 300 °C with 20 °C/min” | [321]
Tap water, river water, palm oil mill effluent | Endosulfsn, dieldrin | 0.0073, 0.0086 | 0.022, 0.025 | 98.6-1-3.5 | 4.61–6.79 | GC-ECD: “The oven was temperature-programed from 120 °C to 190 °C at 40 °C/min, then to 285 °C at 30 °C/min” | [322]
Honey | 2 | 1.0 and 2.0 | 2.0 and 4.0 | 84.1 and 111.7 | 1 and 8 | GC-MS: “The GC oven temperature was programmed at 100 °C, with a heating rate of 20 °C/min to 200 °C (1 min) and with a heating rate of 10 °C/min to 280 °C (1 min), in a total analysis time of 15 min” | [323]
Honey, agricultural soil | Pesticides(10) | Soil: 0.04–0.2; honey: 0.05–0.20 |  | Soil: (64–90); honey: (73–99) |  | GC-ECD: “Initially, the column temperature was set at 90 °C for 1 min and then increased to 180 °C at 30 °C/min. It was then further raised to 260 °C at 4 °C/min and then kept at this temperature for 16 min” | [262]
Honeybee, pollen, honey | 10 |  |  |  | 0.021–0.338. | GC-μECD: “initial furnace temperature is 80 °C (1 min), then with increases of 20 °C min-1 to 240 °C and 5 min wait at this temperature, then with increases of 5 °C min-1 to 270 °C and increases of 20 °C min-1 to 300 °C and kept at this temperature for 3.5 min” | [279]
Honey and propolis | 13 |  | Propolis: 0.49–1.11; Honey: 0.54–1.02 | Propolis: 85.9–105.6; Honey: 89.9–106.3 | Propolis: <14.4; Honey: <12.9 | GC-MS: “Splitless, The GC oven was initially maintained for 2 min at 70 °C, and then the temperature was increased to 150 °C at 25 °C/min. Then, it was increased to 200 °C at 5 °C/min (held for 5 min), increased to 270 °C at 5 °C (held for 2 min), and finally increased to 290 °C at 25 °C/min (held for 5 min)” | [324]
Propolis | 13 | 0.16–0.37 | 0.49–1.11 | 85.9–107.2 | <11.5 | GC-MS: “The temperature program began at 70 °C (2 min), raised to 150 °C at 25 °C/min, ramped to 200 °C for 5 min at 5 °C/min, was further increased to 270 °C (2 min) at 5 °C/min, and lastly 290 °C for 5 min at 5 °C/min” | [325]KKUU
Honey | 14 | 0.01–0.04 | 0.04–0.12 | 81.4–111.4 | 2.0–11.2 | GC-MS/MS-TQ: “Splitless, initial temperature 70 °C (held for 1 min), raised to 160 °C at 10 °C/min, then raised to 240 °C at 2 °C/min, and finally increased to 280 °C at 20 °C/min (held for 6 min)” | [326]
Honey bees, bee bread, honey (HBH) | 18 | IDL: 0.015–0.33; MDL: (HBH): 0.015–0.165; soil&flower: 0.006–0.066 |  | >90 |  | GC-ECD: “The oven temperature program was: 100 °C held for 1 min, followed by an increase of 5 °C/min to 150 °C, held for 1 min, increase 1.5 °C/min to 240 °C, and then 10 °C/min to 300 °C, held for 10 min” | [327]
Low density polyethylene sheet | 22 | 0.00–0.31 |  | 57%-104 |  | GS Micromass MS-MS: “Samples were held at 140 °C for 2 min, ramped up to 180 °C at 10 °C/min, ramped up to 220 °C at 3 °C/min, and ramped up to 315 °C at 10 °C/min and held for 10 min” | [328]

## Table 5
MATRIX | No. ofCOMPOUNDs | LOD | LOQ | RECOVERY | RSD | INSTRUMENT/INSTRUMENTAL CONDITIONS/TOTAL RUN TIME | REFERENCE
Fish | 11 | 1.0–3.0 | 3.3–10.0 | 82.11–115.7 | Single fibre 4.1–7.6Fibre to fibre:8.0–12.0 | GC-ECD: “Split mode at a split ratio of 10, The oven temperature program started at 60 °C for 1 min, followed by being increased at 40 °C min− 1 to 172 °C, held at 172 °C for 2 min, increased at a rate of 1 °C min− 1 to 195 °C, raised at 30 °C min− 1 to 280 °C and then kept at 280 °C for 2 min” | [13]
Fish | 10 | 0.1–0.6 |  | Spiked sample: 88.4 to 98.5 |  | GC-MS: “The oven temperature was programmed as follows: 100 °C for 4 min, 7 °C/min to 310 °C, and 6 min hold at 310 °C. A 2 Μl sample was injected in splitless mode with the split outlet opened after 1.0 min” | [274]
Fish | 16 |  | 0.02–0.05 | 75–105 |  | GC-ECD: “the initial oven temperature of 180 °C was increased by 7.5 °C min− 1 to 220 °C and held for 5 min, then heated to 250 °C at 20 °C min− 1, and held for 10 min” | [329]
Fish | 9 | 0.3 |  | 64–100 |  | GC-MS: “The column temperature was initially set at 70 °C. This temperature was maintained for 3 min, and increased to 150 °C at a rate of 25 °C/min, then increased to 200 °C at a rate of 3 °C/min and then to 280 °C at a rate of 20 °C/min and maintained for 6 min, then increased to 325 °C at a rate of 50 °C/min and held for 1 min” | [330]
fish | 9 | 3 × 10−5-5X10−4 | 0.09–1.8 | 64–110 |  | GC-MS: “Splitless. the oven temperature program was as follows: from 110 to 170 °C at 1.5 °C/min (held for 5 min), from 170 to 226 °C at 2 °C/min (held for 5 min), from 226 to 280 °C at 40 °C/min, and finally kept at 280 °C for 10 min” | [331]
Fish | 15 |  | 0.1–0.7 | 75–92 |  | GC-MS | [332]
Fish | 20 | 0–0.008 |  | 0.024–0.833 rangea |  | GC-ECD: “oven heating ramp to 120 °C (for 1 min) until 240 °C increasing at a rate of 4 °C min−1; EDC at 300 °C; injector at 260 °C″ | [333]
Fish(krill) | 20 | 0.11–1.4 |  | Surrogate: 80.3–137.1 |  | GC-QqQ-MS: “initial temperature at 70 °C (held for 1 min), raised to 160 °C at 10 °C/min, raised to 280 °C at 5 °C/min and held for 5 min. Then it was raised to 300 °C at 20 °C/min and held for 5 min” | [282]
Fish | 8 | 0.01–0.43 | 0.02–1.3 | 72–115 | <25 | GC-MS: “temperature program 80 °C for 2.5 min; then 20 °C/min ramp to 180 °C followed by 5 °C/min ramp to 230 °C and 35 °C ramp to 300 °C (held for 7 min). Total run time was 26.5 min” | [52]
Fish |  | MDL: 0.001 |  |  |  | GC-ECD: “oven temperature: 260 °C starting from 0 to 180 °C for 0.3 min and continued at 5 °C/min to 220 °C, held for 12 min, and continued at 5 °C/min to 260 °C″ | [334]
Fish | 10 | 0.12 | 0.40 | internal standards:78-95 |  | GC-ECD: “The furnace temperature was programmed as follows: 70 °C held for 2 min, ramp at 25 °C/min to 180 °C, held for 1 min, and finally, ramp at 5 °C/min to 300 °C. RT 31.4” | [277]
Fish | 9 |  |  | Ave: 124 |  | GC-MS: “the oven temperature program started at 40 °C which lasted for 2 min, then was increased in turn at 10, 5, and 10 °C/min to 100, 260 and 300 °C respectively, and kept for 10 min” | [335]
Fish | 18 | 0.06–0.45 | 0.25–1.45 | 88.7–99.3 | 3.9–9.4 | GC-MS RT: “10.145–20.941 min” | [336]
Fish | 8 | 0.042–0.179 | 0,14–0.60 | 75.6–124 | 0.61–3.96 | GC-MS: “the oven program was set at 110 °C for 1 min, increased to 220 °C at 10 °C/min, to 300 °C at 5 °C/min where it was held for 10 min” | [337]
fish | 14 | 2–6 |  | 75–144 | <20 | GC-MS/MS: “Splitless, the oven temperature program was as follows: 70 °C for 3 min, up to 150 °C at a rate of 50 °C/min, up to 200 °C at a rate of 3 °C/min maintained for 1 min, up to 280 °C at a rate of 20 °C/min maintained for 5 min and finally, up to 310 °C at a rate of 40 °C/min maintained for 4 min” | [287]
Lipids for Fish muscle | 7 | 0.01 |  | 59–88 | 4–15 | GC-FID | [338]
Fish muscle | 7 | 0.01 |  | 79–89 |  | GC-ECD | [339]
Fish muscle | 9 | 0.1–0.7 |  | 92.6–101.4 |  | GC-MS: “The oven temperature was programmed as follows: 100 °C for 4 min, 7 °C/min to 310 °C, and 6 min hold at 310 °C” | [276]
Fish muscle | 9 |  |  | 56–78 | 7–13 | GC-MS, GC-ECD | [285]
Shellfish | 14 |  | 5 | 95.11–102.17 | 3.57–10.6 | GC-MS/MS: “Splitless, it was initially isothermal for 1 min, raised to 150 °C at 40 °C/min, and finally maintained at 300 °C for 8 min. RT 10.31–13.45” | [340]
shellfish | 5 | 0.003–2.705 | 0.01–9.02 | 70–120 | <10 | GC-MS/MS: “The oven temperature was programmed at 80 °C for 1 min, increased to 150 °C at a rate of 20 °C/min, followed by 5 °C/min ramp to 300 °C (held for 5 min) for the total run time of 39.5 min” | [341]
Shellfish, cephalopods | 10 | Shellfish: 0.10–0.80; cephalopods: 0.21–0.77 | Shellfish: 0.31–2.41; cephalopods: 0.63–2.33 | Shellfish: 83.5–117.4; cephalopod: 79.8–118.4 | Shellfish: 0.3–27.5; cephalopod: 1.2–27.9 | GC-MS/MS: “The oven temperature program was held at 60 °C for 2 min, initially. Then, the temperature was elevated to 165 °C at a rate of 30 °C/min, then increased to 195 °C at a rate of 15 °C/min (held for 1 min), increased from 195 to 210 °C at a rate of 2 °C/min, then elevated to 220 °C at a rate of 5 °C/min and then programmed to 300 °C at 10 °C/min, held for 1.5 min” | [342]
Swordfish | 7 | Muscle: 4 × 10−5-0.02; liver: 2 × 10−5-0.014; gonad: 1 × 10−5- 0.016 |  | 49–8 | 6–28 | GC-HRMS, UPLC-MS/MS | [343]
Lipids of fish | 21 |  | 0.1–0.5 | 65–114 | 3–19 | GC-MS/MS: “50 °C (2.9 min); 30 °C.min−1 to 240 °C; 2 °C.min−1 to 270 °C; and 40 °C.min−1 to 340 °C (12 min)” | [74]
Lipids from fish | 9 | Ave: 0.01 | 0.01–0.06 | Spiked standards:69.8–98.9; Internal standard: 84.1–98.3. |  | GC-MS: “oven temperature: start from 90 °C (0.5 min), increase 7 °C/min; 220 °C (12 min), increase 6 °C/min; 285 °C (7 min), increase 5 °C/min; and 295 °C (6 min) (post run)” | [344]
fish, crab, shrimp, shellfish, and turtle, sediment | 5 | 0.3–1.5 |  | 75–11 | 6.6–17 | GC-MS: “Column initial temperature 50 °C, hold for 3 min; 35 °C·min−1 rise to 220 °C, hold for 1 min; 15 °C· min−1 rise to 300 °C, maintain for 2 min” | [345]
Fish and crustacean | 17 | 0.9–166 |  | 47–114 |  | GC-HRMS | [26]
Water, sediment, Muscle of fish, crustacean, oft tissue of shellfish | 8 | Water: 0.250–2.240; 0.003–0.018; organism: 0.003–0.040 |  | 69–103 | <12 | GC-ECD: “Splitless, the temperature program initially commenced at 80 °C and increased by 20 °C/min up to 200 °C. It then increased to 250 °C at 4 °C/min and was maintained for 2 min; then, it was maintained at 280 °C for 5 min after increasing by 30 °C/min” | [346]
Tilapia muscle | 24 | 0.006–0.087 | 0.020–0.289 | Standard: 68.6–96.2; surrogate: matrix: 73.8–105.6 | <17 | GC-MS/MS: “The GC oven temperature was programmed as follows: 80 °C (5 min)→20 °C/min→160 °C (0 min)→4 °C/min→240 °C (0 min)→10 °C/min→295 °C (2 min)” | [347]
Lipids of pacific salmon fish(muscles, egg, liver, male gonad) | 10 | 0.1–0.6 |  | Spiked samples: 92.6–101.4 |  | GC-MS: “Splitless, the oven temperature was programmed as follows: 100 °C for 4 min, 7 °C/min to 310 °C, and 6 min hold at 310 °C″ | [98]
Muscles, gonad and liver of cockfish | 12 | 0.005–0.547 | 0.017–1.825 | Matrix: 78&85; surrogate: 90&110 |  | GC-ECD: “Splitless, the oven temperature program was set with a 100 °C start, held for 1 min, followed by an increase of 5 °C/min up to 150 °C, held for 1 min, then 1.5 °C/min up to 240 °C, and then 10 °C/min up to 300 °C for 10 min” | [348]
Fish gonad | 16 |  |  | 89–124 | 2–7 | GC-MS: “The oven temperature was programmed to increase from 90 °C with 2 min hold to 180 °C with 2 min hold (25 °C/min), 220 °C with 2 min hold (1.5 °C/min), 275 °C with 2 min hold (3 °C/min), and final 300 °C with 4 min hold (25 °C/min)” | [349]
Muscle, liver, egg, male gonad of salmon | 12 | 0.1–0.5 |  | 94.6–103.7 |  | GC-MS | [350]
muscle, lung, liver, kidney, and blubber samples of dolphins | 19 | 0.3–0.7 |  | Internal standard: 107.3 | <12 | GC-MS: “The column oven temperature was programmed as follows: 70 °C held for 2 min, increased at 3 °C per minute to 270 °C and held for 5 min, and increased at 5 °C per minute to the final temperature of 300 °C and held for 10 min” | [351]
Brain &testes of dolphin | 1 | 0.3–0.7 |  | Spiked: 86.6–123.1; Surrogate: 71.63–128.9 | <12 | GC-MS: “The column oven temperature: 70 °C held for 2 min, increased at 3 °C per min to 270 °C and held for 5 min, and increased at 5 °C per min to a final temperature of 300 °C and held for 10 min” | [352]
Dolphin | 19 | Reporting limit: 0.11–25 |  | 50–150 |  | :GC-MSD “the oven temperature was programmed from 90 °C (1 min hold) to 150 °C at 5 °C/min, to 260 °C at 3 °C/min, and to 320 °C at 20 °C/min (5 min hold)” | [284]
Liver and muscles of fish | 29 | Liver: 0.23–8.62; muscle: 0.09–3.45 |  | Liver: 63& 86; muscle: 59&93 |  | GC-MS: “The oven was programmed as follows: 60 °C for 1 min, then 40 °C/min until 120 °C, and finally 5 °C/min until 285 °C for 0 min (Total time: 35.5 min)” | [255]
Muscle and liver of fish | 10 | 0.1–0.7 | 0.3–2.1 | 73–98.05 | 0.35–5.66 | GC-MS: “The oven temperature was optimized and set as follows: 100 °C hold for 2 min, raised to 280 °C at 10 °C/min, and then hold for 45 min” | [353]
Liver, muscle of fish | 8 |  | Liver: 0.01; 0.03 | 90–99 |  | GC-ECD, GC-MS/MS: “The temperature program of the column oven was initially 90 °C (3 min) to 215 °C (40 min) at 30 °C/min−1 and 275 °C (30 min) at 5 °C/min−1; the injector temperature was 270 °C and the carrier gas was helium (1 mL/min)” | [354]
liver, gills, and muscles of fish | 18 | 0.01–0.02 | 0.04–0.08 | 90.2–103 | 1.2–6.1 | GC-ECD; GS-MS: ECD:“initial column temperature was 80 °C for 6 min, increased to 215 °C at a rate of 15 °C/min (hold for 1 min), then to 230 °C at 5 °C/min, and finally to 290 °C at 5 °C/min (hold for 2 min); MS: The separation temperature program was initially set at 85 °C for 0.3 min, increased to 150 °C (hold for 4 min) at a rate of 30 °C/min, then to 185 °C at a rate of 2 °C/min, and finally to 290 °C (hold for 5 min) at a rate of 4 °C/min” | [355]
Muscles of fish and prawns | 16 | 0.001–0.017 |  | 86–135 |  | GC-MS | [356]
“ater, sediment, lipids from fish | 18 | Water: 4 × 10−5-4.56 × 10−3; sediment: 0.001–0.124; fish: 0.005–0.547 | 1.4 × 10−4-0.01522, sediment: 0.004–415 | Matrix 78 and 85; spiked surrogate: 90 and 110; fish: 0.017–1.825 |  | GC-ECD: “The oven temperature program started at 100 °C, held for 1 min, followed by an increase of 5 °C/min up to 150 °C, held for 1 min, followed by an increase of 1.5 °C/min up to 240 °C, and then by an increase of 10 °C/min up to 300 °C, and held for 10 min” | [283]
Lipid of fish and sediment | 21 | Lipids: 0.7–1.7; sediment: 0.42–3.2 |  | Sediment: 68–115; fish: 74–109 |  | GC X GC-TOFMS | [357]
shrimp | 17 |  |  | 79.7–94.1 | 7.65–16.58 | GC-MS: “The oven temperature was kept at 90 °C for 30s, then increased to 280 °C @ 8 °C min− 1 and then to 300 °C @ 15 °C min− 1 for 2.5 min. The split and carrier gas flow were 50 and 1 mL/min, respectively” | [358]
Zooplanktons, fishes, shrimps, water | 16 | 0.1 and 0.5 |  | 72–108 | 7–15 | GC-MS: “The GC oven temperature was programmed to hold at 80 °C for 1 min and raise the temperature to 180 °C at a rate of 15 °C/min. Once the GC oven attained 180 °C, the temperature was increased to 205 °C at a rate of 3 °C/min and held for 5 min. At a rate of 40 °C increase min−1, the oven temperature finally increased to 265 °C and was held for 20 min” | [359]
Water, shrimps | 14 | MDL: Water: 0.006–0.013; shrimps: 0.95–2.89 | MQL: Water: 0.020–0.044; shrimps: 3.16–9.63 | Water: 60–88; shrimps: 56-96 | 2.9–13.9 | GC-ECD/GC-MS/MS: RT: 16.41–25.19 min | [360]
Blubber (Seal) |  | 0–0.04 |  | 85.5–99 |  | GC-ECD | [361]
lichen |  | 4.36 |  | 56–91 |  | GC-ECNIMS | [362]
Marine mammals | 19 | 0.3–1.2 |  | Surrogate: 71–96; matrix: 93-103 | <20 | GC-MS/MS: “Splitless, initial oven temperature of 100 °C (1 min), ramping at a rate of 11.0 °C/min to 180 °C, then to 260 °C (0.00 min) at 3.0 °C/min, and finally at 20.0 °C/min to a final temperature of 300 °C with a final holding time of 6.00 min” | [363]
Plasma of blue-footed booby | 19 | 7.5&7.8 | 22.5&23.3 |  |  | GC-ECD: “The initial oven temperature was 150 °C. Two temperature gradients were programmed: the first from 5 to 245 °C/min and the second from 10 to 310 °C/min for 5 min” | [364]
| 22 |  |  | Surrogate: 66&72 |  | GC-MS/MS: “The GC oven temperature was set at 100 °C for 0.5 min, then 20 °C/min to 160 °C, 4 °C/min to 290 °C, and finally 10 °C/min to 300 °C, and hold for 10 min” | [365]
shark | 10 |  | 0.0001–0.0274 | 42–115 |  | GC-HRMS | [257]
Plasma of sea turtles | 15 |  | 0.0050 | Spiked blank: 65.6–117; matrice: 62.5–120; internal standard: 53.6–90.8 |  | GC-MS/MS: “The oven temperature program was 50 °C for 1 min, followed by an increase at 20 °C/min to 200 °C and an increase at a rate of 10 °C/min until it reached 300 °C, remaining constant for 5 min” | [366]
Green sea turtle | 20 |  |  | Liver: 81–108; muscle: 56-82 |  | GC-MS: “Splitless, GC column oven temperature was initially at 60 °C held for 1 min, then increased at a rate of 25 °C/min to 160.0 °C. The temperature was then increased to 240.0 °C at a rate of 4.0 °C/min and then finally to 290.0 °C at a rate of 10.0 °C/min where the temperature was held for 11 min” | [367]
Blood plasma of green sea turtle | 11 | 0.18 |  | 100.86–119.14 |  | GC-MS: “The oven temperature profile was programmed from 45 °C to 150 °C at 20 °C/min, and then to 300 °C at 2.5 °C/min” | [267]
Feces, commercial feed, soil of primates | OCPS(21), CUP(29) |  |  | CUP: 63-130 | OCP: 28–54; OCP: 10-22 | OCP: GC-EC:D “The GC oven temperature program was as follows: initial 100 32 °C for 1 min, 1 °C/min to 240 °C, 10 °C/min to 280 °C, and held for 20 min, 80 °C for 0.1 min, 500 °C/min to 280 °C, and held for a final 20 min”. CUP: GC-MS “The GC oven temperature program was as follows: initial 80 °C for 2 min, 30 °C/min to 150 °C, 2 °C/min to 240 °C, 20 °C/min to 300 °C, and held for 5 min. The MS transfer line was held at 280 °C, and the ion source and quadrupole temperatures were both held at 150 °C″ | [368]
Crab | 10 | 0.03–1.08 | 0.1–3.6 | 80.2–119.8 | 1.4–12.9 | GC-MS: “The oven conditions were set as follows: 80 °C hold for 1 min, 80–170 °C at 30 °C/min, then 170–240 °C at 5 °C/min, hold for 1 min, finally 240–300 °C at 20 °C/min” | [369]
Mussels | 4 | 0.01 |  | 60115 | 3–20 | GC-ECD; GC-MS: “the column temperature programme was 90 °C (3 min) to 215 °C (40 min) at 30 °C/min and to 275 °C (30 min) at 5 °C/min” | [370]
Mussels | 14 | 0.02–0.11 | 0.16–0.8 | 61–120 | 0.7–7.0 | GC-MS: “Splitless, the temperature program was as follows: 1 min isothermal program at 90 °C, increased to 180 °C at 30 °C/min, to 300 °C at 5 °C/min and held for 1 min” | [371]
Mussels |  | 0.5 | <500 | 80–120 |  | GC-MS/MS: “GC oven was temperature programmed from 80 °C (held for 2 min) to 300 °C at 12 °C.min−1 and then held for 5 min” | [372]
mussel | 23 | 0.01–0.058 | 0.028–0.93 | 99.6–106 | 1.14–6.60 | GC-MS/MS: “The oven temperature programme began at 70 °C, increased to 240 °C at 20 °C/min, and subsequently increased to the final temperature of 310 °C at 8 °C/min (held for 11 min)” | [373]
Mussel |  | 0.05–0.09 |  | Surrogte: 83.3, 80.5 |  | GC-ECD: “70 °C/2 min to 260 °C at 3 °C/min, and then held for 25 min” | [374]
Frog | 22 |  | 0.13–4 | SRM: 75–110: | <12 | GC-ECD: “Splitless, the oven program was initialised at 100 °C held for 1 min, ramped at 12 °C/min to 180 °C, ramped at 4 °C/min to 240 °C, ramped 10 °C/min to 270 °C and held for 5 min” | [24]
Bivalves | 10 | 0.1–0.6 |  | 92.6 to 101.4 |  | GC-MS/GC-ECD: “The oven temperature was programmed as follows: 100 °C for 4 min, 7 °C/min to 310 °C, and 6 min hold at 310 °C” | [31]
Bird tissue | 9 | 0.3 | 0.9 | 74–103 | <14 | GC-MS: “The column temperature was initially set at 70 °C. This temperature was maintained for 3 min, and increased to 150 °C at a rate of 25 °C/min, then increased to 200 °C at a rate of 3 °C/min and then to 280 °C at a rate of 20 °C/min and maintained for 6 min, then increased to 325 °C at a rate of 50 °C/min and held for 1 min” | [330]
Muscle, whole of seabirds | 20 | IDL: 0.055–0.903 | 0.0–0.772 | 50–120 | <15 | GC-HRMS | [254]
Tissue, liver, brain of bird | 13 | 1 |  | 91–102 |  | GC-ECD: “oven temperature was programmed as 180 °C 3 min; 4 °C/min 260 °C 15 min” | [375]
Feathers of bird | 8 |  | 0.02 | 90–100 |  | GC-ECD: “2 min at 60 °C, gradual heating from 60 to 160 °C at the rate of 20 °C/min; 3 min at 160 °C, gradual heating from 160 to 280 °C at the rate of 2.5 °C/min; 10 min at 280 °C, gradual heating from 280 to 300 °C at the rate of 20 °C/min; and 10 min at 300 °C” | [256]
Feathers of bird | 23 | 0.027 | 0.05 | 97.5 | 4.27 | GC-MS | [376]
Feathers of pigeon | 25 | 0.006–1.3 | 0.3–3 | 61–120 | 1–18 | GC-QqQ-MS | [377]
Feathers of penguins | 6 | 0.011–2.67 |  | Internal std: 69 |  | GC-MS/MS: “Splitless, the column temperature was initially held at 70C for 2 min, raised to 150C at the rate of 25C/min, then to 200C at the rate of 3C/min, and to 280 at a rate of 8C, held for 10 min, finally to 300C at a rate of 100C/min, and held at the final temperature for 5 min. Total analysis time was 47.067 min” | [288]
Penguin tissues | 31 |  | 2.2–4.6 | Sur: 53–136; spike: 113; SRM: 130 |  | GC-ECD: “The column oven was programmed for an initial temperature of 100 °C for 1 min and a rate of 5 °C/min to 140 °C, then held for 1 min at 140 °C, ramped to 250 °C at a rate of 1.5 °C/min, held for 1 min and finally increased at a rate of 10 °C/min to 300 °C and held for 5 min” | [378]
Blood of penguins, short-tailedshearwaters | 14 | MDL: 0.04–0.95 | MQL: 0.12–3.2 | 82–115 |  | GC-MS/MS: “The initial inlet temperature was 90 °C for 0.1 min before ramping to 325 °C at 900 °C/min. The oven temperature was 50 °C for 1 min and then ramped to 320 °C at 25 °C/min with a hold time of 4 min” | [22]
Eagle-owl feathers | 12 |  | 0.05–0.10 | Internal: 71-92 | <30 | GC-MS: “Splitless, the temperature of the DB-5 column was programmed from 90 °C, kept for 1.5 min, then increased with 15 °C/min to 310 °C, kept for 15 min” | [73]
Feathers, liver of owl | 16 | 0.03–0.54 |  | 46-146 (feathers) 86 146 (liver) |  | GC-MS | [379]
Blood cell and feather of Antarctic petrels | 5 | 0.00172–0.305; | 0.00517–1.016 | 45–104 |  | GC-HRMS: “Splitless, the PTV injector was held at 90 °C for 0.1 min, ramped to 320 °C at 5 °C/min with a hold time of 5 min” | [380]
Pellets of black vulture |  | 4–310 |  | >89 |  | GC-MS/MS: “The GC oven temperature was programmed as follows: initial temperature of 70 °C, hold at 70 °C for 2 min, increase the temperature to 300 °C at a rate of 20 °C/min, and hold at 300 °C for 8 min” | [381]
Hamster head | 9 | 0.084–0.725 |  | 87.2–96.1; surrogate: 76.5&121.3 |  | GC-MS: “started with the temperature of 100 °C, held it for 1 min, then increased 43 at 10 °C/min to 220 °C and 20 °C/min to 280 °C, held it for 10 min” | [382]
Liver of wild boar | 4 | 0.008–0.05 | 0.02–1.1 | 86–120 | 3–18 | GC-QqQ-MS/MS | [383]
Water, sediment and waterbird lung tissue | 18 | 0.00001–0.0001 (water), 0.01–0.05 (sediment) (plant) 0.05–0.64 bird. |  | Spiked: 74–102; surrogate: 70&82 | 3.6–11.4 | GC-ECD: “The oven was programmed for an initial temperature of 100 °C, increased to 200 °C at 4 °C/min, then increased to 230 °C at 2 °C/min, and raised to 280 °C at 8 °C/min” | [77]
Bird's liver and lung | 10 | Lung:2.59, Liver: 0.689; fat tissue: 0.309 |  | 50-150; (Matrix spike): 78.1–100; 59.0 (surrogate_ | <20 | GC-MS: “Splitless, oven temperature was set at 80 °C for 1 min, ramped to 240 °C at 20 °C/min, kept at 240 °C for 6 min, then heated to 280 °C 10 °C/min, and held at this temperature for 10 min.RT 11.7O-13.55” | [92]
Bird's egg | 10 | 0.5 |  | 94–105 |  | GC-ECD: “oven temperature was programmed as 180 °C-3 min; 4 °C/min-260 °C-15 min” | [384]
Chicken Brain | 9 | 0,01–0.40 |  | 96–125 |  | GC-ECD and LRMS | [385]
Lipid material, homogenates of muscle, liver and gonad | 19 | 0.005–0.547 | 0.017–1.825 | >90 |  | GC-ECD: “The oven temperature program started at 100 °C, held for 1 min, followed by an increase of 5 °C/min up to 150 °C, held for 1 min, followed by an increase of 1.5 °C/min up to 240 °C, and then by an increase of 10 °C/min up to 300 °C, and held for 10 min” | [386]
Muscle of crocodile | 6 | 0.04–0.16 |  | 75–110 |  | GC-ECD: “The GC oven temperatures were programmed as follows: 100 °C for a 1 min hold; then temperature was ramped at 12 °C.min−1 to 180 °C; followed by an increase of 4 °C.min−1 to 240 °C and then by 10 °C.min−1 to 270 °C″ | [387]
Serum of leopard | 22 | 800–3.4 × 104 |  | 78.8 | <11 | GC-ECD: “Splitless, the oven program initiated at 100 °C held for 1 min, followed by a ramp of 20 °C/min to 200 °C, then changing to a ramp of 6 °C/min until 260 °C held for 4 min” | [25]
Cow and human milk | 50 | 0.15–0.9 | 0.15–3 | 74–121 | 1–18 | UHPLC-MS, GC-ECD: primary temperature, 100 °C held for 1min, then increased with the rate of 3 °C/min to 230 °C. Total run time: 44.33 min. | [388]
Buffaloes and cow's milk | 18 | 0.12–1.54 | 0.42–5.14 | Surrogate: 75&84; spiked: 83–130 | <20 | GC-MS: “for the first 3 min the temperature was 150 °C, 4 °C/min to 290 °C, and the isothermal process was kept for 10min” | [389]
Donkey milk | 7 | 0.01–0.06 | 0.05–0.3 | 81–105 % | <15 % | “The oven temperature was programmed as follows: 110 °C for 1 min, 8 °C/min to 230 °C, hold for 8 min at 230 °C, then 3 °C/min to 300 °C and hold for 6 min at 300 °C″ | [390]
Breastmilk | 8 | 180–780 | 62–2380 | 52.5–117.5 | 1.61–12.81 | GC-ECD: “The initial temperature was 80 °C, ramp at 30 °C min-1 to 180 °C, ramp at 3 °C min-1 to 205 °C, held for 4 min, ramp at 20 °C min-1 to 290 °C, held for 8 min, ramp at 50 °C min-1 to 325 °C. Deltamethrin: the oven temperature was maintained initially at 130°c, held for 1 min, ramp at 30 °C min-1 to 280 °C, held for 16 min and ramp at 50 °C min-1 to 325 °C, held for 3 min. Total time: 27.92 min” | [272]
Human milk | 27 | 150–12730 |  | 76–116.2 | 1.09–15.4 | GC-MS: “initial temperature was set at 70 °C and held for 2 min, increased at 25 °C/min to 150 °C, then immediately increased to 200 °C at a rate of 3 °C/min, finally increased at 8 °C/min to 290 °C, held for 6 min” | [273]
Human milk | 18 | 1.7–4.3 | 5.8–43.1 | 64.7–128.2 |  | GC-MS: “The oven heating program was started at 100.0C (4.00 min), increased from 15.0C min1to 200.0C (0.5 min), at 2.0C min1230.0C (3.0 min), and ending with 15.0C min1to 280.0C, with total time of 27.5 min” | [391]
Milk | 8 | 3700–4800 | 1.2 × 104-1.6 × 104 | 93–97 | <10 | GC-ECD, GC-MS: “The oven program was operated in 6 steps: (1) the starting temperature of the column oven was 40 °C, and this temperature was maintained for 1 min; (2) the temperature was increased to 120 °C at a rate of 20 °C/min, where this temperature was maintained again for 1 min; (3) subsequently, the temperature of the oven was increased to 150 °C at a rate of 10 °C/min, hold 1 min, (4) next at 10 °C/min to 180 °C, hold 1 min; (5) next at 20 °C/min to 200 °C, hold 1 min and (6) finally at 10 °C/min to 290 °C, where it was kept for 2 min (the total running time was 25 min)” | [271]
Breastmilk | 8 | <0.2 |  | >90 | <15 | GC-ECD: “initial oven temperature of 150 °C held for 1 min, ramped at 3 °C/min to 200 °C and then ramped at 8 °C/min to 280 °C held for 10 min” | [392]
breastmilk | 20 | 0.01 | 0.01–0.4 | 69.9–101.3 |  | GC-MS: “oven temperature—start from 90 °C (0.5 min), increase 7 °C/min, 220 °C (12 min), increase 6 °C/min, 285 °C (7 min), increase 5 °C/min, 295 °C (6 min)” | [393]
Water, milk, tea | 8 |  |  | 91–108 | 2.4–9.9 | GC-ECD, GC-MS: “The GC oven temperature program was started at 110 °C and held for 1 min, then increased up to 310 °C at 15 °C/min and held for 1 min. RT 7.30–10.7 min” | [394]
Tap and well water, tea, milk | 5 | 0.002–0.08 | 0.005–0.26 | Tap and well: 84.1–100.9; green tea and milk: 83.4−101.6 | Tap& well: 8.2; green tea& milk: 1.4−8.6 | GC-MS “MRM, the GC oven temperature was set as follows: 70 °C for 1 min, 10 °C min−1 to 180 °C (held for 5 min), 5 °C min−1 to 220 °C, and 30 °C min−1 to 280 °C (held for 8 min)” | [395]
Milk, soil, water | 17 | 0.11 to 0.83 | 0.37–2.75 | 23.35–100.17 |  | GC-ECD: “Splitless, the initial column temperature was 75 °C that was maintained for 1 min, increased at a rate of 25 °C/min to 150 °C, raised at a rate of 6 °C/min to 225 °C, and lastly increased at a rate of 15 °C/min to 290 °C where it was maintained for 10 min. The total analysis time was 31 min” | [396]
Human Serum | 28 | 22.11–128.66 |  | 78–97 | 1.8–13.3 | GC-HRMS: SIM | [40]
Lipid serum of human | 17 | 0.84–25.7 | 3.20–89.9 | 86–120 |  | GC-MS/MS | [41]
Blood | 11 | 0.01–0.5 L |  | Spiked blanks: 93–106, matrix spikes: 94–112 procedural blanks, Surrogate: 63–91 |  | GC-MS: “Splitless mode, column flow rate was set at 1 Ml min− 1 in multiple reaction monitoring mode, with the starting oven temperature at 100 °C (1 min), ramping at 11 °C min− 1 to 180 °C, then 3 ◦Cmin-1 to 260 °C and ultimately to 300 °C at rate of 20 °C min− 1 with final holding time of 6 min” | [397]
Blood | 4 | MDL:0.37–1.8 | MQL: 0.11–0.53 | 82–116 |  |  | [398]
Blood | 9 | 2–5 | 5–40 |  |  | GC-MS/MS | [399]
Serum of man | 15 | MDLs: 0.0113–0.0318 |  |  | >20 | GC-MS/MS: Splitless, | [400]
Blood of man; drinking water | 8 | 0.05 pga |  | 95 |  | GC-ECD | [401]
Blood(serum) of man | 18 | 10–210 |  | 76.0–130.5 |  | GC-MS | [43]
Semen, Blood(serum) of man | 18 | 0.01–0.21 |  | 76.0–130.5 |  | GC-MS | [42]
Blood serum of women | 17 | 0.004–0.11 | 0.02–0.38 | 85–115 | 4–20 | GC-MS/MS: “The oven temperature program was as follows: 70 °C for 3 min, up to 150 °C at a rate of 50 °C min−1, up to 200 °C at a rate of 3 °C min−1 maintained for 1 min, up to 280 °C at a rate of 20 °C min−1 maintained for 5 min and finally, up to 310 °C at a rate of 40 °C min−1 maintained for 4 min” | [402]
Serum of ovarian of human | 9 | 0.002–0.015 | 0.005–0.04 |  |  | GC-MS/MS | [403]
Serum from cord blood | 4 | 0.012 |  | 88–94 |  | GC-MS: “temperature was set on 110 °C for 2 min, then increased by 15 °C/min to 285 °C thus kept for 5 min, and finally increased by 5 °C/min to 300 °C and kept thus for 15 min” | [28]
Human Umbilical cord | 20 |  |  | 3.35–90.07 |  | GC-MS | [261]
Human umbilical cord blood serum | 4 | 0.26–1.12 | 0.86–3.74 | 44.86–71.80 | 1.75–2.43 | GC-MS: “The column temperature was programmed from 50 °C (3 min) to 150° at 30 °C/min, 150 °C to 300° at 10 °C/min (2 min). The carrier gas was helium (purity 99.99 %) with a flow rate of 1 Ml/min. RT: 15.27–22.29 min” | [404]
Plasma/serum | 23 | 0.001–0.029 | 0.002–0.093 | 85.9–109 | 1.97–19.6 | GC-MS/MS: “The oven temperature program was started at 70 °C, increased to 240 °C at 20 °C/min, and continuously increased to 310 °C at 8 °C/min and maintained at the final temperature for 11 min” | [405]
Plasma/serum | 7) | 0.02–0.07 |  |  | 4–40 | GC-MS/MS: “The GC oven program was set to following values: 115 °C (2 min), 4 °C/min, 260 °C (1 min), 30 °C/min, 320 °C (2 min)” | [406]
serum | 26 | 7 × 10−5-0.01344 |  | 30–124 |  | HRMS/HRMS | [407]
Blood serum | 11 |  |  | 74–120 | <20 | GC-MS/MS, UHPLC–MS/MS | [101]
Human tissue (liver, kidney, heart, spleen, lung, brain and abdominal fat) | 20 | 1.0–16 |  | 85–109 | <12 | GTC-MS: “Splitless, the GC oven program started with an isothermal stage at 100 °C for 0.5 min, increasing at 10 °C/min ∼180 °C, and then at 2 °C/min ∼225 °C (2 min), to finally reach 265 °C at a rate of 20 °C/min, which was maintained for 1 min” | [44]
Human blood | 8 | 0.004; 5 × 10−5 nga for detector |  | Ave: 95 |  | GC-ECD | [408]
Human blood | 4,4′DDE, HCB | 0.03 |  | Int. standards: 97.6, 75.62 |  | GC-MS: “Splitless, 90 °C for 2 min followed by an increase of 20 °C/min to 160 °C, then 5 °C/min to 245 °C and held for 5 min, to finally be increased of 20 °C/min up to 300 °C, held 5 min” | [409]
Human blood serum | 19 | 0–0.022 |  | 60–130 |  | GC-ECD: “The oven temperature was programmed from an initial temperature 80 °C for 1 min, and then increasing at 30 °C/min up to 160 (2 min hold) after that, increasing to 260 °C at a rate of 3 °C/min and was maintained at 260 °C for 12 min” | [410]
Serum and venous blood of human | 19 | 0.01–0.21 | 0.03–0.08 | 76.0–130.5; intra/inter batch: 9.0–11.6; biomarkers: 87-102 |  | GC-MS: “The oven temperature program was kept initial temperature at 40 °C for 1.5 min, followed by 25 °C/min to 90 °C for 1.5 min, 25 °C/min to 180 °C, 5 °C/min to 250 °C, and 50 °C/min to 300 °C for 6 min” | [411]
Human placenta | 20 | 0.002–0.17 | 0.006–1.81 | 56–148 |  | GC-MS: “Split, the column temperature raised from the initial temperature of 50 °C (hold for 2 min) to 150 °C (hold for 15 min) at 10 °C/min, then ramped to 240 °C (hold for 30 min) at 3 °C/min, and finally increased to 300 °C (hold for 6 min) at 10 °C/min” | [412]
Placenta sample | 13 | 0.27–0.77 | 0.82–2.30 | 86.2–101.1 | 5.2–10.8 | GC-MS: “Splitless, GC oven program was initiated at 70 °C (with a holding time of 2 min), following temperature increases of 25 °C/min to 150 °C, 5 °C/min to 200 °C (with a holding time of 5 min), 5 °C/min to 270 °C (with a holding time of 5 min), and 25 °C/min to 290 °C (with a holding time of 5 min. RT 44 min” | [280]
Hair | 7 | 0.01–0.1 | 0.02–0.38 | 2.51–200.54 |  | GC-MS: “The starting temperature was 60 °C, held for 1 min, ramped to190°Cat40 °C/min, ramped to 280°Cat5°C/min, held for2 minutes, ramped to 290°Cat10 °C/min and held for 5 min” | [72]
Hair | 25 | 0.1–1.0 | 0.1–25.0 | 67–139 | 1–52 | GC-MS/MS: “It was initially held at 60 °C for 1 min, then raised at a rate of 20 °C/min to 180 °C, held during 1 min. Afterwards, the temperature was increased by 4 °C/min to 240 °C and then by 60 °C/min to 300 °C” | [413]
Hunan Hair | 14 | 0.011–1.38 |  | Int std: 83-180 |  | GC-MS: “The column temperature was initially held at 70 °C for 2 min, raised to 150 °C at the rate of 25 °C/min, then to 200 °C at the rate of 3 °C/min, and to 280 at a rate of 8C, held for 10 min, finally to 300 °C at the rate of 100 °C/min, held at final temperature for 5 min. Total analysis time was 47.067 min” | [289]
Hair, liver |  | Liver: 0.01–0.11; hair: 0.14–1.4 | Liver:0.1–0.50; hair: 0.1–10 | Liver: 74–119; hair: 61- 12 | <20 % | GC-QqQ-MS/MS | [414]
Hair of wild cat | 10 |  | 0.01–0.5 | 94–110 | <10 | GC-ECD: Splitless, “100 °C (1 min), 10 °C/min to 240 °C (1 min), 3 °C/min to 260 °C (1 min), 20 °C/min to 300 °C (10 min)” | [415]
Hair | 7 | 0.01–0.1 | 0.02–0.38 | 2.51–200.54 |  | GC-MS: “The starting temperature was 60 °C, held for 1 min, ramped to190°Cat40 °C/min, ramped to 280°Cat5°C/min, held for2 minutes, ramped to 290°Cat10 °C/min and held for 5 min” | [72]

## Freetext Data
Data | Value | Relevance
Existing literature and methodologies | A systematic review was carried out on 620 articles from Clarivate Web of Science database between 2020 and March 2023. | Relevant in assessing the breadth of existing research and methodologies covered in the paper to identify gaps or limitations for OCP analysis methods.
Quality assurance inclusion criteria | Articles which did not report the quality assurance were discarded. | Relevant as it ensures the focus of the paper is on validated and reliable methodologies, contributing to accurate analysis.
Extraction methods comparison | Techniques such as solid-phase microextraction (SPME), quick component separation, and dispersive liquid-liquid microextraction (DLLME) are compared for various matrices. | Relevant for identifying inconsistencies and providing insights into the best methodologies suitable for specific sample types or analytes.
Inconsistencies in recovery rates | Recovery rates from snow (27%), glacier-ice core (9.67%), water (22.42%), air (38.2%) show variations. | Relevant as it highlights methodological challenges and inconsistencies in achieving accurate and reliable measurements of OCPs.
Relative standard deviation (RSD) | Minimum 5 replicate injections are required; low RSD indicates high precision. | Relevant to provide insights into evaluating technique precision for method validation.
LOD and LOQ calculation techniques | Techniques include 3(S/N), calibration curve slope, laboratory fortified blank; values varied significantly depending on the method used. | Relevant for understanding how the discrepancies impact the sensitivity and reliability of OCP quantification methods.
Instrument techniques comparison | GC-ECD is high-resolution and affordable; GC-MS has low detection limits but challenges with low-concentration OCPs; Liquid chromatography is suitable for polar or thermally unstable compounds. | Relevant for identifying the advantages or limitations of analytical instrumentation to optimize OCP analysis methods.
Validation of accepted criteria | As per SANTE/11813/2017, precision ≤20%, recovery range 70–120%, linearity r > 0.995 are recommended. | Relevant as these criteria provide benchmarks for reliable analysis of OCPs in studies.
Matrix effects | Use of matrix-matched calibration, standard addition, or isotope standards is necessary to offset matrix effects. | Relevant since it ensures accurate results and mitigates potential biases in OCPs analysis due to interference.
Optimization techniques for sensitivity | Experimental methods include Response Surface Methodology (RSM) and one-factor-at-a-time (OFAT) approaches. | Relevant for determining the best practices in enhancing sensitivity and precision of OCP analysis techniques.
Literature Reports | Studies published between 2020 and March 2023 | Provides a timeline for the research considered, which underpins the study design.
Systematic Evaluation | 620 articles screened | Indicates the volume of data reviewed for methodological design relevance.
Inclusion Criteria | Limited to English-language studies | Ensures consistency and accessibility of the reviewed data set, which impacts study design.
Data Extraction Procedures | Centrifuge method used for 2 mL serum samples | Relevant for understanding techniques applied during tissue sampling and analysis validity.
Sample Size for Serum | 2 mL | Provides insight into sample quantity considerations relevant to the PK study design.
LOD for Serum Extraction | 0.004–0.11 ng/g | Indicates precision constraints impacting sampling effectiveness for patient studies.
Recovery Rates in Serum | 85–115% | Provides benchmarks for the reliability of sampling methods in patient populations.
No specific model mentioned | Not explicitly stated in the paper | Relevant for understanding that the study does not address compartmental or physiologically-based modeling for OCPs analysis, emphasizing purely analytical validation techniques.
Acceptance criteria for method validation of OCP analytes | Precision ≤20 %, Recovery 70–120 % | Relevant for evaluating the validity of analytical methods for OCP analysis based on the EU SANTE/11813/2017 document.
Signal-to-noise ratio (S/N) for detecting analytes | 3 (LOD), 10 (LOQ) | Determined by researchers; critical for ensuring the sensitivity of analytical method detection and quantification.
Octanol-water partition coefficient (Kow) | High values | Determines the bioaccumulation and persistence behavior of OCPs in organisms and environmental matrices, essential for analytical accuracy.
Amount of analytes per sample matching lowest calibration point | MDL and MQL | Critical thresholds for determining method sensitivity and performance in detecting trace OCPs.
Standard deviation (σ) | Used for calculated detection limits (LOD: 3σ, LOQ: 10σ) | Parameter assists researchers in estimating sensitivity and quantification capabilities of analytical setups.
Relative recovery (%) | Human tissue: 85–109 %, High-fat matrices: up to 165 % | Validated from specific experimental setups; shows efficiency in extracting OCPs from diverse complex matrices.
Instrument precision | Relative standard deviation (RSD ≤20 %) | Assessed through multiple replicates; optimizes precision in environmental and clinical analyses of OCPs.
Detection limits | <LOD: LOD/2 or <LOQ: LOQ/2 | Suggested by researchers; estimates for non-detection points and improves reported accuracy of OCP levels.
Temperature optimization in high-fat matrices | Methylamine modified graphene d-SPE cleanup approach | Relevant for enhancing LOD and LOQ by optimizing sample properties impacting analytical performance.
Extraction solvent ratios | Hexane/methanol (4:1), Hexane/acetone (1:1) | Improves solubility and extraction efficiency in sample preparation, derived from optimized validation protocols.
Analyte solubility factors | Water content retards solubility | Impacts recovery rates and quality in liquid extraction methodologies validated through experimentation.
Hydrophobicity | OCPs are hydrophobic and lipophilic, leading to bioaccumulation in lipid-rich tissues | Relevance lies in their pharmacokinetics as lipophilicity affects absorption, distribution, and storage in target organisms.
Partition coefficient (Kow) | High octanol-water partition coefficients (Kow) indicate stronger affinity for lipid-based compounds than water | Relevant because partition coefficient influences drug distribution between compartments, particularly lipid-based tissues.
Volatility | OCPs are semi-volatile, and volatilization affects environmental presence and exposure source | Relevant to pharmacokinetics in terms of inhalation exposure and systemic transference due to volatility.
Solubility in solvents | Low water solubility impacts transport and clearance mechanism | Solubility directly affects absorption and bioavailability in organisms exposed to OCPs.
Endocrine-disrupting capabilities | OCPs interfere with hormonal homeostasis and endocrine systems | Relevant in modifying metabolism and pharmacodynamics, causing system-wide biological effects.
Matrix interference | Sample complexity can cause co-eluting matrix compounds to affect recovery and signal detection | Relevant because matrix interferences may mimic or conceal OCP pathways and pharmacological profiles.
Bioaccumulation | OCPs accumulate in the food chain, impacting ecosystem health and human exposure | Influences pharmacokinetics by prolonging systemic retention and altering distribution in lipid compartments.
Half-life and persistence | OCPs resist degradation owing to long half-life (DT50) | Relevant to pharmacokinetics in persistence within biological systems, affecting clearance rates.
Lipophilicity | Lipophilic nature leads to retention in adipose tissues | Relevant because lipophilic retention impacts long-term pharmacokinetics and slow or inconsistent clearance rates.
Metabolic interference | OCPs act as lipid dysregulators and reproductive inhibitors | Directly impacts pharmacokinetics through altered metabolism and reproductive system pathways.
Environmental volatility | Environmental factors like storage and temperature affect OCP levels in samples | Relevant because external factors impact pharmacokinetic profiles upon initial exposure.
Human tissue interaction | Human samples demonstrate specific interference factors due to high lipid content | Relevant as lipid interaction modulates the distribution and retention in human tissues.