## Table 1
Rat | Rise time in the plasma* (min) | Rise time in the ISF** (min)
Rat 1 | 6 | 25
Rat 2 | 11 | 29
Rat 3 | 10 | 59
Rat 4 | 5 | SubQ1: 59SubQ1: 37
Rat 5 | 2 | SubQ1: 35SubQ2: 34
Average | 6.8 | 39.7

## Table 2
Rat identifier | Average and standard deviation of plasma concentration after the 2 μM set point is first reached (μM) | Average and standard deviation of subcutaneous concentration after the 2 μM set point is first reached (μM)
Rat 1 | 2.2 ± 0.4 | 2.0 ± 0.1
Rat 2 | 1.9 ± 0.2 | 2.10 ± 0.07
Rat 3 | 2.2 ± 0.3 | 2.2 ± 0.3
Rat 4 | 2.0 ± 0.4 | SubQ1: 1.8± 0.1SubQ2: 1.9± 0.1
Rat 5 | 2.0 ± 0.2 | SubQ1: 1.80 ± 0.09SubQ2: 1.90 ± 0.14

## Freetext Data
Data | Value | Relevance
Rise time (vein) | 2 to 11 minutes (average 6.8 minutes) | Relevant as it demonstrates the speed with which plasma drug concentrations reach the set point during feedback control, an essential factor for clinical application.
Rise time (ISF) | Rapid equilibration within 30–60 minutes | Relevant as it shows the time required for ISF drug concentrations to closely match plasma concentrations, validating the assumption of using plasma as a proxy for solid tissue concentrations.
Correlation between measured plasma and ISF drug concentrations | RMS deviation of 8–21% | Relevant as it confirms the accuracy of feedback-controlled plasma concentrations in achieving matching drug levels in ISF, supporting clinical feasibility.
Correlation between paired subcutaneous sensors | Pearson correlations: R2 = 0.95–0.99 | Relevant as it indicates strong reproducibility and reliability of subcutaneous sensors for tracking drug concentrations, reinforcing sensor calibration and efficacy in solid tissue contexts.
Systematic deviations between paired subcutaneous sensors | Deviations of 9±2% (Rat 4) and 13±3% (Rat 5) | Relevant as it highlights the sensor-to-sensor fabrication variability, a consideration for technical implementation and calibration in future studies.
Measured plasma vs ISF concentrations after equilibration | Difference of 0.10 to 0.29 μM | Relevant as it supports the argument that maintaining plasma drug levels results in closely similar ISF levels, aligning with therapeutic objectives for drug delivery applications.
Maximum infusion rate required | Varied rates dependent on individual rat physiology and body weight | Relevant as it demonstrates intersubject variability in pharmacokinetics, key to understanding personalized drug infusion rates.
Drug level equilibration mechanism | Likely passive diffusion | Relevant as it provides insight into the transport mechanism between plasma and ISF, underpinning the observed correlation of concentrations and addressing technical assumptions in feedback control systems.
Calibration approach for ISF sensors | Blood used as proxy calibrant | Relevant as it addresses the technical challenge of ISF sensor calibration and validates the suitability of blood as a calibrant medium for sensors in solid tissues.
Sensor baseline noise floors | 0.4 μM in vein, 0.2 μM in subcutaneous space | Relevant as baseline noise levels impact the precision and reliability of drug concentration measurements, critical for technical deployment and clinical decisions.
Patient sample size | 5 adult male Sprague-Dawley rats | Indicates the number and type of subjects used in the study, relevant for assessing the experimental setup and generalizability.
Blood sampling time points | 20–30 min baseline, 2 to 11 min rise time to plasma set point, ongoing monitoring post-equilibration | Provides details on timing for plasma drug monitoring, critical for understanding the pharmacokinetic study design.
Tissue sampling time points | Equilibration within 30–60 min; monitored until ISF concentration sustained ±15% of the plasma set point for 5 min | Describes tissue measurement timing relative to plasma measurements, essential for analyzing tissue equilibration dynamics.
Patient population | Adult male Sprague-Dawley rats (4–5 months old, 400–650 g) | Specifies the subject demographics, relevant for understanding pharmacokinetics in this model organism.
Blood sampling method | Intravenous sensor in the right jugular vein | Provides the method of plasma drug monitoring, critical for pharmacokinetic data collection.
Tissue sampling method | Subcutaneous sensors placed ventrally below the skin, ~1 cm apart for paired measurements | Describes how ISF drug concentration measurements were obtained, important for ensuring the accuracy of the tissue pharmacokinetic analysis.
Model type | Two-compartment pharmacokinetic model | This model was employed by the adaptive feedback control algorithm to fit the individual animal's pharmacokinetics and control the infusion rate for maintaining plasma drug concentrations effectively.
Controller | Adaptive feedback control algorithm | The algorithm used the two-compartment pharmacokinetic model as part of its approach to learn the pharmacokinetics of each animal dynamically and adjust drug infusion rates accordingly.
---
KD | 5 μM | KD is the binding half-point of the aptamer, estimated during in vitro titration experiments targeting doxorubicin in blood plasma to calibrate the aptamer's sensitivity.
nH | Hill Coefficient | Estimated from in vitro titration of sensors to assess the aptamer's drug-binding dynamics.
KDMmax | Maximum signal gain at saturating concentrations | Derived from signal behavior during titration experiments in whole rat blood, used to calibrate sensor responses.
KDMmin | Minimum signal observed without drug | Established during baseline measurements and sensor titration to anchor the sensitivity spectrum of the sensor.
Rise time (Vein) | 2–11 mins | Measured directly during in vivo experiments to evaluate sensor response and feedback control speed in the plasma.
Rise time (ISF) | 30–60 mins | Measured in vivo to determine how quickly interstitial fluid drug levels equilibrated with plasma concentrations.
Sensor noise floor (Plasma) | 0.4 μM | Established in pre-drug, baseline in vivo conditions to inform precision in sensor measurements.
Sensor noise floor (ISF) | 0.2 μM | Measured during baseline ISF sensing for accuracy calibration.
Correlation (Plasma vs ISF) | ~15% deviation after equilibration | Observed directly during live experiments to validate plasma concentration control as a proxy for ISF drug concentration.
Correlation (Paired Subcutaneous Sensors) | R2 = 0.95–0.99 | Derived from simultaneous measurements in the ISF to validate reproducibility and local calibration accuracy within solid tissues.
High inter- and intra-subject variability of doxorubicin | Known to impact clinical outcomes significantly | Demonstrates that individual differences in drug metabolism and elimination influence pharmacokinetics.
Kidney failure's impact on doxorubicin elimination rate | Reduced elimination rate to zero due to doxorubicin-induced kidney failure | Highlights how organ dysfunction can directly alter pharmacokinetic processes.
Adaptive feedback control algorithm | Adjusts to individual rat’s pharmacokinetics using real-time data | Shows the importance of personalized approaches to account for variable pharmacokinetics among subjects.
Drug infusion rate variability within animals | Infusion rate required to maintain plasma drug concentration dropped to 0 | Suggests dynamic physiological changes in animals over time affecting pharmacokinetics.
Leaking of drug from the pump | Slight upward drift in drug concentration hypothesized due to pump leakage | Points to external experimental factors influencing pharmacokinetics in controlled settings.