## Table 1
Transitions | Precursor (M + H)+ | Product ion | Dwell (ms) | Declustering potential (DP) | Collision energy (CE)
Carfentanil | 395 | 335 | 62 | 60 | 30
Carfentanil | 395 | 113 | 62 | 60 | 43
Carfentanil | 395 | 105 | 62 | 60 | 63
Caffeine | 195 | 138 | 62 | 60 | 15
Caffeine | 195 | 110 | 62 | 60 | 20
Caffeine | 195 | 83 | 62 | 60 | 41
Sufentanil | 387 | 238 | 62 | 30 | 23
Sufentanil | 387 | 355 | 62 | 30 | 23
Sufentanil | 387 | 111 | 62 | 30 | 52

## Table 2
Expected conc. (ng/ml)_Expected conc. (ng/ml) | Day 1_Mean conc. (ng/ml) | Day 1_Accuracy (%) | Day 2_Mean conc. (ng/ml) | Day 2_Accuracy (%) | Day 3_Mean Conc. ng/ml) | Day 3_Accuracy (%)
0.046 | 0.09 | 193 | 0 | 0 | 0.042 | 91.25
0.069 | 0.2 | 292.4 | 0.068 | 98.53 | 0.065 | 94.62
0.23 | 0.21 | 91.8 | 0.24 | 104.25 | 0.27 | 116.43
0.69 | 0.74 | 107.5 | 0.68 | 98.76 | 0.72 | 104.8
2.3 | 2.34 | 101.9 | 2.31 | 100.25 | 2.17 | 94.35
6.9 | 6.83 | 98.9 | 6.72 | 97.44 | 6.73 | 97.51
36.8 | 36.73 | 99.8 | 37.55 | 102.04 | 37.77 | 102.63
46 | 46.06 | 100.1 | 45.42 | 98.74 | 45.27 | 98.41

## Table 3
Expected conc.(ng/ml)_Expected conc.(ng/ml) | Day 1_Mean conc.(ng/ml) | Day 1_Accuracy(%) | Day 2_Mean conc.(ng/ml) | Day 2_Accuracy(%) | Day 3_Mean Conc.(ng/ml) | Day 3_Accuracy(%)
46 | 46.9 | 101.93 | 46.2 | 100.39 | 45.6 | 99.21
138 | 135.1 | 97.87 | 137.5 | 99.6 | 141.3 | 102.39
368 | 359.5 | 97.69 | 364.5 | 99.04 | 354.3 | 96.29
522 | 573 | 103.8 | 563.8 | 102.1 | 559.7 | 101.4
736 | 731.9 | 99.45 | 724.9 | 98.5 | 750.9 | 102.02
920 | 913.3 | 99.28 | 923.3 | 100.4 | 908.1 | 98.71

## Table 4
Expected conc.(ng/ml)_Expected conc.(ng/ml) | Day 1_Mean conc.(ng/ml) | Day 1_Accuracy(%) | Day 2_Mean conc.(ng/ml) | Day 2_Accuracy(%) | Day 3_Mean conc.(ng/ml) | Day 3_Accuracy(%)
7.9 | 7.3 | 92 | 3.3 | 41.9 | 11.4 | 143.5
11.9 | 10.9 | 92 | 11.9 | 86.1 | 13.1 | 110.4
39.6 | 42.3 | 106.8 | 39.6 | 110.1 | 42.0 | 106
59.4 | 62.3 | 104.8 | 59.4 | 92.8 | 59.3 | 99.8
198 | 205.6 | 103.8 | 198.0 | 100.9 | 178.8 | 90.3
597 | 601 | 100.7 | 597.0 | 104.2 | 565.3 | 94.7
1584 | 1579 | 99.7 | 1584 | 105.7 | 1571 | 99.2
1980 | 2003 | 101.1 | 1980 | 102.5 | 1914 | 96.7
5940 | 5886 | 99.1 | 5940 | 97.5 | 6367 | 107.2
7920 | 7802 | 98.5 | 7920 | 101.8 | 7453 | 94.1
11 880 | 12 340 | 103.9 | 11 880 | 100.9 | 11 970 | 100.7
15 840 | 15 310 | 96.6 | 15 840 | 94.5 | 16 240 | 102.5
19 800 | 20 020 | 101.1 | 19 800 | 103 | 19 480 | 98.4

## Table 5
Expected conc.(ng/ml)_Expected conc.(ng/ml) | Day 1_Measured conc.(ng/ml) | Day 1_Accuracy (%) | Day 2_Measured conc.(ng/ml) | Day 2_Accuracy (%) | Day 3_Measured conc.(ng/ml) | Day 3_Accuracy (%)
1.54 | 1.65 | 107.1 | 1.52 | 98.77 | 1.59 | 103.07
1.54 | 1.51 | 98.3 | 1.51 | 98.20 | 1.66 | 108.00
1.54 | 1.45 | 94.3 | 1.77 | 115.17 | 1.57 | 101.88
1.54 | 1.48 | 96.1 | 1.44 | 93.44 | 1.42 | 92.24
1.54 | 1.39 | 90.2 | 1.45 | 93.94 | 1.38 | 89.90
| Mean accuracy = 97.20 | Mean accuracy = 97.20 | Mean accuracy = 99.90 | Mean accuracy = 99.90 | Mean accuracy = 99.02 | Mean accuracy = 99.02
| %CV = 6.46 | %CV = 6.46 | %CV = 8.88 | %CV = 8.88 | %CV = 7.73 | %CV = 7.73
40.48 | 36.15 | 89.31 | 39.50 | 97.57 | 38.48 | 95.06
40.48 | 36.27 | 89.60 | 39.68 | 98.03 | 39.92 | 98.63
40.48 | 36.79 | 90.87 | 38.34 | 94.72 | 41.62 | 102.81
40.48 | 36.04 | 89.03 | 40.41 | 99.82 | 38.59 | 95.32
40.48 | 37.17 | 91.83 | 40.03 | 98.89 | 40.07 | 98.98
| Mean accuracy = 90.13 | Mean accuracy = 90.13 | Mean accuracy = 97.81 | Mean accuracy = 97.81 | Mean accuracy = 98.16 | Mean accuracy = 98.16
| %CV = 1.31 | %CV = 1.31 | %CV = 1.97 | %CV = 1.97 | %CV = 3.23 | %CV = 3.23
625.6 | 601.8 | 96.20 | 639.54 | 102.23 | 611.90 | 97.81
625.6 | 597.1 | 95.45 | 609.06 | 97.36 | 595.30 | 95.15
625.6 | 601.6 | 96.16 | 654.88 | 104.68
625.6 | 622.6 | 99.53 | 615.65 | 98.41 | 597.80 | 95.55
625.6 | 606.4 | 96.94 | 634.43 | 101.41 | 591.40 | 94.53
625.6 | Mean accuracy = 96.86 | Mean accuracy = 96.86 | Mean accuracy = 100.82 | Mean accuracy = 100.82 | Mean accuracy = 95.76 | Mean accuracy = 95.76
625.6 | %CV = 1.64 | %CV = 1.64 | %CV = 6.46 | %CV = 6.46 | %CV = 1.49 | %CV = 1.49

## Table 6
Expected conc.(ng/ml)_Expected conc.(ng/ml) | Day 1_Measured conc.(ng/ml) | Day 1_Accuracy (%) | Day 2_Measured conc.(ng/ml) | Day 2_Accuracy (%) | Day 3_Measured conc.(ng/ml) | Day 3_Accuracy (%)
132.78 | 117.4 | 88.38 | 152.4 | 114.79 | 126.6 | 95.31
132.78 | 118.1 | 88.97 | 147.9 | 107.6 | 131.5 | 99.03
132.78 | 116.7 | 87.92 | 156.7 | 118.03 | 129.3 | 97.39
132.78 | 113.0 | 85.13 | 135.3 | 101.87 | 131.0 | 98.63
132.78 | 113.2 | 87.23 | 135.7 | 102.2 | 123.0 | 92.62
132.78 | Mean accuracy = 87.13 | Mean accuracy = 87.13 | Mean accuracy = 108.9 | Mean accuracy = 108.9 | Mean accuracy = 96.6 | Mean accuracy = 96.6
132.78 | %CV = 2.08 | %CV = 2.08 | %CV = 6.72 | %CV = 6.72 | %CV = 2.75 | %CV = 2.75
1742 | 1522 | 87.37 | 1895 | 108.77 | 1615 | 92.71
1742 | 1563 | 89.72 | 1867 | 107.12 | 1655 | 94.99
1742 | 1536 | 88.16 | 1806 | 103.62 | 1751 | 100.49
1742 | 1542 | 88.51 | 1891 | 108.51 | 1638 | 94.00
1742 | 1623 | 93.17 | 1850 | 106.18 | 1766 | 101.36
1742 | Mean accuracy = 89.39 | Mean accuracy = 89.39 | Mean accuracy = 106.84 | Mean accuracy = 106.84 | Mean accuracy = 96.71 | Mean accuracy = 96.71
1742 | %CV = 2.55 | %CV = 2.55 | %CV = 1.95 | %CV = 1.95 | %CV = 4.08 | %CV = 4.08
13 464 | 12 660 | 94.02 | 13 711 | 101.84 | 14 950 | 111.02
13 464 | 12 670 | 94.10 | 12 867 | 95.57 | 13 590 | 100.94
13 464 | 12 720 | 94.47 | 13 891 | 103.17
13 464 | 12 670 | 94.13 | 13 475 | 100.08 | 13 200 | 98.06
13 464 | 12 260 | 91.08 | 13 375 | 99.34 | 13 780 | 102.34
13 464 | Mean accuracy = 93.56 | Mean accuracy = 93.56 | Mean accuracy = 100.00 | Mean accuracy = 100.00 | Mean accuracy = 103.09 | Mean accuracy = 103.09
13 464 | %CV = 1.49 | %CV = 1.49 | %CV = 2.89 | %CV = 2.89 | %CV = 5.41 | %CV = 5.41

## Table 7
Analyte | Expected conc. (ng/ml) | Mean conc. (ng/ml) | Accuracy (%) | %CV
Carfentanil | 1.54 | 1.52 ± 0.11 | 98.7 | 7.25
Carfentanil | 40.48 | 38.74 ± 1.82 | 95.7 | 4.71
Carfentanil | 625.60 | 628.10 ± 61.86 | 100.4 | 9.85
Caffeine | 132.78 | 129.50 ± 13.50 | 97.5 | 10.4
Caffeine | 1742.0 | 1707.0 ± 130.0 | 98 | 7.63
Caffeine | 13 464.0 | 13 663.0 ± 12.1 | 101.5 | 12.12

## Freetext Data
Data | Value | Relevance
--- | --- | ---
LC-MS/MS Method Development | Quantitative liquid chromatography/hybrid triple quadrupole‐linear ion trap MS/MS | Relevant as it describes the analytical method established for detecting carfentanil and caffeine.
Analytes of Interest | Carfentanil and Caffeine | Relevant because these were the compounds quantified and assessed in this study.
Target Application | Human epidermis absorption kinetics of Carfentanil | Relevant as it establishes the focus on absorption and permeation properties through human skin.
Carfentanil Potency | 10,000 times more potent than morphine | Relevant as it underlines the clinical and occupational significance of studying carfentanil absorption.
Solvent Effects | Organic solvents decreased the absorption rate of carfentanil | Directly relates to the study's key finding about how solvent composition affects absorption.
Study Model | Live human epidermal model (RhE) | Important as it indicates the experimental setup for observing percutaneous absorption.
Internal Standard Used | Sufentanil | Relevant to the analytical validation process for quantifying target compounds in LC-MS/MS.
Carfentanil LOD | 230 pg/ml | Indicates method sensitivity, essential for applicability in detecting low levels of carfentanil.
Caffeine LOD | 12 ng/ml | Establishes method sensitivity for the marker used for epidermal viability monitoring.
Calibration Linearity | Carfentanil and Caffeine linear over at least 4 orders of magnitude | Relevant as it confirms the method's suitability for quantitative analysis across a wide range of concentrations.
Analytical Accuracy | Calibration accuracy within ±20% for both compounds | Highlights method reliability for quantifying target analytes.
Precision (%CV) | ≤15% CV for all analytes | Demonstrates consistent reproducibility of measurements.
Human Lethal Dose of Carfentanil | ~20 µg | Relevant to understanding the extreme potency and occupational hazard associated with trace exposures.
Franz Cell System | In vitro experimental setup | Relevant because it describes the specific apparatus used for studying percutaneous absorption.
Regulatory and Safety Guidance | NIOSH discourages alcohol‐based sanitizers for responders | Relates to practical applications and concerns from prior assumptions influencing the study.
Carfentanil Calibration Range | 230 - 4600 pg/ml (low); 46 - 920 ng/ml (high) | Indicates method scope for quantifying analyte concentrations relevant to exposure studies.
Caffeine Calibration Range | 12 - 19,800 ng/ml | Relevant for assessing epidermal viability, validating the experimental model.
MS Instrument Used | SCIEX 4000 QTRAP | Specifies the tool used for analytical measurements, essential for method reproducibility.
Major MS/MS Mode | Multiple Reaction Monitoring (MRM) | Relevant as it details the technique used for specific and sensitive quantification of the analytes.
Conclusion on Alcohol Solvents | Reduced carfentanil permeation, contrary to initial concerns | Relevant finding indicating lower occupational exposure risk when using alcohol-based sanitizers.
Use of in vitro Franz Cell system | The PK study utilized an in vitro Franz Cell system incorporating live, reconstructed human epidermis (RhE) with intact stratum corneum | Relevant because the design of the study using a human epidermis model provided insights into carfentanil's absorption kinetics necessary for determining its percutaneous permeability coefficients and other pharmacokinetic measures.
Sampling process | Samples were drawn from the lower chamber before dosing and at different time points after dosing | Relevant because sampling from the lower chamber at multiple intervals allowed researchers to evaluate the absorption and permeability kinetics over time.
Human epidermis preparation | The live, cultured human epidermis was suspended stratum corneum side up at an air/liquid interface | Relevant because this preparation allowed accurate study of drug absorption as it mimics human skin's exposure conditions.
Cell viability marker | Caffeine was employed as a cell viability marker in the in vitro model | Relevant because caffeine's use ensured the reliability of the epidermal cells during the experiment and supported the study design's validity.
Dosing solvents evaluated | Carfentanil was formulated in four vehicles: water, ethanol, and two ethanol-containing hand sanitizers | Relevant because solvent composition is crucial to determining its impact on percutaneous absorption kinetics in pharmacokinetics studies.
Patient sample model | Live, reconstructed human epidermis model used; no in vivo human subjects | Relevant because the in vitro human epidermal model provided a controlled environment to study carfentanil’s absorption without involving human populations.
Time points | Samples were collected at various time intervals post-dosing | Relevant because multiple sampling intervals allowed the measurement of changes in absorption kinetics over time.
Model | In vitro reconstructed human epidermal (RhE) model | The study utilized this model for assessing percutaneous absorption kinetics of carfentanil in a human skin analog system.
Analysis System | In vitro Franz Cell system | The Franz Cell system was central to the experimental setup, ensuring accurate modeling of the skin absorption process by maintaining conditions such as air/liquid interface and viable epidermal cell layers.
LC/MS-MS Integration | LC/hybrid triple quadrupole-linear ion trap system | This system was validated and employed in the study to quantitatively analyze the pharmacokinetics of carfentanil permeating the reconstructed epidermal model.
Caffeine marker | Caffeine viability marker use | While not directly the mechanistic model, caffeine was an indicator of epidermal cell culture health and supported model reliability during solvent experiments.
CUR | 55 psi | This is the curtain gas parameter for optimal MS/MS performance, relevant to the specificity and sensitivity of the validated LC-MS/MS method.
IS | 5500 V | This ionspray voltage parameter was optimized for electrospray ionization, enhancing accuracy and reproducibility in carfentanil detection.
TEM | 700 °C | Turbo gas temperature was calibrated for stable ion generation, ensuring the suitability of the triple quadrupole instrument performance.
GS1 | 60 psi | The gas one pressure setting aids in optimizing ion flow and collision efficiency during MS/MS methodology validation.
GS2 | 60 psi | Gas two pressure works alongside GS1 to maintain instrument reliability in quantitative high-sensitivity analysis.
ihe | ON | Interface heater was activated ensuring efficient desolvation in the MS process for achieving valid experimental results.
CAD voltage | MED | Collision-assisted dissociation voltage optimized the fragmentation of analytes for accurate multiple reaction monitoring (MRM).
EP | 10 V | Entrance potential was adjusted for precise precursor ion selection during MS/MS analysis.
CXP | 9 V | Collision cell exit potential was determined for accuracy in detecting fragment ions, vital for specificity validation.
Carfentanil calibration curve dynamic concentration range | 230–4600 pg/ml (low) and 46–920 ng/ml (high) | These concentration ranges were estimated during intra- and inter-day validation reflecting the method's dynamic range for carfentanil quantification.
Caffeine calibration curve dynamic concentration range | 12–19,800 ng/ml | This range was validated to evaluate epidermal cell viability, directly relevant to assessing the skin model.
Declustering potential (DP) | 50 V | Used in product ion scanning, optimizing precursor ion declustering for sensitivity in detecting carfentanil and caffeine.
Carfentanil LOD | 230 pg/ml | Determined via signal-to-noise ratio, informing method sensitivity.
Carfentanil LOQ | 230 pg/ml | Defined from accuracy and precision criteria, emphasizing carfentanil quantification thresholds.
Caffeine LOD | 12 ng/ml | Signal-to-noise estimation, critical for assessing caffeine's validation as a viability marker.
Caffeine LOQ | 12 ng/ml | Delineates threshold specificity and ensures correct control during the human epidermis model validation.
Collision energy for carfentanil | 0‐120 V | Estimated via ramp experiments; critical for optimizing fragmentation patterns in MS/MS.
Collision energy for caffeine | 0-60 V | Empirically optimized for generating reliable product ions, aiding model validation.
Counterintuitive influence of alcohol modifiers | Alcohol modifiers appeared to reduce permeation rates and peak concentrations, prolonging permeation lag times | The finding illustrates how solvent polarity impacts carfentanil permeability, affecting pharmacokinetics.
Solvent composition | Carfentanil formulated in ethanol-containing solvents had slower partitioning compared to water-alone formulations | Demonstrates how solvent composition can modulate absorption kinetics, key to understanding drug pharmacokinetics.
Reduced solvent polarity | Alcohol-containing solvents exhibited reduced polarity compared to water, slowing carfentanil partitioning | Highlights the physicochemical properties influencing drug absorption rates, relevant to pharmacokinetic considerations.
Law enforcement hand sanitizer use | Alcohol-based hand sanitizers do not enhance dermal absorption rates for carfentanil | Indicates occupational exposure concerns tied to altered drug pharmacokinetics may be exaggerated.