## Table 1
Component | Retention time (min) | Positive ion mode MRM transition trace (m/z) | Cone voltage (V) | Collision energy (eV)
Pemetrexed | 2.71 | 428.08 > 281.06 | 14 | 18
Pemetrexed Q | 2.71 | 428.08 > 163.00 | 14 | 36
[13C5]‐Pemetrexed | 2.71 | 433.10 > 281.01 | 14 | 20

## Table 2
Validation sample | Nominal concentration (μg/L) | Within‐day accuracy (n = 5) (%) | Within‐day imprecision (RSD%) | Between‐day accuracy (n = 15) (%) | Between‐day imprecision (RSD%)
LLOQ | 0.0252 | 74.2 | 6.04 | 77.4 | 5.39
QC low | 0.0602 | 88.4 | 4.03 | 90.2 | 0.78
Alternative LLOQ | 0.075 | 107 | 6.05 | 103 | 3.38
QC alternative low | 0.12 | 90.6 | 4.57 | 98. 6 | 6.77
QC medium | 1.2 | 108 | 8.79 | 104 | 3.99
QC high | 12 | 102 | 2.86 | 101 | 0
HLOQ | 25.2 | 105 | 5.93 | 102 | 3.23
Average |  | 96.5 |  | 96.5

## Table 3
Matrix | Condition | Time interval | Mean concentration compared to nominal concentration (%)
Stock solutions | –40°C | 24 months | 102 ± 0.3 (n = 3)
Spiked plasma | Room temperature 4°C –40°C Three freeze–thaw cycles | 7 days 7 days 14 months | 99.3 ± 2.4 (n = 6) 94.9 ± 6.3 (n = 6) 108 ± 4.8 (n = 6) 107 ± 5.5 (n = 6)
Spiked whole blood | Room temperature 4°C | 3 days 3 days | 97.9 ± 12.2 a (n = 6) 102 ± 6.7 (n = 6)

## Freetext Data
Data | Value | Relevance
Carryover threshold | Less than 20% of the LLOQ for pemetrexed and 5% of the IS | Relevant to ensuring assay accuracy and minimal interference during validation.
Accuracy range | Accuracy for the assay in plasma was 96.5 and 96.5% | Shows reliability and performance of the assay methodology for quantifying pemetrexed.
Precision coefficients | Within‐day and between‐day precision in coefficients of variations was <8.8% | Demonstrates consistency and reproducibility of the assay.
Extraction recovery | Recovery was 59 ± 1% for pemetrexed and 55 ± 5% for the IS | Indicates effectiveness of the solid-phase extraction method used in sample preparation.
Stability interval for processed plasma samples | Stable for 2 days in a cooled autosampler at 10°C | Relevant to sample handling and storage conditions to ensure integrity during analysis.
Detection run time | Total run time of 9.5 minutes for detection | Relevant to optimizing analytical throughput while ensuring sensitivity and specificity.
Calibration range | 0.0250–25.0 μg/L pemetrexed | Set to capture the expected concentration range in plasma post-microdosing study.
Instrument used for detection | Xevo TQ‐s micro tandem mass spectrometer with UPLC system | Relevant as the technical equipment utilized in achieving sensitivity for low concentrations.
Hemolysis effects | 9.24% deviation (QC low) in hemolyzed plasma | Indicates the assay’s capacity to handle varying plasma conditions.
Lipemic plasma impact | 5.71% deviation (QC low) with 150 mg/dL intralipid | Relevant to understanding effects of matrix variability on assay performance.
ClinicalTrials.gov Identifier | NCT03655834 | Provides the identifier for the clinical microdosing study mentioned in the paper, which can give access to further details about the study design.
PK Study Design | Pharmacokinetic curve with blood sampling at 0.5–1, 1–2, 3–4, and 6–8 hours post-dose | Details the specific time points for plasma sampling after dosing, critical for understanding the timeline of sample collection and drug concentration assessment.
Pemetrexed Dose | 100 μg | Key to understanding the microdose used in the study, which highlights that it is several orders below therapeutic levels and suitable for microdosing PK analysis.
Patient Population | Six oncological patients without prior pemetrexed treatment | Indicates the specific population involved in the study, which is relevant for interpreting PK data and avoiding confounding factors like prior exposure.
Blood Type Sample Source | Drug-free EDTA-plasma from blood bank donors | Demonstrates the use of plasma from oncology patients not treated with pemetrexed, ensuring baseline data without interference by prior exposure.
Duration of Sample Storage | Samples stored at −40°C, stability confirmed | Ensures that sample integrity was maintained throughout the study period and confirms reliable PK measurements post-administration.
Monte Carlo simulation | Previously developed population pharmacokinetic model (Latz et al., 2006) | Monte Carlo simulation using an established model was used to predict expected drug concentrations for setting the calibration range.
Calibration | Quadratic regression model (weighted 1/x²) | The calibration curve construction employed a quadratic model to accurately fit pemetrexed concentration data, essential for pharmacokinetic evaluation.
Population PK model | Correlation with therapeutic exposure and toxicity (Latz et al., 2006) | The study utilized a population pharmacokinetic model that established links between drug exposure, toxicity, and efficacy, applicable to individualizing doses.
Accuracy | 96.5% | Relevant because assay validation accuracy confirms reliability of the method to detect pemetrexed in plasma consistently.
Within-day precision | <8.8% coefficient of variation | Relevant because within-day precision is crucial for determining variability in a short timeframe during the assay validation.
Between-day precision | <8.8% coefficient of variation | Relevant because between-day precision helps assess the assay's reproducibility over multiple days of analysis.
Extraction recovery | 59 ± 1% for pemetrexed | Relevant for assessing the efficiency of solid-phase extraction, crucial for accurate quantification of pemetrexed.
Extraction recovery of internal standard | 55 ± 5% for [13C5]‐pemetrexed | Relevant because the consistency in recovery of internal standards aids reliable normalization of results.
Processed sample stability | 102 ± 7.41% for 2 days at 10°C | Relevant because analyzed samples need to remain stable during processing and storage.
Calibration curve regression coefficient | 0.9978 ± 0.0004 | Relevant for defining linearity and sensitivity of the assay in the validated concentration range.
Analyte MRM transition for quantification | m/z 428.08/281.06 | Relevant because method detection parameters determine assay accuracy for pemetrexed quantification.
Internal standard MRM transition | m/z 433.10/281.01 | Relevant because optimized transitions ensure robust detection of the [13C5]‐pemetrexed IS.
Retention time of pemetrexed | 2.8 min | Relevant because retention time is fundamental for identification and separation of the analyte during chromatography.
pH of eluent | pH of 3 | Relevant because more acidic conditions optimize sensitivity and reduce noise during detection.
Flow rate during chromatography | 0.3 mL/min | Relevant for steady mobile phase transport contributing to the reproducibility and efficiency of the assay.
Column temperature | 40°C | Relevant as temperature consistency directly affects column performance and analyte separation.
Mobile phase composition | 85% A (0.1% formic acid in water), 15% B (0.1% formic acid in acetonitrile) | Relevant because it determines pemetrexed elution efficiency and peak stability.
Autosampler storage temperature | 10°C | Relevant to ensuring the stability of processed samples during extended runs.
Freeze-thaw cycle stability | 107 ± 5.5% after three cycles | Relevant for proving pemetrexed's stability in plasma samples subject to multiple freeze-thaw processes.
Intralipid concentration impact on recovery | <15% deviation at 150 and 300 mg/dL concentrations | Relevant to ruling out interference of lipemic plasma in pemetrexed quantification.
Hemolysis impact on recovery | <15% deviation at QC concentrations | Relevant for confirming assay compatibility with hemolyzed samples.
Monte Carlo simulation | Latz et al., 2006 | Relevant because the concentration range (0.0250–25 μg/L) was estimated based on predictions from this pharmacokinetic population model.
SPE cartridge batch variability | <15% difference between batches | Relevant for ensuring reproducibility and reliability of solid-phase extraction.
Desolvation gas flow | 950 L/h | Relevant because proper gas flow optimizes electrospray efficiency and signal response in MS/MS detection.
The pH of the eluent | Lower pH of 3 yielded more response and less noise than pH of 7 | Relevant because pH influences assay sensitivity, impacting accurate measurements of pharmacokinetics.
Sample pretreatment method | Solid-phase extraction (SPE) was selected due to pemetrexed’s polar nature | Relevant because the sample preparation method directly affects the reliability of pharmacokinetic data.
Extraction recovery | 59 ± 1% for pemetrexed and 55 ± 5% for its internal standard | Relevant as low recovery rates could affect quantification accuracy in pharmacokinetics studies.
Carryover issue | Reduced by using new long needle tips for every standard or sample | Relevant since carryover could result in inaccurate pharmacokinetic measurements.
Plasma type | Validation performed in defrosted filtered EDTA–plasma | Relevant because plasma choice and preparation method impact stability and quantification in pharmacokinetic analysis.
Autosampler storage temperature | Processed samples were stable for 2 days at 10°C with ascorbic acid added | Relevant because storage conditions influence sample stability during pharmacokinetic analysis.
Matrix effects | IS-normalized matrix factor had <15% variability | Relevant as matrix variability impacts assay sensitivity and pharmacokinetics measurement.
Lipemic plasma | Nominal concentrations showed deviations of 5.71% and 4.22% with lipemic plasma | Relevant because lipemia introduces potential errors in pharmacokinetic quantification.
Hemolyzed plasma | Deviation of 9.24% and 4.32% observed for QC low and QC high concentrations | Relevant since hemolysis can interfere with drug quantification and pharmacokinetics.
Processed sample degradation | Stabilized by adding ascorbic acid to prevent oxidation | Relevant due to potential for degradation influencing observed pharmacokinetic profiles.