## Table 1
Entrance potential (V) | 10.0
Curtain gas (psi) | 15.0
Collision gas | Medium
Ionspray voltage (V) | 5500.0
Temperature (°C) | 600.0
Ion source gas 1 (psi) | 40.0
Ion source gas 2 (psi) | 40.0

## Table 2
Q1 | Q3 | Time (ms) | ID | DP (V) | CE (V) | CXP (V)
253.011 | 182.1 | 150 | Phenytoin | 71 | 27 | 10
263.152 | 192.088 | 150 | d10‐Phenytoin | 31 | 37 | 12
269.051 | 198.1 | 150 | 4‐HPPH | 71 | 25 | 6

## Table 3
Analyte in plasma_Analyte in plasma | Nominal conc. (ng/mL)_Nominal conc. (ng/mL) | Interday (n = 3)_Mean (ng/mL) | Interday (n = 3)_Accuracy (%) | Interday (n = 3)_Precision (CV%)
Phenytoin | 15.6 | 15.3 | 98.1 | 4
Phenytoin | 62.5 | 62.2 | 99.6 | 3.1
Phenytoin | 250 | 255.2 | 102.1 | 4.5
4‐HPPH | 15.6 | 16.8 | 107.6 | 4.3
4‐HPPH | 62.5 | 65 | 104 | 3.2
4‐HPPH | 250 | 272.2 | 108.9 | 3.6

## Table 4
Unnamed: 0_level_0_Analyte in plasma | Unnamed: 1_level_0_Nominal conc. (ng/mL) | Intraday 1 (n = 3)_Mean (ng/mL) | Intraday 1 (n = 3)_Accuracy (%) | Intraday 1 (n = 3)_Precision (%CV) | Intraday 2 (n = 3)_Mean (ng/mL) | Intraday 2 (n = 3)_Accuracy (%) | Intraday 2 (n = 3)_Precision (%CV) | Intraday 3 (n = 3)_Mean (ng/mL) | Intraday 3 (n = 3)_Accuracy (%) | Intraday 3 (n = 3)_Precision (%CV)
Phenytoin | 15.6 | 15.8 | 92.6 | 5.1 | 14.5 | 92.6 | 5.1 | 15.75 | 100.8 | 3
| 62.5 | 59.6 | 102.7 | 5.9 | 64.2 | 102.7 | 5.9 | 62.9 | 100.7 | 1.2
| 250 | 243.8 | 100.4 | 2.9 | 250.9 | 100.4 | 2.9 | 270.9 | 108.4 | 1.1
4‐HPPH | 15.6 | 17.1 | 109.1 | 1.9 | 17.5 | 112.3 | 5.5 | 15.8 | 101.4 | 4
| 62.5 | 65.2 | 104.4 | 4.5 | 67.4 | 107.8 | 7.3 | 62.3 | 99.7 | 3.5
| 250 | 266.9 | 106.8 | 4.8 | 285.8 | 114.3 | 2.7 | 263.9 | 105.6 | 1.4

## Table 5
Analyte in brainstem_Analyte in brainstem | Nominal conc. (ng/g)_Nominal conc. (ng/g) | Intraday (n = 3)_Mean (ng/g) | Intraday (n = 3)_Accuracy (%) | Intraday (n = 3)_Precision (CV%)
Phenytoin | 46.9 | 48.9 | 104.3 | 6
Phenytoin | 187.5 | 197.8 | 105.5 | 5.8
Phenytoin | 750 | 739.5 | 98.6 | 1.6
4‐HPPH | 46.9 | 50.2 | 107 | 4.2
4‐HPPH | 187.5 | 200.8 | 107.1 | 4.6
4‐HPPH | 750 | 744.4 | 99.3 | 4.1

## Table 6
Matrix_Matrix | Concentration (ng/mL)_Concentration (ng/mL) | % Recovery_Phenytoin | % Recovery_4‐HPPH | Matrix effects (%)_Phenytoin | Matrix effects (%)_4‐HPPH
Plasma | 15.6 | 100.9 ± 0.5 | 99.2 ± 1.4 | −2.1 ± 1.5 | −1.9 ± 1.1
Plasma | 62.5 | 99.0 ± 0.7 | 97.9 ± 1.3 | −2.0 ± 1.1 | −1.7 ± 0.3
Plasma | 250 | 104.2 ± 3.3 | 99.8 ± 0.3 | −7.8 ± 2.6 | −3.3 ± 0.6
Brain | 46.9 | 101.7 ± 3.4 | 98.5 ± 2.5 | −5.5 ± 2.6 | −3.1 ± 2.2
Brain | 187.5 | 103.7 ± 3.1 | 100.5 ± 2.0 | −4.3 ± 0.7 | −3.5 ± 0.9
Brain | 750 | 110.6 ± 1.4 | 104.1 ± 1.5 | −7.8 ± 2.3 | −7.4 ± 0.8

## Freetext Data
Data | Value | Relevance
--- | --- | ---
Calibration range (plasma) | 7.81-250 ng/mL | Relevant for establishing the sensitivity and quantification limits of the developed LC-MS/MS method in plasma samples.
Calibration range (brain) | 23.4-750 ng/g | Relevant for establishing the sensitivity and quantification limits of the LC-MS/MS method in brain tissue.
Major metabolite | 5‐(4‐hydroxyphenyl)‐5‐phenylhydantoin (4‐HPPH) | Relevant because the metabolite is crucial for investigating phenytoin's metabolism and distribution.
Analyte detection method | LC-MS/MS | Relevant as the paper discusses the development and validation of this method for simultaneous quantification.
Mobile phase A | 0.1% formic acid in Milli Q water | Relevant for optimizing chromatographic separation during LC-MS/MS analysis.
Mobile phase B | 0.1% formic acid in 2:1 acetonitrile:methanol | Relevant for optimizing chromatographic separation and achieving sensitivity in the assay.
Lower Limit of Quantification (LLOQ) plasma | 7.81 ng/mL | Relevant for determining the lowest concentration detectable with acceptable accuracy and precision in plasma.
Lower Limit of Quantification (LLOQ) brain | 23.4 ng/g | Relevant for determining the lowest concentration detectable with acceptable precision in brain homogenates.
Extraction solvents | Acetonitrile and tert‐butyl methyl ether (TBME) | Relevant for the optimized sample preparation method enabling efficient analyte recovery.
Validation criteria | Specificity, sensitivity, accuracy, precision | Relevant for ensuring the analytical method meets scientific and experimental rigor.
MRM transitions for phenytoin | 253.011/182.100 | Relevant for the LC-MS/MS method's sensitivity and specificity in detecting phenytoin.
MRM transitions for 4‐HPPH | 269.051/198.100 | Relevant for ensuring accurate detection and quantification of the major metabolite.
Intravenous dose | 25 mg/kg | Relevant for evaluating phenytoin distribution and plasma/brain ratios in an in vivo preclinical study.
Intranasal dose | 20 µL of 100 mg/mL solution | Relevant for studying intranasal delivery and brain targeting efficacy.
Matrix recovery rate | >98% | Relevant to demonstrate the effectiveness of the sample preparation method.
Primary route of clearance | Hepatic metabolism via CYP2C9 and CYP2C19 | Relevant for understanding phenytoin's pharmacokinetics and metabolism pathways.
Plasma phenytoin concentration (IV) | 6.96 ± 1.07 µg/mL | Relevant for validating the LC-MS/MS method and quantifying pharmacokinetics post-intravenous administration.
Brain phenytoin concentration (IV) | 7.57 ± 1.09 µg/g | Relevant for assessing brain penetration and distribution after intravenous administration.
Ratio of phenytoin in brain/plasma (IV) | 1.10 ± 0.12 | Relevant for understanding drug distribution to the brain compared to plasma levels post-intravenous delivery.
Brain 4‐HPPH concentration (IV) | 0.21 ± 0.03 µg/g | Relevant for studying the limited brain penetration of the major metabolite.
Plasma phenytoin concentration (IN) | 203 ng/mL | Relevant to demonstrate the LC-MS/MS method's capability after low-dose intranasal delivery.
Brainstem concentration (IN, phenytoin) | 337 ng/g | Relevant for assessing regional brain distribution following intranasal delivery.
Plasma 4‐HPPH concentration (IN) | 424 ng/mL | Relevant for understanding the metabolite's systemic presence post-intranasal administration.
Brain to plasma ratio (4‐HPPH, IV) | 0.06 ± 0.01 | Relevant for evaluating the metabolite's barrier crossing and comparison with systemic exposure.
Brain to plasma ratio (4‐HPPH, IN) | 0.1 in all regions | Relevant to assess hydrophobicity/metabolism differentials in intranasal vs systemic routes.
Instrument details | Agilent 1290 HPLC and AbSciex QTRAP 5500 | Relevant for describing the analytical platform used to characterize phenytoin pharmacokinetics.
Blood samples collected post-mortem | At 65 minutes following drug administration | The time point for sampling was chosen to evaluate phenytoin and 4-HPPH concentrations after drug metabolism and distribution, ensuring data on the kinetic profile of the drug and its major metabolite were gathered.
Brain tissue sampling | Obtained at 65 minutes after drug administration, immediately post-decapitation | Brain sampling provided data on phenytoin and 4-HPPH distribution in the blood-brain barrier-relevant compartment, critical for assessing CNS pharmacokinetics.
Blood collected | In sodium heparin-coated tubes and centrifuged afterward | Ensures plasma extraction with minimal degradation or change in analyte concentration, essential for reliable quantification.
Brain regions sampled | Olfactory bulbs, main brain, brainstem | Specific brain regions were analyzed to explore direct and regional distribution differences, highly relevant for studying nose-to-brain delivery.
Number of rats for validation | Not explicitly stated in the excerpt | Number of animals impacts the power and reproducibility of PK studies; however, the exact count isn't provided.
Sampling from two administration routes | Intravenous and intranasal phenytoin solutions | Differentiating systemic exposure (intravenous) from potential localized (intranasal) delivery is crucial in the context of evaluating targeting efficiency.
Rat population | Male Wistar rats 300–380 g from specific pathogen-free conditions | Ensures consistency in metabolic and physiological responses to phenytoin across a controlled animal model, critical for reproducibility of PK outcomes.
LC-MS/MS Analytical Method | Validation and Application for Phenytoin Quantification | Relevant as the paper primarily focuses on validating a sensitive LC-MS/MS method, not pharmacokinetic modeling.
Seizure Model Validation | In vivo validation of phenytoin effects | Relevant as it demonstrates real-time pharmacodynamic assessments rather than relying on compartment models.
Intranasal Delivery Investigation | Using phenytoin solution to explore routes for drug delivery | Relevant to studying phenytoin distribution without explicitly engaging in compartment or mathematical modeling approaches.
Pharmacokinetic Profile | Saturable metabolism of phenytoin, dose-dependent elimination | Relevant for understanding the complexity in phenytoin’s elimination profile without indicating a compartment model usage.
Drug Distribution Analysis | Brain/plasma ratios of phenytoin and metabolite | Relevant for empirical tissue partitioning data, suitable for validating analytical methods, but does not imply compartment modeling.
Fraction Matrix Effect | Plasma and brain matrix studies analyzed | Relevant in determining drug/metabolite recovery within tissue types for method sensitivity, but no model category applied.
Phenytoin | 5,5‐diphenylhydantoin | Relevant as this is the primary drug analyzed.
Phenytoin Sodium Concentration for Stock Solution | 1 mg/mL dissolved in methanol | Relevant as this detail is foundational for preparing experimental stock solutions.
Calibration Curve Range Estimation, Rat Plasma | 7.81 to 250 ng/mL | Relevant for understanding the sensitivity and calibration parameters for detecting phenytoin in plasma.
Calibration Curve Range Estimation, Brain Tissue | 23.4 to 750 ng/g | Relevant for understanding the calibration curve sensitivity for brain tissue.
Mobile Phase A | 0.1% formic acid in Milli Q water | Relevant as a key parameter for chromatographic separation.
Mobile Phase B | 0.1% formic acid in 2:1 acetonitrile:methanol | Relevant as an optimization parameter for generating clear separation peaks.
Flow Rate | 0.25 mL/min | Relevant for the gradient chromatography conditions.
Column for LC-MS/MS | Kinetex EVO 5 μm 100 Å C18 (150 × 2.1 mm) | Relevant as the stationary phase affects analyte separation.
Autosampler Temperature | Maintained at 20°C | Relevant for sample integrity during injection.
Injection Volume | 5 μL | Relevant for achieving optimized loading of samples for MS analysis.
Gradient Mobile Phase Starting Ratio | 80% A, 20% B | Relevant to elution parameters.
Elution Peak Times | Phenytoin: 5.16 min | Relevant to identifying the retention time for the drug’s chromatographic peak.
Elution Peak Times | 4‐HPPH (metabolite): 4.06 min | Relevant to identifying the retention time for the primary metabolite chromatographic peak.
Precursors/Product Ion Pair for Phenytoin (MRM) | 253.011/182.100 | Relevant to ion monitoring optimization in mass spec detection.
Column Temperature | Maintained at 40°C | Relevant in improving chromatographic performance.
LLOQ in Plasma | 7.81 ng/mL | Relevant for the lower sensitivity limit of detection for plasma.
LLOQ in Brain Tissue | 23.4 ng/g | Relevant for the lower sensitivity limit of detection for brain tissue.
Analyte Recovery Rate in Plasma | >98% | Relevant regarding the method’s recovery efficiency for reliable quantification.
Analyte Recovery Rate in Brain Tissue | >98% | Relevant to assess extraction efficiency in complex tissues like the brain.
Accuracy and Precision Limit | ±15% | Relevant as this validates consistent reliability of the method developed.
Mass Spectrometer Type | AbSciex QTRAP 5500 | Relevant for LC-MS/MS hardware used in the development.
Analyte Support Buffer | 2 mL/g of Milli Q water for brain | Relevant for ensuring reproducibility in sample homogenization before analysis.
Metabolite Precursor/Product Ion Pair (MRM) | 269.051/198.100 for 4‐HPPH | Relevant for ensuring specificity in detecting the metabolite.
Autosampler Needle Cleaning Solvent | 90% methanol in water | Relevant for preventing sample cross-contamination.
Brain or Plasma Ratios | Phenytoin: 1.10± 0.12 | Relevant as the partition coefficient was estimated using validated assay.
Phenytoin's water solubility | Poor | Phenytoin's poor water solubility impacts its bioavailability and formulation challenges, which are critical factors in influencing its pharmacokinetics.
Phenytoin metabolism pathway | Saturable metabolism by CYP2C9 and CYP2C19 | Saturable metabolism leads to nonlinear and dose-dependent pharmacokinetics, significantly influencing drug clearance and steady-state plasma levels.
Major metabolite | 4‐HPPH (5‐(4‐hydroxyphenyl)‐5‐phenylhydantoin) | The conversion of phenytoin to 4‐HPPH is saturable, impacting drug elimination rates and interpatient variability in pharmacokinetics.
Drug transporter | Substrate for P‐glycoprotein (P‐gp) | P‐gp expresses in the brain capillary endothelial cells and limits phenytoin's central nervous system penetration by ejecting it back to the plasma, affecting brain bioavailability.
Therapeutic range elimination properties | Nonlinear dose-dependent elimination | Dose-dependent pharmacokinetics complicate achieving and maintaining therapeutic levels without overdosing.
Interpatient variability | Significant | Variability in phenytoin metabolism due to genetic and other factors strongly affects pharmacokinetic behavior across patients.
Matrix effects on analytes | Ion enhancement | Matrix effects, such as ion enhancement, may alter bioanalytical measurement accuracy, indirectly influencing observable pharmacokinetics.
Protein binding | Highly protein-bound drug | High protein binding capacity impacts the free (active) drug concentration, a crucial determinant of pharmacokinetics.
Formic acid vs. ammonium acetate | Improved ionization with formic acid | Ionization efficiency in LC-MS/MS analyses can influence accurate pharmacokinetic measurement outcomes for in vivo studies.
Nasal vs intravenous pathways | Possible nasal metabolism of phenytoin | Observed metabolite profiles after intranasal delivery suggest metabolism at the delivery site, influencing the pharmacokinetic profile relative to systemic administration.
Brain-to-plasma concentration ratio | Different for phenytoin and 4-HPPH | The variability in brain-to-plasma ratios suggests differential pharmacokinetics due to solubility and permeability properties.