## Table 1
0
Performing pharmacokinetic studies in paediatric patients is difficult due to several challenges, such as low blood volume and venipuncture-related pain. We successfully developed an innovative intravascular microdialysis catheter that can be inserted into small intravenous lines for paediatric patients, which overcomes these challenges.
This proof-of-concept study is encouraging and opens the way to further experiments leading towards future use of microdialysis in paediatric patients.

## Table 2
Unnamed: 0 | 67 IV Microdialysis CatheterRef. no.: 8050092 | Modified 67 IV Microdialysis Catheter
CE-certification | Yes | No
Membrane | Membrane | Membrane
Material | Polyarylethersulphone | Polyarylethersulphone
Total length | 30 mm | 60 mm
Length within bloodstream | 30 mm | 15 mm
Diameter | 0.6 mm | 0.6 mm
Shaft | Shaft | Shaft
Material | Polyurethane | Polyurethane
Length | 46 mm | 10 mm
Diameter | 0.9 mm | 0.9 mm
Inlet tube | Inlet tube | Inlet tube
Material | Polyurethane | Polyurethane
Length | 200 mm | 200 mm
Diameter | 1.0 mm | 1.0 mm
Outlet tube | Outlet tube | Outlet tube
Material | Polyurethane | Polyurethane
Length | 45 mm | 95 mm
Diameter | 1.0 mm | 1.0 mm

## Table 3
Unnamed: 0_level_0_Unnamed: 0_level_1 | Recovery/loss_Physiological saline | Recovery/loss_Human albumin
Forward dialysis | Forward dialysis | Forward dialysis
0–60 min | 27.8% | 48.4%
60–120 min | 28.3% | 54.2%
120–180 min | 29.7% | 55.6%
Mean ± SD: | 28.6% ± 1.0% | 52.7% ± 3.8%
Retrodialysis | Retrodialysis | Retrodialysis
0–60 min | 16.4% | 50.1%
60–120 min | 12.4% | 45.0%
120–180 min | 11.3% | 42.3%
Mean ± SD: | 13.4% ± 2.7% | 45.8% ± 4.0%

## Table 4
Characteristic | Unnamed: 1
Number of subjects | 10
Age (years) | 27 (25–28)
Weight (kg) | 77.6 (72–80)
Height (cm) | 178 (176–182)
eGFR (mL/min) | 120.5 (107–122)
Albumin (g/L) | 45.5 (45–47)
Total protein (g/L) | 71.7 (69–72)
Plasma protein binding (%) | 49.7 (45–52)

## Table 5
Loss rates (%)_Loss rates (%) | Subject ID_1 | Subject ID_2 | Subject ID_3 | Subject ID_4 | Subject ID_5 | Subject ID_6 | Subject ID_7 | Subject ID_8 | Subject ID_9 | Subject ID_10
MD probe 1 | 17 | 15.9 | 16 | 22.5 | 24.4 | 19.8 | 13.3 | 13.1 | 20.7 | 26.8
MD probe 2 | 16.9 | NAa | 17 | 21.4 | NAa | 19.7 | 20.4 | 18.6 | 15.5 | 16.4

## Table 6
Vancomycin PK parameters | MD | Plasma | p value
t 1/2 (h) | 10.6 ± 6.7 | 6.3 ± 0.8 | 0.022t
C max (mg/L) | 28.5 ± 7.8 | 28.2 ± 4.9 | 0.832W
AUClast (h⋅mg/L) | 87.82 ± 19.1 | 98.2 ± 9.1 | 0.130W
AUC∞ (h⋅mg/L) | 97.56 ± 22.0 | 103.9 ± 9.9 | 0.369W
Cl (L/h) | 10.9 ± 2.7 | 9.7 ± 0.95 | 0.089t

## Table 7
Parameter | Estimate | 95% CI | IIVa
CL [L/h] | 4.74 | 4.47–5.02 | 9.0 %CV
V1 [L] | 10.8 | 9.27–12.5 | 21.0 %CV
Q [L/h] | 4.95 | 4.2–5.8
V2 [L] | 19 | 17.8–20.2
Recovery [–] | 0.175 | 0.156–0.193 | 15.7 %CV
FUMD [–] | 0.48 | 0.40–0.59 | 27.7 %CV
FUUF [–] | 0.51 | 0.50–0.53
σ Plasma Prop. [%CV] | 13.9 | 11.3–17.9
σ UF Prop. [%CV] | 7.1 | 5.6–8.6
σ MD Prop. [%CV] | 34.7 | 30.2–39.2
σ Recovery Prop. [%CV] | 5.5 | 4.1–7.2

## Freetext Data
Data | Value | Relevance
--- | --- | ---
Mean Protein Binding for Vancomycin | Mean PPB of 49% ± 1.7% | Relevant for evaluating the unbound drug fraction, which is critical for pharmacokinetic analysis using microdialysis.
Plasma PK Model Type | Two-compartment model | Relevant as it describes the drug's plasma pharmacokinetics, necessary for the pharmacometric modeling and comparison with microdialysis data.
In vivo Probe Loss Rate (Mean) | 18.6% ± 3.7% | Relevant for understanding the microdialysis probe's performance in vivo and the adjustments needed for concentration calculations.
Correlation Between MD and Plasma Unbound Drug Concentrations | r = 0.783, p < 0.0001 | Relevant to assess if microdialysis accurately reflects unbound drug levels in plasma.
Agreement Mean Bias (MD vs. Plasma) | 0.19 mg/L | Relevant for the degree of difference between the two measurement methods for concentration values.
AUC0-24 MD vs. Plasma | MD: 87.8 ± 19.1 mg·h/L, Plasma: 98.2 ± 9.1 mg·h/L | Relevant to compare pharmacokinetic parameters derived from microdialysis and plasma sampling, informing methodological reliability.
Study Drug Administered | Single IV dose of 1000 mg vancomycin | Relevant because vancomycin serves as the model drug for evaluating microdialysis as a method for drug PK analysis.
Safety Related to Modified MD Catheter | No adverse effects due to catheter; one mild rash due to vancomycin | Relevant to assess the feasibility and safety of the catheter system for future pediatric applications.
Calibration Method Used | Retrodialysis technique | Relevant as it ensures accurate calibration of microdialysis probes for in vivo drug concentration measurements.
Flow Rate for Calibration in Vivo | 2.0 µL/min | Relevant because flow rate impacts probe recovery and concentration measurement accuracy in microdialysis.
Concentration Measurement Correlation Method | Bland-Altman and Pearson correlation | Relevant for validating microdialysis as a comparable method to plasma sampling.
MD Recovery in vitro | 52.7% ± 3.8% (human albumin setting) | Relevant because in vitro experiments demonstrated stability and feasibility of microdialysis probes in controlled settings.
Ethical Approvals and Patient Consent | Ethics Committee of Medical University of Vienna; written consent obtained | Relevant to ensure compliance with legal and ethical requirements for conducting clinical studies, especially when extending to pediatric populations.
Development and Design of Modified MD Catheters | Shaft shortened to 10 mm, membrane extended to 60 mm | Relevant because the catheter modifications are specific innovations enabling its use in smaller pediatric intravenous lines.
Limitations Identified for MD Probe Performance | Two MD probes excluded due to loss rates <10% | Relevant as it highlights variability and areas for improvement in device consistency for drug sampling.
Patient sample size | 10 healthy adult male volunteers | Indicates the number of participants involved in the study, crucial for understanding the study's scale and statistical robustness.
Blood sampling time points | Baseline, 30 min, 1, 2, 3, 4, 5, 6, 8, 12, and 24 h post-drug administration | Highlights the specific time intervals chosen to monitor drug concentration over time, relevant for understanding pharmacokinetic profiles.
Microdialysate sampling time intervals | 15–45 min, 45–75 min, 1.5–2.5 h, 2.5–3.5 h, 3.5–4.5 h, 4.5–5.5 h, 5.5–6.5 h, 7.5–8.5 h, 11.5–12.5 h, 23.5–24.5 h | Specifies the time intervals for drug concentration measurements using microdialysis, relevant for evaluating unbound drug profiles in real-time.
Patient population | Healthy adult volunteers with a median age of 27.3 ± 5.1 years and weight of 75.5 ± 6.4 kg | Describes the demographics of the study participants, relevant for understanding the generalizability of the PK findings.
Vancomycin dose administered | Single intravenous dose of 1000 mg over 60 minutes | Defines the drug regimen applied to participants, fundamental information for PK modeling.
Microdialysis probe insertion location | Two 20-gauge venous catheters in forearm veins | Explains where and how the MD probes were placed, critical for understanding tissue sampling methodology.
Calibration method | Retrodialysis technique with a 10 μg/mL calibration solution in human albumin 1% | Describes how the MD probes were calibrated in vivo, relevant for ensuring accuracy in drug concentration measurements.
Baseline sampling | Conducted prior to study drug administration | Ensures initial drug concentration measurements are zero, critical for PK calculations.
Final sampling | Both plasma and microdialysate samples collected at 24 h post-drug administration | Marks the endpoint of data collection, ensuring completion of the 24-hour concentration-time profile.
Study design | Prospective, open-label, single-center PK study | Indicates the methodological approach used, relevant for interpreting study reliability and bias control.
Plasma PK Model | Two-compartment | The plasma pharmacokinetics of vancomycin were best described by a two-compartment model, which allows for more precise characterization of vancomycin's distribution and clearance in humans.
Pharmacometric Modelling Integration | Two-compartment model integrated with calibration for microdialysis probes | Incorporating the microdialysis data into the two-compartment pharmacometric model allowed for effective validation of the novel sampling method, ensuring its potential relevance in pediatrics.
Model Parameter Validation | Model-based estimates for recovery were in high agreement with observed values | This validates the accuracy of the two-compartment model for predicting microdialysis-based recovery, reinforcing the reliability of the method.
Unbound Drug Fraction Representation | Constant fraction for both ultrafiltrate and microdialysis | The two-compartment model included a fixed fraction unbound, important for comparing ultrafiltration and microdialysis methods to ensure accurate representation.
Mean recovery (forward dialysis, Experiment A) | 28.6% ± 1.0% | Relevant to understanding in vitro recovery efficiency using physiological saline.
Mean loss (retrodialysis, Experiment A) | 13.4% ± 2.7% | Relevant to assessing drug loss during retrodialysis with physiological saline as perfusion.
Mean recovery (forward dialysis, Experiment B) | 52.7% ± 3.8% | Demonstrates improved recovery in vitro using human albumin 1% in perfusion solution.
Mean loss (retrodialysis, Experiment B) | 45.8% ± 4.0% | Indicates stability and reproducibility of loss rates with human albumin 1%.
Mean in vivo MD loss rate (Retrodialysis technique) | 18.6% ± 3.7% | Reflects in vivo calibration performance of the MD probes in the clinical study.
Plasma protein binding (PPB, 2.5 mg/L) | 46.1% ± 0.02% | Relevant to determining unbound drug fraction at low vancomycin concentrations.
Plasma protein binding (PPB, 50 mg/L) | 52.0% ± 0.02% | Crucial for understanding unbound fraction at high vancomycin levels.
Fraction unbound (MD estimate) | 0.48 (95% CI 0.40–0.59) | Comparison of unbound concentration estimations from MD in the pharmacometric model.
Fraction unbound (Ultrafiltration estimate) | 0.51 (95% CI 0.50–0.53) | Validates agreement between ultrafiltration and MD-derived unbound fractions.
CL (Clearance) | 8.95 L/h (95% CI not provided) | Estimated from the pharmacometric model for plasma PK.
V1 (Central volume of distribution) | 8.90 L (95% CI not provided) | Derived from the pharmacometric model as a key PK parameter.
V2 (Peripheral volume of distribution) | 23.8 L (95% CI not provided) | Indicates peripheral distribution compartment in a two-compartment model.
Q (Intercompartmental clearance) | 7.40 L/h (95% CI not provided) | Represents drug distribution rate between compartments in the pharmacometric model.
Recovery estimate of MD probes (Pharmacometric model) | 17.6% | Validates measured in vivo recovery rates against predictions using integral sampling.
Mean AUClast (Plasma) | 98.2 ± 9.1 mg·h/L | Represents total exposure from time 0 to 24 hours, calculated from plasma sampling.
Mean AUClast (MD) | 87.8 ± 19.1 mg·h/L | Represents total exposure from time 0 to 24 hours, calculated using MD, relevant for PK parameter comparison.
Mean Cmax (Plasma) | 56.5 ± 5.7 mg/L | Indicates the maximum concentration achieved in plasma.
Mean Cmax (MD) | 52.0 ± 16.3 mg/L | Reflects peak concentration measured in MD samples, used to assess agreement with plasma sampling.
Mean t1/2 (Plasma) | 5.02 ± 0.49 hours | Terminal half-life derived from unbound plasma concentrations using noncompartmental analysis.
Mean t1/2 (MD) | 5.57 ± 1.01 hours | Terminal half-life derived from unbound drug measured in MD samples, showing a slight statistical difference from plasma.
Physiological differences in paediatrics | Maturation-dependent differences in organ function, body weight, body water, and protein composition | Relevant because these physiological differences continuously change with growth in paediatrics, affecting PK.
Age-related ontogenetic changes | Below and above 2 years, paediatric ontogenetic changes are unique compared to adults | Relevant because immature organ function and protein binding differences influence drug metabolism and disposition.
Plasma protein binding | Mean protein binding of vancomycin is 49% with SD of 1.7% | Relevant as dynamic protein binding affects the unbound active drug fraction and PK interpretation.
Drug dosing practices in paediatrics | Weight-based down-scaling from adult doses | Relevant because this approach may not account for unique paediatric PK parameters, risking toxicity or underexposure.
Limited blood volume in paediatrics | Physiological limitation in neonates and infants | Relevant as sparse blood sampling hampers proper PK study design and sampling density.
Pain and feasibility | Venipuncture-related pain and practical burden in paediatric patients | Relevant because it hinders repeat sampling necessary for PK analysis.
Microdialysis variables | Recovery rates depend on perfusion solution (physiological saline vs. human albumin) | Relevant because MD probe performance directly influences PK measurement accuracy.
Flow rate in microdialysis | Variability in recovery increases with higher flow rates | Relevant to optimize MD protocols for precise PK measurements.
Intra-study variability | Variability among MD probes and patients; low loss rates in vivo compared to in vitro | Relevant for understanding discrepancies and enhancing reproducibility in PK.
Unbound drug concentration dynamics | Observed differences between active (unbound) and total (bound) drug concentrations | Relevant for focusing on the pharmacologically active component in PK studies.
Vancomycin antimicrobial mechanism | AUC/MIC primarily drives efficacy | Relevant as PK/PD relationships depend on accurately measured PK parameters.
Variable plasma protein binding between concentrations | PPB of vancomycin at 2.5 mg/L: 46.1%, at 50 mg/L: 52% | Relevant for understanding concentration-dependent binding as a PK-altering factor.
Limitations of predictive models in paediatrics | Assumptions and sparse data in pharmacometric models | Relevant because it shows limits of current PK prediction methods in children.