## Table 1
Analyte | Precursor ion(m/z) | Product ion(m/z) | Collision energy(eV) | Dwell time(msec) | Retention time(min)
MPA | 321.10 321.10 | 275.00 285.00 | − 15 − 20 | 28 28 | 5.58
MPA-d3(internal standard) | 324.00 341.00 | 210.00 210.00 | − 22 − 35 | 28 28 | 5.58
MPAG(only monitored) | 514.21 514.21 | 207.04 303.13 | − 38 − 18 | 28 28 | 5.2

## Table 2
Statistic_Statistic | Before conversion (cross-validation)_MPAVAMS-LC–MS/MS vs MPAWB-LC–MS/MS | Before conversion (cross-validation)_MPAVAMS-LC–MS/MS vs MPAPL-LC–MS/MS | Before conversion (cross-validation)_MPAPL-HPLC–DAD vs MPAPL-LC–MS/MS | Before conversion (cross-validation)_MPAVAMS-LC–MS/MS vs MPAPL-HPLC–DAD | After conversion (clinical validation)_MPAVAMSHt-LC–MS/MS vs MPAPL-LC–MS/MS | After conversion (clinical validation)_MPAVAMSc-LC–MS/MS vs MPAPL-LC–MS/MS | After conversion (clinical validation)_MPAVAMSc_DAD-LC–MS/MS vs MPAPL-HPLC–DAD
Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula)
Passing-Bablok-regression formula | CVAMS = 0.90·CWB + 0.17 | CPL = 1.60·CVAMS + 0.08 | CHPLC = 1.01·CPL + 0.11 | CHPLC = 1.56·CPL + 0.17 | CPL = CVAMSHt — 0.13 | CPL = CVAMSc — 0.01 | CVAMSc_DAD = 1.01·CHPLC —0.02
Intercept (A) | 0.17 (− 0.06 to 0.29) | 0.08 (− 0.14 to 0.29) | 0.11 (− 0.24 to 0.30) | 0.17 (− 0.22 to 0.38) | − 0.13 (− 0.35 to 0.03) | − 0.01 (− 0.27 to 0.19) | − 0.03 (− 0.35 to 0.31)
Slope (B) | 0.90 (0.79–1.00) | 1.60 (1.44–1.80) | 1.01 (0.91–1.13) | 1.56 (1.38–1.78) | 1.00 (0.92–1.09) | 1.00 (0.90–1.11) | 1.01 (0.89–1.15)
Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias)
%-mean difference bias (− 1.96SD; + 1.96SD) | − 6.80% (− 61.50 to 47.90%) | 54.07% (48.62–59.53%) | + 4.08% (− 1.21 to 9.37%) | 50.00% (1.60–98.50%) | + 10.90% (5.24–16.56%) | 4.76% (− 1.11 to 10.64%) | − 1.01% (− 8.26 to 6.24%)
% of paired samples fulfilled the EMA criteria (< 20%) | 70% | 0% | 12% | 12% | 88% | 86% | 72%
% of paired samples fulfilled the IATDMCT criteria (< 15%) | 60% | 0% | 6% | 4% | 78% | 72% | 67%
Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors
MPE | − 0.05 | − 0.92 | − 0.12 | − 0.81 | − 0.12 | − 0.01 | 0.00
MPPE [%] | − 5.50 | 85.00 | 1.90 | 67.54 | 7.03 | 9.20 | 2.60
RMSE [%] | 15.00 | 92.00 | 12.00 | 81.54 | 12.14 | 12.21 | 13.00
MAPE [%] | 13.69 | 66.65 | 8.25 | 62.42 | 14.02 | 13.38 | 20.24
Correlation coefficients | Correlation coefficients | Correlation coefficients | Correlation coefficients | Correlation coefficients | Correlation coefficients | Correlation coefficients | Correlation coefficients
Pearson’s (R 2) | 0.92 (0.87–0.96; 95% CI) | 0.92 (0.86–0.95; 95% CI) | 0.93 (0.88–0.96; 95% CI) | 0.90 (0.84–0.95; 95% CI) | 0.95 (0.85–0.97; 95% CI) | 0.92 (0.87–0.96; 95% CI) | 0.91 (0.85–0.95; 95% CI)
Spearman (SRCC) | 0.81 (0.68–0.89; 95% CI) | 0.88 (0.80–0.93; 95% CI) | 0.92 (0.86–0.95; 95% CI) | 0.83 (0.73–0.90; 95% CI) | 0.91 (0.84–0.95; 95% CI) | 0.88 (0.80–0.93; 95% CI) | 0.84 (0.73–0.90; 95% CI)
Intraclass (ICC) | 0.91 (0.85–0.95; 95% CI) | 0.84 (0.73–0.90; 95% CI) | 0.96 (0.93–0.98; 95% CI) | 0.83 (0.74–0.91; 95% CI) | 0.94 (0.90–0.97; 95% CI) | 0.92 (0.87–0.95; 95% CI) | 0.91 (0.85–0.95; 95% CI)

## Freetext Data
Data | Value | Relevance
Patients included in the study | 50 pediatric patients | Relevant for understanding the study population, focusing on pediatric renal transplant recipients treated post-transplantation.
Matrix types for MPA determination | Whole blood (WB), plasma (PL), and VAMS | Relevant for assessing the development and validation of methods specific to different biological matrices.
Therapeutic drug monitoring method developed | LC–MS/MS | Relevant as it highlights the method applicable for precise MPA quantification and monitoring in the study population.
Calibration range for MPA determination | 0.10–15 µg/mL | Relevant for assessing the sensitivity and applicability of the developed LC–MS/MS method.
Validation adherence | EMA and IATDMCT guidelines | Relevant for ensuring the study complies with international bioanalytical method validation standards.
Cross-validation methods compared | WB-LC–MS/MS, PL-LC–MS/MS, VAMS-LC–MS/MS, and HPLC–DAD | Relevant for confirming the interchangeability and accuracy of the methods used.
Regression formula for VAMS-to-plasma MPA concentration | y = 1.60x + 0.08 | Relevant for enabling the conversion of VAMS sample results into plasma-equivalent concentrations for clinical interpretation.
Stability of VAMS-loaded samples | Up to 6 weeks at RT in the dark | Relevant for ensuring accurate MPA concentration measurement over time for practical storage and transportation conditions.
Clinical validation highlights | Formula successfully validated | Relevant for confirming that the VAMS sampling approach can reliably replace venipuncture methods for MPA monitoring post-transplantation.
Obtained methods compatibility | Satisfactory equivalence between LC–MS/MS and HPLC–DAD | Relevant for supporting the transition from traditional methods to LC–MS/MS in routine TDM practices.
Patient sample size | 50 pediatric renal transplant recipients | Critical for assessing the study population and validity for the PK analysis in pediatric post-transplant settings.
Sampling method | WB and PL samples via venipuncture, VAMS samples via fingerstick | Relevant for understanding the study design and the sampling techniques used for PK analysis.
Sampling time points | Samples obtained before the next daily dose of immunosuppressive agents | Defines trough concentration measurement timing essential for PK evaluation.
Clinical setting | Children’s Memorial Health Institute, Warsaw | Indicates the study site, relevant for study context and institutional expertise in pediatric transplantation.
Study duration | Sampling conducted between November 2022 and February 2023 | Important to contextualize the temporal scope of patient monitoring for PK evaluation.
The model used for the study | Regression model | The regression model was developed and validated for converting volumetric absorptive microsampling (VAMS) mycophenolic acid (MPA) concentrations to plasma (PL) concentrations, which is a critical component for the interpretation of therapeutic drug monitoring (TDM).
Alternative model explored | Hematocrit-based model | The hematocrit-dependent formula was evaluated for converting VAMSMPA to PLMPA concentrations, ensuring methodological accuracy for drugs primarily binding with plasma proteins and providing a comparative baseline against the regression model.
Cross-validation process for model | Passing-Bablok regression | Passing-Bablok regression was employed to evaluate equivalency between methods and was instrumental in establishing the new regression model for conversion.
Clinical validation performance of model | Conversion formula based on regression | The Passing-Bablok regression-based model showed successful clinical validation, confirming the reliability and clinical applicability of the regression approach for MPA concentration conversions.
Regression application | y = 1.60x + 0.08 for MPA concentration | This formula provided a direct conversion of MPA concentration in VAMS to plasma levels, facilitating comparison between matrices and supporting individualized therapeutic monitoring strategies.
Retention Time (RT) | 5–7 minutes | Relevant as it describes the time intervals for analyte elution during chromatographic separation, critical for method validation and identification.
Quantitative MRM pair (MPA) | 321.10 > 275.00, −15 collision energy | Relevant because MRM transition was used for precise quantitation of mycophenolic acid, essential in LC-MS/MS validation.
Qualitative MRM pair (MPA) | 321.10 > 285.00, −20 collision energy | Relevant to confirm analyte identity and enhance method specificity.
Quantitative MRM pair (MPA-d3) | 324.00 > 210.00, −22 collision energy | Relevant for internal standard determination, ensuring accuracy in MPA quantification.
Qualitative MRM pair (MPA-d3) | 324.00 > 210.00, −35 collision energy | Ensures internal standard specificity, further validating method precision.
Interface Temperature | 150 °C | Relevant since optimized detector conditions influence ionization efficiency, critical for accurate quantitation.
Desolvation Line (DL) Temperature | 250 °C | Impacts ion transport and is crucial for maintaining method robustness.
Heat Block (HB) Temperature | 400 °C | Relevant for analyte desolvation and ion formation during the mass spectrometry process.
Drying Gas Flow Rate | 5 L/min | Relevant for efficient desolvation, a necessary parameter for spectrometric analysis.
Nebulizing Gas Flow Rate | 3 L/min | Crucial for stable spray formation in electrospray ionization, affecting method reproducibility.
Heating Gas Flow Rate | 14 L/min | Impacts ion evaporation processes, contributing to method accuracy.
Ammonium Fluoride Concentration in Mobile Phase | 2.5 mM | Relevant for enhancing ion formation, facilitating chromatographic separation and quantification.
Formic Acid Addition in Mobile Phase | 0.05% | Enhances protonation of analytes, critical for consistent ionization in mass spectrometry.
Mobile Phase Composition | Pure water (Phase A), methanol (Phase B) | Relevant for achieving optimal separation and sensitivity during LC-MS/MS analysis.
Flow Rate | 0.50 mL/min | Influences resolution and retention, optimizing chromatographic performance for MPA.
Injection Volume | 10 µL | Ensures reproducibility and uniform sample delivery, critical in LC-MS/MS method validation.
Autosampler Cooler Temperature | 5 °C | Ensures sample stability during analysis, improving method reliability.
Calibration Range | 0.10–15.0 µg/mL | Relevant for establishing analytical quantitation limits and supporting method validation.
Correlation Coefficient (R2) | 0.9984 ± 0.0013 for WB, 0.9994 ± 0.0013 for PL | Indicates method linearity, a fundamental validation parameter.
Slope and Intercept of Calibration Curve | WB: 0.9821x + 0.011, PL: 1.004x − 0.037 | Relevant for evaluating quantitative accuracy and calibration reliability.
Autosampler Stability Tested Duration | 1, 3, and 5 days | Determines sample stability under storage conditions, ensuring method robustness.
Loaded VAMS Stability Duration | 6 weeks (RT & 4 °C, dark), 8 weeks (−20 °C) | Relevant for evaluating usability and storage conditions for VAMS in real-world applications.
Hematocrit Effect (HT) | Pearson’s r = −0.02596 (−0.31 to 0.27, 95% CI) | Indicates negligible influence of HT, validating VAMS method appropriateness for MPA determination.
Matrix Effect (ME) | < ±15% for WB, PL, and VAMS matrices | Relevant for confirming the lack of interference, supporting method precision.
Absolute Recovery (AR) | Calculated across WB, PL, and VAMS samples | Provides insight into extraction efficiency, critical for MPA quantification reliability.
Effect of Food-Drug Interactions | Pharmacokinetics of MPA can be influenced by food-drug interactions. | This factor impacts the absorption and efficacy of mycophenolic acid when administered, requiring monitoring in pediatric post-transplant care.
Inter- and Intra-individual Variability | MPA has high inter-individual and intra-individual variability among patients. | Variability affects the dosing and therapeutic drug monitoring required for personalized care.
Age-specific Effects | Children younger than 6 years exhibit potential adverse effects/toxicities with MPA. | Developmental pharmacokinetics may require careful monitoring for this age group to mitigate toxicity.
Binding to Plasma Proteins | MPA exclusively binds with plasma proteins, affecting its concentration in VAMS capillary blood. | This binding characteristic must be accounted for in sample analysis and drug level conversion protocols.
Hematocrit Level | Hematocrit influence was assessed but no correlation was found in MPA concentration differences from VAMS samples. | Although found negligible, hematocrit levels are typically considered in studies as they can alter drug sampling outcomes.
Storage Stability | Stability of MPA in VAMS was dependent on storage conditions, impacting long-term sample integrity. | Pharmacokinetics can be influenced by sample degradation, affecting reliability of therapeutic drug monitoring.
Limited Plasma Availability | MPA dilution in whole blood compared to plasma alters observed concentrations. | Reflects sampling matrix-specific pharmacokinetics, critical for accurate drug monitoring methods.