## Table 1
Item | N | Minimum | Maximum | Mean | Std. Deviation
Kerma Average (mGy) | 25 | 3.93 | 17.53 | 8.87 | 4.42
HVL average (mmAl) | 25 | 0.33 | 0.68 | 0.49 | 0.13
Calculated MGD (mGy) | 25 | 0.95 | 4.1 | 2.31 | 1.02

## Table 2
Level MGD(mGy) | Number of facilities | Percentage (%) | Minimum MGD(mGy) | Maximum MGD(mGy)
> 3.0 | 8 | 32 | 3.03 | 4.10
<3.0 | 17 | 68 | 0.95 | 2.75
Total | 25 | 100 | – | –

## Table 3
Equipment parameters | EMGD | KVP | HVL | mGy_mAs
MGD | 1.000 | −0.017 | −0.088 | 0.259
KVP | −0.017 | 1.000 | 0.263 | 0.162
HVL | −0.088 | 0.263 | 1.000 | 0.384
mGy_mas | 0.259 | 0.162 | 0.384 | 1

## Table 4
Unnamed: 0 | Unnamed: 1 | KVP | HVL | mGy per mAs | MGD
KVP | Correlation Coefficient | 1.000 | 0.286 | 0.313 | −0.057
KVP | Sig. (2-tailed) |  | 0.166 | 0.128 | 0.788
HVL | Correlation Coefficient | 0.286 | 1.000 | 0.609** | −0.049
HVL | Sig. (2-tailed) | 0.166 |  | 0.001 | 0.815
mGy per mAs | Correlation Coefficient | 0.313 | 0.609** | 1.000 | −0.087
mGy per mAs | Sig. (2-tailed) | 0.128 | 0.001 |  | 0.678
MGD | Correlation Coefficient | −0.057 | −0.049 | −0.087 | 1.000
MGD | Sig. (2-tailed) | 0.788 | 0.815 | 0.678

## Table 5
Authors | Study Type | Results (MGD) | Methodology | Model | Sample Size | Country
Alghamdi et al 202419 | Retrospective | 1.2–2.5 mGy | Examined the radiation dose in mammography across various hospitals in Saudi Arabia. | Patient | 510 | Saudi Arabia
Hegian et al 202345 |  | 1.6 to 1.7 mGy | Estimate the MGD-CBT relationship based on patient age in Jordan |  | 3465 | Jordan
Asadollahzadeh et al 202446 | RetrospectiveSingle center | 1.0 and 2.8 mGy | Correlate MGD with breast tissue density and assess the variability of dose in relation to different breast thicknesses. | Patient | 400 | Iran
Suliman et al 202314 | Retrospective | 0.56 to 2.89 mGy for CC and 0.48 to 2.08 mGy for MLO, with averages of 1.36 mGy and 1.19 mGy, respectively | Measures the (MGD) from mammography in 496 breasts across eight clinics in Sudan | Patient | 496 | Sudan
Abdulwahid et al 202342 | Retrospective | 0.80 mGy and 0.82 mGy for the craniocaudal (CC) projection and between 0.89 mGy and 0.971.8 mGy for the mediolateral oblique (MLO) projection | Determine local (DRLs) for mammography procedures in Dubai at different ranges of breast thickness | Patient | 2599 | UAE
Al naemi, H. et al 201641 | Retrospective | The average MGD of 1.885 mGy and 1.84 mGy for the MLO and CC193 views respectively | To measure the (MGD) from CC and MLO views from mammography patients in (HMC) in Qatar and to establish (DRL) for the country by applying the (QC) protocol for the Digital mammography units. | Patient | 4085 | Qatar
Parmaksız et al 202043 | Retrospective | 1.1–1.5 and 1.6–2.0 mGy for CC and MLO projections | Establish national diagnostic reference levels for mammography examinations | Patient | 25624 | Turkey
Xu et al 202444 | Prospective | 1.67 mGy for CC views and 1.71 mGy for MLO views | Determine the national diagnostic reference levels for digital mammography | Patient | 4769 | China
Current Study 2023 | Perspectivemulti-center | 0.95–4.10 mGy, mean: 2.31 mGy, SD: 1.02 mGy, 32% centers exceeded 3.0 mGy | Analysis of mammography centers, compliance with ACR standards | Phantom | 25 mammography centers | Jordan

## Table 6
MGD | Accredited | Non- accredited | Total
> 3.0mGy | - | 8 | 8
<3.0mGy | 9 | 8 | 17
Total | 9 | 16 | 25

## Table 7
Accreditation status | Mean (mGy) | Number | Std. Deviation | Minimum(mGy) | Maximum(mGy)
Accredited | 1.6 | 9 | 0.42 | 0.95 | 2.11
Non-accredited | 2.7 | 16 | 1.02 | 1.15 | 4.1
Total | 2.3 | 25 | 1.02 | 0.95 | 4.1

## Table 8
Level MGD | MoH | University | Private | Total
> 3.0 (mGy) | 3 | – | 5 | 8
≤3 (mGy) | 3 | 4 | 10 | 17
Total | 6 | 4 | 15 | 25

## Table 9
Health care sector | Percent | Mean (mGy) | Std. Deviation | Minimum (mGy) | Maximum (mGy)
MoH | 24% | 2.84 | 1.07 | 0.99 | 3.79
University | 16% | 2 | 0.54 | 1.59 | 2.75
Private | 60% | 2.18 | 1.08 | 0.95 | 4.1
Total | 100% | 2.31 | 1.02 | 0.95 | 4.1

## Freetext Data
Data | Value | Relevance
Mammography units evaluated | 25 | Relevant as it outlines the sample size and scope for assessing MGD in mammography units in Jordan.
Average MGD | 2.3 mGy | Relevant to determine if mammography units in Jordan are within the ACR’s recommended dose limit.
MGD range | 0.95-4.10 mGy | Critical for identifying the variability and instances where MGD exceeds ACR limits, impacting patient safety.
Percentage of units exceeding MGD limit | 33% | Relevant as it highlights non-compliance with safety standards, underscoring the need for improvement in radiation dose management.
MGD in non-accredited facilities | 2.7 mGy | Relevant as it demonstrates the disparity in quality control between accredited and non-accredited centers.
MGD in accredited facilities | 1.6 mGy | Relevant to showcase the effectiveness of accreditation in adhering to radiation safety standards.
Accreditation rate among participants | 36% | Relevant as it indicates the proportion of facilities meeting established quality and safety benchmarks.
Primary mammography technology | Full Field Digital Mammography (44.4%) | Relevant to understand the prevailing technology and its influence on MGD values in Jordan.
Mo/Mo target-filter combination prevalence | 64% | Relevant as target/filter combinations influence radiation exposure and MGD results.
Fully accredited mammography center performance | MGDs below 3.0 mGy | Relevant to show the efficacy of accreditation in maintaining safety and adherence to guidelines.
Non-accredited center failure rate | 50% exceeding 3.0 mGy | Relevant to highlight the shortcomings in QA and radiation exposure management in non-accredited facilities.
HVL influence on MGD | Weak negative correlation (r = -0.088) | Relevant in assessing the technical factors affecting dose levels in mammography practices.
Private healthcare sector failure rate | 30% (5 out of 15 machines) exceeding ACR guidelines | Relevant to highlight variability in compliance across sectors, focusing on QA gaps in private clinics.
MGD in Ministry of Health (MoH) centers | 2.84 mGy average, 50% failure rate | Relevant to identify sectors with the most significant non-compliance, driving targeted improvements in MGD control.
University hospital MGD average | 2.0 mGy | Relevant as it demonstrates effective quality management and compliance in this sector.
Significant positive correlation | HVL and mGy per mAs (ρ = 0.609) | Relevant as it underscores the interdependence of equipment parameters critical to dose optimization.
Primary accreditation authority | HCAC (Health Care Accreditation Council) | Relevant as it contextualizes who ensures compliance with mammography standards in Jordan.
Funding strategy recommendation for accreditation | Financial incentives and subsidies | Relevant as proposed solutions to increase facility compliance with accreditation standards, improving overall radiation safety.
Comparison to international standards | Jordan's MGD aligns with broader trends but has areas for improvement | Relevant for situating Jordan's results in a global context to highlight strengths and challenges.
Number of mammography facilities targeted | 82 | Relevant as it sets the initial scope of the study and explains the selection process for mammography units.
Response rate | 30.5% leading to 25 participating sites | Demonstrates the achieved sample size and aligns with the question concerning patient population and inclusion criteria.
Patient population design | Random sampling of mammography units, including both accredited and non-accredited facilities | This indicates the study's approach to encompassing a representative population, critical for interpreting results.
Mammography unit technologies evaluated | CR, FFDM, and DBT, and their combinations | Relevant as it shows variations in equipment, impacting potential differences in radiation exposure.
Sampling tools | ACR phantom, RaySafe X2 MAM dosimeter | These devices were used to replicate clinical settings for dose measurements, addressing the question about sampling methodology.
Measurements conducted | Entrance kerma and HVL measurements | Relates to blood/tissue sampling analog in mammography and how radiation doses were assessed.
Phantom type and specifications | ACR-Accredited Full-Field Phantom simulating a 4.2 cm compressed breast | Shows the control breast tissue/tissue-equivalent sampling conditions aligned with the study goals.
Dosimetry methodology | Based on the ACR’s 2018 Digital Mammography QC Manual | Relevant as it demonstrates the standardized protocol for gathering dosimetry data from the sampled units.
Statistical analysis tools | IBM SPSS Statistics 27, descriptive/inferential statistical methods | Shows how sample data was analyzed to derive conclusions about the patient population's safety parameters.
Approval for sampling and ethical considerations | IRB approval (No. 43/160/2023), ethical compliance guaranteed | Relevant for confirming ethical standards followed during patient-equivalent sampling.
Accreditation status impact | 33% exceeded MGD limits, especially in non-accredited centers | Shows the relationship between sampling locations (accredited/non-accredited) and findings.
Phantom Model | ACR-Accredited Full-Field Phantom | The phantom serves as a standardized model to simulate a compressed breast for evaluating the performance of digital mammography systems.
Composition | 50% glandular and 50% adipose tissue | The phantom's composition mimics the anatomical structure used in dose calculations, making it analogous to breast tissue for better dose estimation.
Dimensions | 311.2×190.5×41.3 mm3 | The phantom covers the entire image detector area and minimizes scattering, ensuring accurate simulation for MGD calculation.
Thickness | 42 mm | Represents a compressed breast thickness, critical for applying dosimetry models in mammography.
Breast Thickness | 4.2 cm | Relevant to MGD calculation, as breast thickness directly influences radiation dose measurement.
Glandular Tissue to Fat Ratio | 50:50 | Foundational for MGD calculations to ensure measurement consistency with phantom testing.
Incident Air Kerma (K) Range | 3.93 mGy to 17.53 mGy | Essential for MGD estimation as it reflects radiation exposure levels at the phantom’s surface.
Mean Glandular Dose (MGD) Range | 0.95 mGy to 4.10 mGy | Critical for assessing compliance with the American College of Radiologists' standards.
Average MGD | 2.31 mGy | Central value for evaluating adherence to dose guidelines.
Half-Value Layer (HVL) | Provided by RaySafe X2 dosimeter | Necessary for determining X-ray beam quality and calculating MGD via conversion factors.
Voltage (kVp) | 28 kVp | Represents operating conditions of mammography units during data collection, linking to dose measurement.
RaySafe X2 Dosimeter Range | 1 μGy to 99.99 Gy | Ensures the dosimeter's capacity for accurate air kerma and HVL readings relevant to MGD estimation.
Measured Equipment Parameters | kVp, HVL, mGy/mAs | These are factors contributing to MGD computation and optimization.
Conversion Factors | Derived from Dance et al (Monte Carlo simulations) | Determines MGD estimation accuracy by relating air kerma to glandular dose using phantoms.
Phantom Type | ACR-Accredited Full-Field Phantom | Used to simulate breast structure during mammography and link measured values to MGD.
Equipment Variants | CR, FFDM, DBT | Reflect technology variations in mammography units that affect dose and imaging protocol.
Target/Filter Combos | Mo/Mo, W/Rh, Rh/Mo | Influences X-ray spectra and thus dose calculation for breast tissue.
Calibration Procedure | Conducted at Radiation Application Lab, Royal Scientific Society | Ensures the dosimeter's reliability for accurate HVL and MGD estimations.
Calculation Method | ACR Digital Mammography QC Manual (2018) | A standardized guideline critical for ensuring consistent MGD measurement methods.
Breast thickness | Breast thickness influences the accuracy of MGD measurement. | Variability in breast thickness is a factor that could influence how radiation is absorbed and distributed, indirectly impacting MGD and reflecting on tissue kinetics in imaging.
Imaging protocols | Variability in imaging protocols affects dose levels. | Different protocols dictate exposure settings, which can affect how radiation interacts with breast tissue, influencing derived dosage metrics.
Equipment characteristics | Variations in machine technologies and specifications influence MGD. | Equipment-specific variables, such as target/filter materials, beam qualities, and calibration settings, can alter the radiation dose and thereby affect its pharmacokinetic behavior in tissue.
Tube efficiency | Lower mGy per mAs values lead to reduced MGD. | Efficient X-ray production (linked to beam generation and filtration) correlates with optimized dosage, which is analogous to pharmacokinetic efficiency in drug administration.
HVL (Half-Value Layer) | Weak negative correlation between MGD and HVL. | HVL represents the penetrating quality of X-ray beams, correlating with dose absorption levels, akin to physiological absorption factors in pharmacokinetics.
Patient-specific anatomy | Breast composition affects dose distribution. | Variability in tissue density (glandular vs. adipose tissue) affects MGD and is comparable to patient-specific differences in drug absorption and metabolism.