## Table 1
Unnamed: 0_level_0_Unnamed: 0_level_1_Unnamed: 0_level_2 | Unnamed: 1_level_0_Unnamed: 1_level_1_GLY Conc. (ng/mL) | MET Conc. (ng/mL)_LLOQ (50)_Degree of Signal Suppression of GLY by MET (%) | MET Conc. (ng/mL)_LQC (150)_Degree of Signal Suppression of GLY by MET (%) | MET Conc. (ng/mL)_MQC (600)_Degree of Signal Suppression of GLY by MET (%) | MET Conc. (ng/mL)_HQC (2400)_Degree of Signal Suppression of GLY by MET (%) | MET Conc. (ng/mL)_ULOQ (3200)_Degree of Signal Suppression of GLY by MET (%) | MET Conc. (ng/mL)_LLOQ (50)_Degree of Signal Suppression of MET by GLY (%) | MET Conc. (ng/mL)_LQC (150)_Degree of Signal Suppression of MET by GLY (%) | MET Conc. (ng/mL)_MQC (600)_Degree of Signal Suppression of MET by GLY (%) | MET Conc. (ng/mL)_HQC (2400)_Degree of Signal Suppression of MET by GLY (%) | MET Conc. (ng/mL)_ULOQ (3200)_Degree of Signal Suppression of MET by GLY (%)
Nominal concentration | LLOQ (20) | 87 | 88 | 74 | 72 | 66 | 93 | 99 | 92 | 111 | 113
Nominal concentration | LQC (60) | 89 | 86 | 74 | 73 | 72 | 94 | 90 | 91 | 101 | 100
Nominal concentration | MQC (240) | 93 | 91 | 80 | 77 | 71 | 106 | 105 | 103 | 95 | 100
Nominal concentration | HQC (960) | 87 | 90 | 81 | 71 | 69 | 106 | 95 | 108 | 88 | 110
Nominal concentration | ULOQ (1280) | 92 | 90 | 79 | 74 | 73 | 113 | 98 | 106 | 89 | 111
10-fold dilution for MET and GLY | LLOQ (20) | 106 | 91 | 101 | 76 | 73 | 89 | 111 | 107 | 114 | 85
10-fold dilution for MET and GLY | LQC (60) | 102 | 106 | 103 | 82 | 81 | 109 | 105 | 91 | 115 | 87
10-fold dilution for MET and GLY | MQC (240) | 105 | 93 | 105 | 88 | 89 | 90 | 103 | 97 | 102 | 90
10-fold dilution for MET and GLY | HQC (960) | 96 | 95 | 96 | 92 | 89 | 95 | 111 | 99 | 95 | 94
10-fold dilution for MET and GLY | ULOQ (1280) | 89 | 92 | 92 | 88 | 93 | 93 | 110 | 106 | 96 | 104
20-fold dilution for MET and GLY | LLOQ (20) | 107 | 109 | 91 | 71 | 75 | 91 | 110 | 95 | 93 | 100
20-fold dilution for MET and GLY | LQC (60) | 114 | 101 | 102 | 82 | 92 | 110 | 95 | 92 | 86 | 105
20-fold dilution for MET and GLY | MQC (240) | 102 | 97 | 99 | 96 | 87 | 99 | 107 | 89 | 93 | 99
20-fold dilution for MET and GLY | HQC (960) | 87 | 101 | 108 | 92 | 91 | 106 | 114 | 98 | 92 | 94
20-fold dilution for MET and GLY | ULOQ (1280) | 89 | 106 | 87 | 90 | 88 | 99 | 113 | 89 | 86 | 94
Plasma samples for MET and GLY | LLOQ (20) | 88 | 82 | 76 | 79 | 77 | 94 | 100 | 96 | 97 | 104
Plasma samples for MET and GLY | LQC (60) | 92 | 82 | 85 | 84 | 81 | 92 | 107 | 107 | 102 | 106
Plasma samples for MET and GLY | MQC (240) | 90 | 87 | 83 | 80 | 81 | 104 | 106 | 88 | 91 | 110
Plasma samples for MET and GLY | HQC (960) | 94 | 90 | 93 | 83 | 85 | 100 | 96 | 98 | 94 | 106
Plasma samples for MET and GLY | ULOQ (1280) | 103 | 97 | 93 | 95 | 88 | 91 | 97 | 91 | 91 | 101

## Table 2
Unnamed: 0_level_0_Unnamed: 0_level_1 | Unnamed: 1_level_0_Nominal Conc.(ng/mL) | Intar Run (n = 6)_Calculated Conc.(ng/mL) | Intar Run (n = 6)_Accuracy(%) | Intar Run (n = 6)_CV(%) | Inter Run (n = 18)_Calculated Conc.(ng/mL) | Inter Run (n = 18)_Accuracy(%) | Inter Run (n = 18)_CV(%)
Non-isotope internal standard method | 20 | 20.8 ± 2.6 | 3.8 | 12.7 | 20.2 ± 0.9 | 0.8 | 4.3
Non-isotope internal standard method | 60 | 63.5 ± 7.5 | 5.9 | 11.7 | 59.4 ± 4.7 | −1.1 | 7.8
Non-isotope internal standard method | 240 | 234.4 ± 19.3 | −2.3 | 8.3 | 235.6 ± 18.5 | −1.8 | 7.9
Non-isotope internal standard method | 960 | 904.0 ± 75.8 | −5.8 | 8.4 | 981.3 ± 86.1 | 2.2 | 8.8
Isotope internal standard method | 20 | 21.8 ± 1.4 | 9.0 | 6.3 | 20.4 ± 1.2 | 2.0 | 6
Isotope internal standard method | 60 | 61.3 ± 1.8 | 2.2 | 3 | 60.4 ± 1.0 | 0.7 | 1.7
Isotope internal standard method | 240 | 244.7 ± 10.3 | 2.0 | 4.2 | 241.7 ± 5.1 | 0.7 | 2.1
Isotope internal standard method | 960 | 932.5 ± 35.5 | −2.9 | 3.8 | 943.0 ± 9.3 | −1.8 | 1

## Freetext Data
Data | Value | Relevance
Two drugs studied to evaluate signal interference | Metformin (MET) and glyburide (GLY) | The focus of the study was to model the potential impact of co-eluting drugs on LC-MS/MS signal interference for accurate drug quantification.
Effect of MET on GLY signal | Can suppress GLY’s signal by up to 30% | Demonstrates the primary issue under investigation, where co-eluting drugs impact accuracy in quantification.
Mention of stable isotope-labeled internal standard (SIL-IS) effectiveness | GLY-d11 effectively corrected the suppression of GLY by MET | Relevant as a proposed solution to mitigate quantification errors caused by signal suppression.
Matrix effect mechanism discussed | Charge competition and change in surface tension of droplets during electrospray ionization | Explains the root cause of signal interference observed in LC-MS/MS during drug co-analysis.
Dilution as a mitigation strategy | Reduces suppression but affects sensitivity and precision | Relevant as a partial solution explored in the paper; highlights its limitations in overcoming suppression fully.
Chromatographic separation findings | Effective but not universally applicable, can lead to new co-eluting drugs | Relevant in demonstrating practical difficulties and limitations in addressing co-elution issues through separation alone.
Precision and accuracy deviations without SIL-IS | Deviations of up to about 25% for GLY without SIL-IS correction | Highlights the quantification errors caused by signal suppression, which is a core focus of the paper.
Pharmacokinetic study impact due to suppression | GLY concentration curve is lower than theoretical value without SIL-IS correction | Relevant as a real-world example of how suppression errors affect drug monitoring and pharmacokinetic studies.
Recovery rates for MET and GLY in sample pretreatment | ~25% (MET) and ~62% (GLY) | Useful context for understanding how drug behavior during sample preparation influences quantification and supports finding regarding residual co-elution effects.
Proton affinity relevance | MET’s higher proton affinity led to suppression of GLY | Provides a mechanistic explanation for the observed signal interference between the two drugs.
Blood Collection Time | Various time points (see Table S2) | Relevant as it defines the time points for simulated pharmacokinetic study measurements.
Patient Population | Simulated samples using previously studied pharmacokinetics settings | Relevant because it specifies that simulated biological samples were used instead of actual patient samples for ethical reasons and to avoid intersubject variability.
Sample Size | Simulated samples rather than actual patients, preparation based on literature doses | Relevant as it indicates the approach taken for the PK study design to reduce animal use and variability factors.
Sampling Time Points | Based on previous pharmacokinetic studies—exact times derived from references | Relevant as it ensures the simulated sampling matches realistic pharmacokinetics profiling used in practice.
Patient Samples | Blank rat plasma harvested from healthy adult rats | Relevant because it identifies the source of the simulated matrix used in the analysis.
Oral Dose for MET | 45 mg/kg simulated human equivalent | Relevant for defining dosage settings for the simulated pharmacokinetic study framework.
Oral Dose for GLY | 10 mg/kg simulated human equivalent | Relevant for establishing the expected concentration-time data for GLY in the simulated PK study results.
GLY and MET pharmacokinetics analyzed with a simulated model | Simulated pharmacokinetics model | The study created simulated biological samples to mimic in vivo pharmacokinetic scenarios, useful for understanding signal suppression impacts under controlled conditions.
Deviation in GLY measurements without SIL-IS correction being 25% | Simulated pharmacokinetic model validation | Illustrates how the simulated pharmacokinetic model highlights the effect of signal suppression on drug measurements, emphasizing its use in method validation.
Simulated oral doses set for MET (45 mg/kg) and GLY (10 mg/kg) | Simulated pharmacokinetics design | Defines the ad-hoc pharmacokinetic model parameters applied to simulate combined drug analysis scenarios, enabling specific evaluations within the study goals.
Blood collection time and drug concentrations aligned with previous pharmacokinetic studies | Simulated pharmacokinetics model configuration | The configuration of the simulated model aligns it with prior validated pharmacokinetic modeling, ensuring credible comparisons and findings.
Spray Voltage | 4500 V | Relevant to setup of MS conditions for LC-MS/MS analysis, contributing to optimization.
Vaporizer Temperature | 350 °C | Crucial for MS detection in ESI positive ion mode, enabling efficient analyte ionization.
Capillary Temperature | 350 °C | Ensures stable evaporation of solvent within the ESI source, supporting detection reliability.
Sheath Gas Pressure | 40 psi | Optimizes the dispersion of analyte ions within the MS interface.
Aux Gas Pressure | 20 psi | Facilitates solvent evaporation and droplet formation during ionization.
Collision Pressure | 1.0 mTorr | Ensures appropriate fragmentation for mass spectrometry analysis of compounds.
Mobile Phase A | 2 mM ammonium acetate in water adjusted to pH 5.3 | Supports analyte retention and elution behavior in LC-MS/MS.
Mobile Phase B Ratio | 65% ACN | Parameter adjusted for co-elution versus separation during chromatographic setup.
Flow Rate | 0.25 mL/min | Specifies LC conditions affecting peak shapes and analysis throughput.
Injection Volume | 2 μL | Defines the sample amount enabling reproducibility in LC-MS/MS quantification.
Retention Time of MET and GLY | 2.16 min | Chromatographic condition used to simulate co-elution for examining ion suppression effects.
Retention Time of SIL-IS GLY-d11 | 0.01 min earlier than GLY | Demonstrates differential retention critical for SIL-IS application in suppression correction.
Signal Suppression Ratio Threshold | < 85% | Used for assessing the influence of co-elution between MET and GLY on quantification accuracy.
Dilution Factor | 10-fold and 20-fold | Explored for mitigating suppression effects in quantitative analysis, derived from experimental validation.
Calibration Curve Range for GLY | 20–1280 ng/mL | Estimated from bioequivalence pharmacokinetic requirements, defining quantification limits.
Concentration Ratio for ESI Ion Competition | Biased by MET with high proton affinity | Mechanism estimation based on Enke model to explain suppression phenomena.
Recovery for MET | ~25% | Reflects plasma sample adjustment impacts on quantitative pharmacokinetics.
Recovery for GLY | ~62% | Indicates sample preparation efficiency, affecting matrix influence on outcomes.
Matrix Effect Metrics | No significant effect for MET and GLY | Derived from sample pretreatment validation studies ensuring reliable analysis conditions.
Composition of blank biological matrix | Lack of co-elution of drugs or different concentrations | Relevant for method validation in pharmacokinetics as the matrix composition may not reflect actual clinical sample interference.
Signal suppression | High concentration of MET suppresses GLY by up to 66% | Suppression can directly affect the accuracy of pharmacokinetic analysis when determining drug concentrations.
Dilution | Reduces MET concentration to mitigate suppression effects | Attempts to address pharmacokinetic inaccuracies caused by interference but has limitations affecting sensitivity.
Stable-isotope-labeled internal standard | Corrects signal interference from co-eluting drugs | Used to improve quantitative accuracy in pharmacokinetics studies influenced by signal suppression.
Co-elution of drugs | MET and GLY have overlapping chromatographic peaks | Co-elution demonstrates signal suppression that can impact pharmacokinetic data accuracy.
Chromatographic separation | Effective in preventing co-elution but introduces new potential overlapped peaks | Relevant for pharmacokinetics analysis as incomplete separation may lead to misinterpretation of drug levels.
Endogenous matrix effect | Charge competition in droplets during ionization | Significantly affects quantitative drug concentration analysis, influencing pharmacokinetics interpretation.
Dilution multiples | 10- or 20-fold dilution affects sensitivity and precision | While reducing suppression, it challenges accurate pharmacokinetic analysis at lower drug concentrations.
Use of reserpine vs GLY-d11 | Reserpine caused deviation (~25%) in GLY results without SIL-IS correction | Highlights the impact of internal standards on pharmacokinetic data accuracy.
Adjustment of mobile phase pH | pH of aqueous phase adjusted to 5.3 (±0.1) with acetic acid | Regulates retention behavior of drugs, directly influencing quantitative consistency in pharmacokinetics.
Proton affinity | MET has higher proton affinity due to guanidine group | Explains signal suppression mechanism impacting accurate pharmacokinetic analysis.
Recovery rates | MET: ~25%; GLY: ~62% after plasma pretreatment | Influences measured drug concentrations critical for pharmacokinetic evaluations.