## Table 1
0
Performing pharmacokinetic studies in paediatric patients is difficult due to several challenges, such as low blood volume and venipuncture-related pain. We successfully developed an innovative intravascular microdialysis catheter that can be inserted into small intravenous lines for paediatric patients, which overcomes these challenges.
This proof-of-concept study is encouraging and opens the way to further experiments leading towards future use of microdialysis in paediatric patients.

## Table 2
Unnamed: 0 | 67 IV Microdialysis CatheterRef. no.: 8050092 | Modified 67 IV Microdialysis Catheter
CE-certification | Yes | No
Membrane | Membrane | Membrane
Material | Polyarylethersulphone | Polyarylethersulphone
Total length | 30 mm | 60 mm
Length within bloodstream | 30 mm | 15 mm
Diameter | 0.6 mm | 0.6 mm
Shaft | Shaft | Shaft
Material | Polyurethane | Polyurethane
Length | 46 mm | 10 mm
Diameter | 0.9 mm | 0.9 mm
Inlet tube | Inlet tube | Inlet tube
Material | Polyurethane | Polyurethane
Length | 200 mm | 200 mm
Diameter | 1.0 mm | 1.0 mm
Outlet tube | Outlet tube | Outlet tube
Material | Polyurethane | Polyurethane
Length | 45 mm | 95 mm
Diameter | 1.0 mm | 1.0 mm

## Table 3
Unnamed: 0_level_0_Unnamed: 0_level_1 | Recovery/loss_Physiological saline | Recovery/loss_Human albumin
Forward dialysis | Forward dialysis | Forward dialysis
0–60 min | 27.8% | 48.4%
60–120 min | 28.3% | 54.2%
120–180 min | 29.7% | 55.6%
Mean ± SD: | 28.6% ± 1.0% | 52.7% ± 3.8%
Retrodialysis | Retrodialysis | Retrodialysis
0–60 min | 16.4% | 50.1%
60–120 min | 12.4% | 45.0%
120–180 min | 11.3% | 42.3%
Mean ± SD: | 13.4% ± 2.7% | 45.8% ± 4.0%

## Table 4
Characteristic | Unnamed: 1
Number of subjects | 10
Age (years) | 27 (25–28)
Weight (kg) | 77.6 (72–80)
Height (cm) | 178 (176–182)
eGFR (mL/min) | 120.5 (107–122)
Albumin (g/L) | 45.5 (45–47)
Total protein (g/L) | 71.7 (69–72)
Plasma protein binding (%) | 49.7 (45–52)

## Table 5
Loss rates (%)_Loss rates (%) | Subject ID_1 | Subject ID_2 | Subject ID_3 | Subject ID_4 | Subject ID_5 | Subject ID_6 | Subject ID_7 | Subject ID_8 | Subject ID_9 | Subject ID_10
MD probe 1 | 17 | 15.9 | 16 | 22.5 | 24.4 | 19.8 | 13.3 | 13.1 | 20.7 | 26.8
MD probe 2 | 16.9 | NAa | 17 | 21.4 | NAa | 19.7 | 20.4 | 18.6 | 15.5 | 16.4

## Table 6
Vancomycin PK parameters | MD | Plasma | p value
t 1/2 (h) | 10.6 ± 6.7 | 6.3 ± 0.8 | 0.022t
C max (mg/L) | 28.5 ± 7.8 | 28.2 ± 4.9 | 0.832W
AUClast (h⋅mg/L) | 87.82 ± 19.1 | 98.2 ± 9.1 | 0.130W
AUC∞ (h⋅mg/L) | 97.56 ± 22.0 | 103.9 ± 9.9 | 0.369W
Cl (L/h) | 10.9 ± 2.7 | 9.7 ± 0.95 | 0.089t

## Table 7
Parameter | Estimate | 95% CI | IIVa
CL [L/h] | 4.74 | 4.47–5.02 | 9.0 %CV
V1 [L] | 10.8 | 9.27–12.5 | 21.0 %CV
Q [L/h] | 4.95 | 4.2–5.8
V2 [L] | 19 | 17.8–20.2
Recovery [–] | 0.175 | 0.156–0.193 | 15.7 %CV
FUMD [–] | 0.48 | 0.40–0.59 | 27.7 %CV
FUUF [–] | 0.51 | 0.50–0.53
σ Plasma Prop. [%CV] | 13.9 | 11.3–17.9
σ UF Prop. [%CV] | 7.1 | 5.6–8.6
σ MD Prop. [%CV] | 34.7 | 30.2–39.2
σ Recovery Prop. [%CV] | 5.5 | 4.1–7.2

## Freetext Data
Data | Value | Relevance
Vancomycin PK studies in pediatric patients | Unavailable for most antibiotics | Highlights the gap in pediatric PK data that the study aims to address.
Challenges of PK studies in pediatrics | Low blood volume and venipuncture-related pain | Explains barriers in conducting PK studies directly in pediatric populations.
Microdialysis (MD) potential use | Promising method for continuous and volume-neutral PK sampling | MD offers a minimally invasive solution for measuring unbound drug concentrations, suitable for pediatric challenges.
Study drug | Vancomycin | Representative drug chosen for proof-of-concept due to its common usage in pediatric populations.
Modified MD catheter | Fits smaller intravenous lines | Designed to enable MD application in pediatric patients.
In vitro MD experiment results | High recovery/loss with human albumin | Validated MD setup and calibration stability as a base for clinical testing.
In vivo MD experiment setting | Calibration via retrodialysis | Ensured accurate analysis for unbound drug concentration measurement.
Unbound drug fraction determination | Through ultrafiltration and MD | Focuses on drug's active fraction, corresponding to its efficacy and safety in pediatric dosing models.
Comparison method | Agreement with Bland-Altman and PK parameters | Relevant to validating the new MD catheter for replacing blood sampling in pediatric PK studies.
Plasma protein binding (PPB) in vancomycin | Mean 49%; range 46–52% | Demonstrated consistency with previous literature for unbound drug fraction reliability.
Demographics of study subjects | Mean age 27.3, eGFR within normal range | Adult volunteer data relevant as a stepping stone in validating MD probe before pediatric application.
Safety | No safety issues with MD device | Critical for feasibility of application in the vulnerable pediatric population.
Study results | Good alignment between MD and plasma measurements | Supports MD as a reliable method for collecting PK data while overcoming ethical and practical barriers.
Bland-Altman plot bias | Mean bias 0.19 mg/L | Provides quantitative evidence of the MD catheter's measurement accuracy.
PK parameters comparison | Significant difference in half-life only | MD's measured PK parameters show relatively low variability vs. traditional methods, supporting pediatric use.
Impact of PPB concentration | No evidence of significant variation | Reinforces the reliability of unbound drug measurements via MD even under dynamic changes.
Clinical study format | Open-label, prospective | Ensures transparency and serves as a proof-of-concept for pediatric applicability.
Next step for MD in pediatrics | Further refinement and testing in specific pediatric contexts | Represents the study's ultimate goal of addressing the pediatric PK challenges discussed.
Patient Sample Size | 10 healthy adult volunteers | The study was conducted in a controlled population to validate the modified MD catheters before transitioning to pediatric studies.
Blood Sampling Time Points | Baseline and post-drug administration at 30 min, 1, 2, 3, 4, 5, 6, 8, 12, and 24 hours | Frequent blood sampling was performed to capture the key pharmacokinetic phases of the drug.
Microdialysate Sampling Time Points | Baseline and at 15–45 min, 45–75 min, 1.5–2.5 h, 2.5–3.5 h, 3.5–4.5 h, 4.5–5.5 h, 5.5–6.5 h, 7.5–8.5 h, 11.5–12.5 h, and 23.5–24.5 h | These intervals were chosen to correspond with blood sampling points, ensuring comparability of the methods.
Patient Population | Healthy adult male volunteers | This population allowed for the controlled evaluation of the novel MD catheter in a stable physiological environment without confounding factors such as underlying illness.
Drug Dose Administered | Single intravenous dose of 1000 mg vancomycin | Standardizing the dose and administration ensured consistency and relevance to the study objectives for validating vancomycin PK measurement methods.
Calibration Method | Retrodialysis technique conducted in vivo after the last microdialysate sample over a 30 min run-in followed by 2×30 min sampling | Calibration ensured accurate measurement and adjustment of MD-derived drug concentrations.
Venous Access | Four intravenous lines, including one for drug infusion and three for blood and microdialysate sampling | The setup allowed robust sampling protocol implementation and ensured precise PK data collection.
Model Used | Two-compartment model | Plasma PK was best described by a two-compartment model, highlighting its relevance in estimating the pharmacokinetics of vancomycin in the study.
Unbound Fraction Model | Constant fraction for both ultrafiltrate and MD | This model was used to describe the unbound drug fraction, essential for understanding pharmacokinetic differences between plasma sampling and microdialysis methods.
Recovery Estimation Within the Model | Integral sampling approach by Tunblad et al. | This method was implemented for raw microdialysis measurements to accurately estimate recovery, which supports the validation of the MD probes for PK sampling.
Vancomycin Clearance (Cl) | 9.7 L/h | Estimated from plasma sampling and reflects the drug elimination rate in HVs; essential for pharmacokinetics modeling and dose optimization.
Terminal Half-life (t1/2) | 4.95 h | Calculated from plasma sampling; crucial for understanding drug duration and interval administration.
PPB at 2.5 mg/L | 46.1% | Protein binding percentage determined via ultrafiltration; important for interpreting unbound, active drug availability.
PPB at 50 mg/L | 52.0% | Protein binding percentage calculated with higher concentration; examines any dose-dependent binding variability.
Mean PPB Overall | 49%, SD 1.7% | Aggregated protein binding value determined across concentrations; instrumental in evaluating unbound drug levels.
Vancomycin Recovery (in vitro, experiment B) | 52.7%, SD 3.8% | In vitro forward dialysis recovery estimate; validates the modified microdialysis probe.
Vancomycin Loss (in vitro, experiment B) | 45.8%, SD 4.0% | Loss rate from retrodialysis experiments in vitro; crucial for back-calculating accurate unbound concentrations.
In vivo Loss Rate | 18.6%, SD 3.7% | Estimated using retrodialysis technique; critical for correcting microdialysate concentrations to actual unbound drug levels in plasma.
Mean AUC0–24 (MD) | 87.8 mg·h/L | Calculated from microdialysis data; reflects total drug exposure over time, essential for efficacy predictions.
Cmax (Plasma Sampling) | 74.6 mg/L | Value derived from plasma concentration; the peak drug concentration helps assess dosage adequacy.
Fraction Unbound (MD) | 0.48, CI 0.40–0.59 | Determined in pharmacometric modeling using raw MD measurements; reflects active drug proportion available for therapeutic effect.
Fraction Unbound (Ultrafiltration) | 0.51, CI 0.50–0.53 | Based on ultrafiltration; mirrors physiological unbound drug estimations aligning closely with MD data.
Recovery Estimate (Pharmacometric Model) | 17.6% | Derived from model-based analysis; confirms consistency between observed and predicted values for microdialysis performance.
Visual Predictive Check Overview | Slight underprediction at 8 and 12 h (MD) | Indicates adequacy of pharmacometric model fitting plasma and MD data; assesses prediction reliability.
Low blood volume | Limited in paediatric patients | Relevant because it restricts the amount of blood that can be sampled for pharmacokinetic studies.
Plasma Protein Binding (PPB) | Vancomycin PPB determined to be 46.1% at 2.5 mg/L and 52.0% at 50 mg/L | Relevant because PPB affects drug distribution and availability of the active unbound drug fraction.
Developmental differences in organ function | Significant differences below the age of 2 years | Relevant as these differences influence drug metabolism and clearance in paediatric patients compared to adults.
Body weight and composition | Continually changing during childhood | Relevant because weight and body composition adjustments are critical for accurate drug dosing and pharmacokinetic modelling.
Flow rate in microdialysis | Lower flow rate increases recovery | Relevant because flow rates influence the measured concentrations and variability in pharmacokinetic studies using microdialysis.
Concentration-dependent protein binding | PPB changes with drug concentrations | Relevant because it affects the unbound drug fraction used in pharmacokinetic calculations.
Patient variability | Accounted for 46.7% in observed variability | Relevant because inter-individual variability greatly impacts pharmacokinetic predictions and therapy optimization.
Modified MD catheter design | Ensures adequate recovery in pediatric settings | Relevant as catheter design influences drug measurement reliability and applicability in paediatric pharmacokinetics.
Sampling interval duration | MD samples represent averages over 30-60 min | Relevant because it impacts the precision and temporal resolution of concentration measurement in pharmacokinetics.
Dynamic changes in plasma proteins | Observed in critically ill patients | Relevant because plasma protein variations affect drug binding and pharmacokinetics, especially in vulnerable populations.