## Table 1
0 | 1
Matrix Size | 128 × 128
Collimator | Medium Energy General Purpose
Views per Head | 60 (120 in total)
Time per View | 30 s
Body Contour | Enabled
Photopeak | 208.0 keV ± 10 %
Lower Scatter Window | 181.3 keV ± 3 %
Upper Scatter Window | 236.4 keV ± 3 %

## Table 2
Organ_Unnamed: 0_level_1 | STL Volume (mL)_Unnamed: 1_level_1 | Filling volume (% of STL volume)_Scan 1 | Filling volume (% of STL volume)_Scan 2 | Filling volume (% of STL volume)_Scan 3 | Filling volume (% of STL volume)_Scan 4 | Filling volume (% of STL volume)_Scan 5 | Filling volume (% of STL volume)_Scan 6
Liver | 1306.7 | 100.08(1) | 99.84(1) | 100.04(1) | 99.95(1) | 100.13(1) | 99.89(1)
Spleen | 124.9 | 99.43(2) | 99.45(2) | 99.40(2) | 99.21(2) | 99.29(2) | 99.23(2)
Tumour | 15.9 | 99.6(1) | 99.4(1) | 99.2(1) | 99.5(1) | 99.7(1) | 94.8(1)
Left kidney cortex | 90.9 | 97.69(2) | 96.48(2) | 98.92(2) | 98.14(2) | 98.33(2) | 98.40(2)
Left kidney medulla | 42.9 | 99.05(5) | 97.19(5) | 102.04(5) | 97.47(5) | 97.72(5) | 97.19(5)
Right kidney cortex | 75.7 | 99.29(3) | 98.89(3) | 99.59(3) | 98.65(3) | 99.48(3) | 98.72(3)
Right kidney medulla | 35.7 | 96.20(6) | 97.06(6) | 96.56(6) | 96.45(6) | 97.32(6) | 97.60(6)
Background | 8297 | 98.25(3) | 98.25(3) | 98.25(3) | 98.35(3) | 98.28(3) | 98.24(3)

## Freetext Data
Data | Value | Relevance
Design of the 3D printed phantom | Based on ICRP 110 Female Computational Phantom | Relevant for MRT validation as it mimics a patient undergoing PRRT.
STL files availability | The STL files for the phantom are available online | Relevant as institutions can print and use their own phantom for validation purposes.
Type of radionuclide used | 177Lu DOTATATE | Relevant since data generation and validation focus on this isotope for PRRT dosimetry.
Filling process accuracy | Mean deviation of 0.6(8)% | Relevant for validating activity concentrations against the pharmacokinetic model.
SPECT/CT time points | T=1 h, 4 h, 24 h, 40 h, 72 h, 144 h | Relevant for creating clinically-relevant validation datasets.
Camera used for imaging | GE Discovery 670 SPECT/CT | Relevant as it affects the consistency and compatibility of imaging data.
Calibration phantom types | Cylindrical Jaszczak and IEC NEMA PET body phantoms | Relevant for quantitative imaging calibration techniques.
Activity measurement methodology | Linked to UK NPL standard | Relevant for ensuring traceability and reproducibility in multi-centre studies.
Dose quantification method | Model solved using MATLAB RK4 algorithm | Relevant for ensuring the accuracy of dose calculations.
Multi-centre comparison | Dataset used for multicentre MRT dosimetry exercises | Relevant as it supports inter-institutional validation efforts.
Phantom shell material | Polylactic acid (PLA) | Relevant for establishing tissue equivalence in SPECT imaging studies.
Activity concentration | Agreed with pharmacokinetic model | Relevant for ensuring the reliability of quantitative imaging for dosimetry.
Pharmacokinetic Model | Implemented using SAAM II software for kinetic analysis | The model provides activity concentrations over time, critical for dosimetry validation.
Patient Population | Mimicked using a 3D printed phantom based on ICRP 110 female computational phantom | This simulates patient biological behavior for PRRT dosimetry accuracy validation.
Activity Concentration Timepoints | T=1 h, 4 h, 24 h, 40 h, 72 h, and 144 h | Represent common clinical imaging timepoints in post-therapy scenarios.
Activity Estimation | Administered activity assumed at 7.4 GBq | Reflects typical PRRT therapeutic activity, supports realistic dataset development.
Liver Tumor Inclusion | Added a 15.9 mL spherical compartment filled with activity | Represents tumorous uptake in patients for accurate dosimetry testing.
Organs Sampled | Liver, kidneys (cortex and medulla), spleen, background tissue | Key organs for PRRT dosimetry and OARs.
Activity Measurement | Verified using radionuclide calibrator decay correction | Ensures precision in activity measurement aligning to PK model outputs.
Validation Methodology | Dispensed activities matched PK model predictions within uncertainties (mean deviation 0.6%) | Confirms adherence to simulated patient readings for dosimetry validity.
Model Type | Pharmacokinetic model | The paper describes the use of a pharmacokinetic model to generate activity concentrations and validate MRT dosimetry systems, establishing its relevance to dosimetry workflows.
Compartment Description | Kidney medulla activity concentration set as one third of kidney cortex | This demonstrates the compartment-specific modeling approach used, suggesting a simplified compartmental model relevant for phantom design.
Software for Implementation | SAAM II software for kinetic analysis | This highlights the tool used for implementing the pharmacokinetic model, crucial for achieving accurate activity concentration predictions.
Compartmental Adjustments | Kidney medulla and renal pelvis combined into one compartment | Relevant to understanding how the compartments were modeled and simplified for creating the anthropomorphic phantom and drug distribution simulations referred to in the study.
SPECT acquisition parameters | 2 detectors, 2×128 projections over 360°, 60s per projection, photopeak energy window: 208 keV ± 10%, lower scatter window: 170 keV ± 10%, upper scatter window: 236 keV ± 10% | These acquisition parameters are crucial as they define the imaging setup used to collect data for validating the Molecular Radiotherapy Dosimetry system. They ensure consistent data collection for accurate validation comparisons.
Filled activity concentration | Pharmacokinetic model with activity corrected to T=1 h, 4 h, 24 h, 40 h, 72 h, 144 h | These activity concentrations are estimated from a pharmacokinetic model and used for multi-time-point imaging, ensuring that dosimetry calculations have contextually relevant inputs tied to realistic clinical scenarios.
Reconstruction algorithm | Ordered Subset Expectation Maximization (OSEM) algorithm, 4 iterations, 10 subsets | This reconstruction method aligns activity imaging with standards required for dosimetric analysis, ensuring accurate representation of the phantom data relative to clinical practice.
Scatter correction method | Triple Energy Window (TEW) method | Scatter correction is a necessary parameter for minimizing image distortions, directly impacting the validation and accuracy of dosimetry calculations.
Activity measure calibration factor | Linked to UK National Physical Laboratory (NPL) primary standard | Calibration of activity measurement ensures traceability and consistency across studies, a prerequisite for intercenter dosimetry comparison and validation exercises.
Phantom materials | Polylactic acid (PLA) with tissue equivalency, Perspex© for housing | Parameters like material tissue equivalence impact quantitative imaging accuracy and dosimetry validation relevance, simulating real patient conditions.
Pharmacokinetic solving algorithm | Fourth-Order Runge–Kutta method (ΔT=1 min) | Ensures precise resolution of activity concentration time curves, critical for validating dosimetry systems against temporal dose delivery metrics.
Measured insert volumes | Left kidney cortex: 94.0%-98.0%, medulla: 89.3%-96.0%, liver: 97.5%–102.4%, spleen: 97.8%–100.0%, tumour: 94.0%–103.0% of STL volumes | Accurate volume measurements correspond to the estimated pharmacokinetic model parameters, directly affecting dose estimation fidelity in dosimetry validation.
Recovery curve methodology | Partial volume correction method | Facilitates corrections for spill-out partial volume effects, directly impacting quantitative imaging and calibration accuracy as part of dosimetry system validation.
Activity concentration | Accuracy of dispensed activity concentration matched pharmacokinetic model within measurement uncertainties | Demonstrates how accurate activity measurements directly relate to pharmacokinetics described in the model and contribute to validation.
Compartment volumes | Adjusted in the pharmacokinetic model to match STL volumes of organs | Influences the pharmacokinetics by ensuring relevant organ-specific data is accurately modeled in calculations.
Kidney medulla activity | Set to be one-third of the activity concentration in the kidney cortex | Specific ratios used in pharmacokinetics modeling are critical to accurately represent organ functionality in dose calculations.
Elliptical phantom background activity | Matched to the rest of body in the pharmacokinetic model | Reflects systemic background activity which affects pharmacokinetic distribution.
Radioactive decay | Corrected for physical decay of 177Lu | Intrinsic factor influencing pharmacokinetics that needs to be accounted for during modeling and measurement.
Filling procedure | Precision in filling compartments ensured minimal air bubble trapping | Accurate compartment filling is crucial for reflecting realistic pharmacokinetic data.
Scan time points | Chosen post-therapy imaging times: 1h, 4h, 24h, 40h, 72h, 144h | Relevant clinical imaging timepoints directly influence pharmacokinetic modeling and validation accuracy.
Compartments and organs | Defined as kidney cortex, medulla, liver, spleen, and spherical tumor | Specific compartment design affects pharmacokinetics as organs influence dose distributions.