## Table 1
Analyte | Precursor ion(m/z) | Product ion(m/z) | Collision energy(eV) | Dwell time(msec) | Retention time(min)
MPA | 321.10 321.10 | 275.00 285.00 | − 15 − 20 | 28 28 | 5.58
MPA-d3(internal standard) | 324.00 341.00 | 210.00 210.00 | − 22 − 35 | 28 28 | 5.58
MPAG(only monitored) | 514.21 514.21 | 207.04 303.13 | − 38 − 18 | 28 28 | 5.2

## Table 2
Statistic_Statistic | Before conversion (cross-validation)_MPAVAMS-LC–MS/MS vs MPAWB-LC–MS/MS | Before conversion (cross-validation)_MPAVAMS-LC–MS/MS vs MPAPL-LC–MS/MS | Before conversion (cross-validation)_MPAPL-HPLC–DAD vs MPAPL-LC–MS/MS | Before conversion (cross-validation)_MPAVAMS-LC–MS/MS vs MPAPL-HPLC–DAD | After conversion (clinical validation)_MPAVAMSHt-LC–MS/MS vs MPAPL-LC–MS/MS | After conversion (clinical validation)_MPAVAMSc-LC–MS/MS vs MPAPL-LC–MS/MS | After conversion (clinical validation)_MPAVAMSc_DAD-LC–MS/MS vs MPAPL-HPLC–DAD
Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula) | Passing–Bablok analysis (regression formula)
Passing-Bablok-regression formula | CVAMS = 0.90·CWB + 0.17 | CPL = 1.60·CVAMS + 0.08 | CHPLC = 1.01·CPL + 0.11 | CHPLC = 1.56·CPL + 0.17 | CPL = CVAMSHt — 0.13 | CPL = CVAMSc — 0.01 | CVAMSc_DAD = 1.01·CHPLC —0.02
Intercept (A) | 0.17 (− 0.06 to 0.29) | 0.08 (− 0.14 to 0.29) | 0.11 (− 0.24 to 0.30) | 0.17 (− 0.22 to 0.38) | − 0.13 (− 0.35 to 0.03) | − 0.01 (− 0.27 to 0.19) | − 0.03 (− 0.35 to 0.31)
Slope (B) | 0.90 (0.79–1.00) | 1.60 (1.44–1.80) | 1.01 (0.91–1.13) | 1.56 (1.38–1.78) | 1.00 (0.92–1.09) | 1.00 (0.90–1.11) | 1.01 (0.89–1.15)
Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias) | Bland–Altman bias analysis (based on plot and %-differences bias)
%-mean difference bias (− 1.96SD; + 1.96SD) | − 6.80% (− 61.50 to 47.90%) | 54.07% (48.62–59.53%) | + 4.08% (− 1.21 to 9.37%) | 50.00% (1.60–98.50%) | + 10.90% (5.24–16.56%) | 4.76% (− 1.11 to 10.64%) | − 1.01% (− 8.26 to 6.24%)
% of paired samples fulfilled the EMA criteria (< 20%) | 70% | 0% | 12% | 12% | 88% | 86% | 72%
% of paired samples fulfilled the IATDMCT criteria (< 15%) | 60% | 0% | 6% | 4% | 78% | 72% | 67%
Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors | Predictive analysis performance factors
MPE | − 0.05 | − 0.92 | − 0.12 | − 0.81 | − 0.12 | − 0.01 | 0.00
MPPE [%] | − 5.50 | 85.00 | 1.90 | 67.54 | 7.03 | 9.20 | 2.60
RMSE [%] | 15.00 | 92.00 | 12.00 | 81.54 | 12.14 | 12.21 | 13.00
MAPE [%] | 13.69 | 66.65 | 8.25 | 62.42 | 14.02 | 13.38 | 20.24
Correlation coefficients | Correlation coefficients | Correlation coefficients | Correlation coefficients | Correlation coefficients | Correlation coefficients | Correlation coefficients | Correlation coefficients
Pearson’s (R 2) | 0.92 (0.87–0.96; 95% CI) | 0.92 (0.86–0.95; 95% CI) | 0.93 (0.88–0.96; 95% CI) | 0.90 (0.84–0.95; 95% CI) | 0.95 (0.85–0.97; 95% CI) | 0.92 (0.87–0.96; 95% CI) | 0.91 (0.85–0.95; 95% CI)
Spearman (SRCC) | 0.81 (0.68–0.89; 95% CI) | 0.88 (0.80–0.93; 95% CI) | 0.92 (0.86–0.95; 95% CI) | 0.83 (0.73–0.90; 95% CI) | 0.91 (0.84–0.95; 95% CI) | 0.88 (0.80–0.93; 95% CI) | 0.84 (0.73–0.90; 95% CI)
Intraclass (ICC) | 0.91 (0.85–0.95; 95% CI) | 0.84 (0.73–0.90; 95% CI) | 0.96 (0.93–0.98; 95% CI) | 0.83 (0.74–0.91; 95% CI) | 0.94 (0.90–0.97; 95% CI) | 0.92 (0.87–0.95; 95% CI) | 0.91 (0.85–0.95; 95% CI)

## Freetext Data
Data | Value | Relevance
--- | --- | ---
Matrix for MPA determination | EDTA plasma (PL), heparinized plasma (HP), serum (SR) | Relevant as these are recommended matrices for MPA determination per IATDMCT guidelines.
Analytical method validated | LC–MS/MS method for MPA determination in whole blood, plasma, and VAMS | Relevant as the study aimed to validate UHPLC-MS/MS for MPA across different matrices.
Calibration range for LC–MS/MS | 0.10–15 µg/mL | Relevant for setting the validated measurement range for trough MPA concentrations according to the EMA guidelines.
Pediatric patients included | 50 | Relevant as it highlights the study's focus on the pediatric demographic post-renal transplantation.
Sample types collected | Whole blood (WB), plasma (PL), VAMS samples | Relevant since the study tests comparability and conversion between these sampling approaches.
Sampling technique for VAMS | Fingerstick with Mitra™ device | Relevant as VAMS simplifies sampling in pediatric patients and avoids venipuncture.
Conversion formula used | MPAVAMS / [1- HCT] = MPAPL | Relevant for ensuring clinical interpretation of MPA levels collected via VAMS.
Regression model used | Passing-Bablok regression for VAMS to PL conversion | Relevant as the study aims to establish reliable conversion methods.
Stability of VAMS samples | Stable for 42 days at RT, 4°C in the dark | Relevant for VAMS use in TDM, enabling remote and convenient sample storage.
Main clinical reason for TDM | Prevent toxicity and acute or chronic graft rejection | Relevant for understanding the requirement for precise MPA monitoring post-transplantation.
Hematocrit effect on VAMS | Correlation coefficient = −0.02596 indicates no significant impact | Relevant for confirming VAMS suitability in pediatric patients without HCT adjustment.
Interchangeability confirmed | VAMS-LC–MS/MS and WB-LC–MS/MS methods showed equivalence | Relevant for supporting clinical validation of VAMS as an alternative sampling method.
Established regression formula | y = 1.60x + 0.08 for VAMS to PL conversion | Relevant as it provides a clinically validated approach to extrapolate MPA levels.
Predictive performance criteria | MPE, MAPE, MPPE <15% | Relevant for establishing statistical reliability of concentration comparisons across methods.
Clinical validation result | Formula successfully validated for independent sample group | Relevant as clinical validation is necessary for real-world application of the findings.
Patient demographic focus | Pediatric renal transplantation | Relevant because the study seeks to optimize methods specifically for this vulnerable population.
Comparison with methods | Cross-validated LC–MS/MS and HPLC-DAD methods | Relevant for confirming analytical reliability and method compatibility.
Correlation coefficients between methods | Pearson: R2 = 0.92–0.95 | Relevant for assessing statistical agreement across different testing approaches.
Fifty pediatric renal transplant recipients participated in this study. | 50 | Indicates the size of the patient sample used in the pharmacokinetic study, which is key for statistical and clinical reliability.
The children were treated at the Kidney Transplantation Outpatient Clinic, Children’s Memorial Health Institute (CMHI), Warsaw. | Renal transplant pediatric patients | Explains the population on which the study was focused, i.e., pediatric kidney transplant recipients, pertinent for understanding the application of the PK study.
Samples of WB and PL were stored at –20 °C pending the LC–MS/MS assay. | -20°C | Highlights the preservation conditions of the blood samples, which impact sample integrity and PK data reliability.
Samples were collected at pre-dose (trough concentration, C0) during regular post-transplant follow-up visits. | C0 (trough concentration) sampling | Specifies the timing of the blood sampling, crucial for determining pharmacokinetics and therapeutic drug monitoring.
VAMS samples (finger puncture by a lancet) were simultaneously collected and analyzed during the next two weeks. | VAMS collection method | Details an alternative minimally invasive collection method for capillary blood, relevant to the study design and pediatric viability.
MPA levels targeted were in the range of 0.10–15 µg/mL for trough concentration analysis. | 0.10–15 µg/mL | Defines the target analytical range for MPA measurement, which informs the tools and precision used in the PK analysis.
The patients were on regimens including TAC, MMF, and glucocorticosteroids. | TAC, MMF, and corticosteroids | Indicates the co-medication used, which may influence pharmacokinetics and relevant study interpretations.
Samples collected were WB, PL, and VAMS. | WB, PL, VAMS | Specifies the bio-matrices used, essential for understanding the study's comparative analysis and PK results.
The pediatric participants and their parents or legal guardians provided informed consent before sampling. | Informed Consent | Ethical consideration ensuring appropriate inclusion of such a vulnerable population in the study.
This study was conducted between November 2022 and February 2023. | Nov 2022 - Feb 2023 | Provides the temporal scope of the data collection, contextualizing the design of the PK analysis over time.
Samples obtained during routine post-transplant follow-ups. | Routine clinical visits | Describes the real-world setting in which samples were collected, enhancing the relevance to daily clinical practice.
Model Type | Not explicitly stated (specific model information not included) | The paper validates LC–MS/MS methods for the quantification of Mycophenolic Acid (MPA) in different matrices, but it does not employ or describe a pharmacokinetic model such as 1-compartment, 2-compartment, or physiologically based models (PBPK), nor an IVIVE framework.
Purpose | Method validation and cross-validation of MPA concentrations in pediatric post-transplant patients | The focus of the research was on method development and clinical validation for TDM, rather than on pharmacokinetic modeling.
Context | Analytical method development and concentration corrections | The study corrects and validates MPA concentration data using mathematical regression formulas, but no pharmacokinetic modeling frameworks are described.
Calibration range | 0.10–15.0 µg/mL | The calibration range was established for MPA quantification to align the LC-MS/MS method with therapeutic drug monitoring in pediatric patients post-renal transplantation.
Calibration curve | Linear 1/x weighting | Demonstrates how calibration curve fitting was accomplished for optimal quantification accuracy of MPA in different matrices.
Correlation coefficient (R2) for WB-LC–MS/MS | 0.9984 ± 0.0013 | Ensures high linearity and robustness of the analytical method for measuring MPA in whole blood.
Correlation coefficient (R2) for PL-LC–MS/MS | 0.9994 ± 0.0013 | Confirms the accuracy of the analytical method for MPA determination in plasma, which is the recommended matrix.
Accuracy range for WB and PL | Within 85-115% | Validates that the LC-MS/MS methods meet the EMA’s accepted range for bioanalytical accuracy in WB and PL matrices.
Precision for WB and PL | < 15% and < 20% for LLOQ | Precision was validated for both whole blood and plasma, adhering to EMA guidelines for reproducibility.
Percentage mean difference bias for WB/VAMS | −6.80% | Confirms the acceptability of the methods when comparing WB and VAMS matrices for MPA analysis, using EMA and IATDMCT standards.
Validation method | EMA guidelines | Demonstrates adherence to internationally recognized standards for bioanalytical method validation.
Stability at RT for VAMS-loaded samples | Six weeks | Highlights the sample-storage conditions and timeframe applicable for MPA quantification using VAMS devices.
Passing-Bablok regression slope for VAMS-LC-MS/MS vs WB | 0.90 (0.79–1.00, 95% CI) | Establishes comparability between these sampling techniques, critical for clinical validation of VAMS.
Passing-Bablok regression slope for VAMS-LC-MS/MS vs PL (regression adjusted) | 1.00 (0.90–1.11, 95% CI) | Ensures converted VAMS results match plasma-derived MPA levels for clinical equivalency.
Predictive performance parameters | MPPE: 9.20%, MAPE: 13.38%, RMSE: 12.21% | Indicates how MPA concentrations predicted in VAMS align with plasma via a regression formula, reflecting good predictive validity.
Autosampler stability | Five days at 5°C | Measure of how samples maintain integrity during routine lab operations, validated for VAMS, WB, and PL.
Matrix effects for WB | Fulfilled EMA criteria | Quantifies how different matrices affect analyte recovery and precision, critical for method reliability.
MPA stability in VAMS at −20°C | Eight weeks | Ensures the robustness of VAMS for use in extended timeframes, useful for clinical and research settings.
Conversion formula from VAMS to PL | y = 1.60x + 0.08 | Provides a validated regression equation for interpreting MPA concentrations in plasma equivalent contexts, a necessity for therapeutic monitoring adjustments.
HT independence correlation for MPA in VAMS | −0.02596; 95% CI: −0.3171 to 0.2696 | Confirms MPA recovery from VAMS is not affected by hematocrit, eliminating a confounding variable.
Hematocrit level | No observed correlation with MPA recovery from VAMS sampler (correlation coefficient: -0.02596) | Relevant because hematocrit level was evaluated as a potential influencing factor on pharmacokinetics but deemed not significant in the study.
Food-drug interactions | Confounding factors in MPA exposure and dose relationships | Relevant because food-drug interactions are listed as a factor that can affect MPA pharmacokinetics, underscoring the need for monitoring.
Age | Children younger than 6 years require closer attention due to potential adverse effects or toxicities | Relevant since patient age influences drug metabolism and efficacy/adverse effect profile, requiring individualized treatment adjustments.
Inter- and intra-individual variability | Observed high variability in MPA pharmacokinetics between individuals | Relevant as variability necessitates therapeutic drug monitoring for precise dosage adjustment in pediatric transplant recipients.
Plasma protein binding | MPA binds almost exclusively to plasma proteins | Relevant because the drug's pharmacokinetics are strongly tied to protein binding properties, affecting distribution and efficacy.
Stability conditions | MPA stability in VAMS was variable depending on storage conditions (RT, 4°C, dark/light exposure) | Relevant as storage stability can influence pharmacokinetics if samples degrade before analysis.
Co-medication with TAC | TAC is a common co-medication affecting immunosuppressive pharmacokinetics | Relevant because co-administered drugs like TAC can alter MPA metabolism and exposure.
Sample matrix type | MPA levels in WB were lower than in plasma due to dilution effects | Relevant since matrix type directly influences measured pharmacokinetics values and interpretation.
Adherence to therapy | Adherence influences pharmacokinetic variability and is assessed via monitoring | Relevant because non-adherence can complicate the pharmacokinetic profile and necessitate monitoring.