## Table 1
Blood group (DEHP released, mg)_1a | Blood group (DEHP released, mg)_2a | Blood group (DEHP released, mg)_3a | Water/ethanol mixture group (DEHP released, mg)_Unnamed: 3_level_1
0.11 | 0.068 | 0.073 | 2.84

## Table 2
Concentration (μg/mL) | Mean (Peak area) | Standard deviation (SD) | Relative standard deviation (RSD%)
Very low (0.1)a | 7693 | 299 | 3.88
Low (1.0) | 215724 | 3347 | 1.55
Intermediate (50.1) | 14429286 | 502984 | 3.48
High (200.6) | 48864521 | 1841486 | 3.76

## Table 3
Spiked concentration μg/mL_Spiked concentration μg/mL_Unnamed: 0_level_2 | Spiked sample recovery_Spiked in ethanol/water (n = 3)_Mean recovery (%) | Spiked sample recovery_Spiked in ethanol/water (n = 3)_RSD (%) | Spiked sample recovery_Spiked in blood (n = 3)_Mean recovery (%) | Spiked sample recovery_Spiked in blood (n = 3)_RSD (%)
0.1 | 94.23 | 2.33 | 92.43 | 3.88
1 | 94.23 | 4.51 | 91.04 | 6.71
10 | 101.53 | 8.8 | 98.9 | 0.57
100 | 89.58 | 7.16 | 98.15 | 2.24
180, 200a | 94.6 | 5.02 | 95.35 | 7.86

## Table 4
Group | Mean (n = 6) μg/mL | Total RSD (%) within the two groups (n = 12)
Group 1 | 13.17 | 9
Group 2 | 12.66

## Table 5
Medical devices | Extraction media | DEHP released, mg
Transfusion set for single usea | Water/ethanol mixture | 2.84
Transfusion set for single usea | blood | 0.068-0.073b
Leukcyte-reducing filter for single use | Water/ethanol mixture | 1.2
Tube system for haemodialysis use | Water/ethanol mixture | 11.6