## Table 1
Source temperature (°C) | 500
Dwell time per transition (ms) | 150
Nebulizer gas (psi) | 12
Auxiliary gas (L/min) | 8
Curtain gas setting | 10
Collision gas setting | 12
IonSpray voltage (V) | 2000
Declustering potential (V) | 35
Focusing potential (V) | 210
Collision energy (eV) | 45
Collision Cell Exit Potential (V) | 25
Resolution for Q1 and Q3 | Unit
Mode of analysis | Positive
Ion transition for erlotinib, m/z | 394→278
Ion transition for IS, erlotinib-d6 , m/z | 400→278

## Table 2
Extraction method | PPT a | PPT a.1 | PPT a.2 | LLE b | LLE b.1 | LLE b.2 | SLE c | SLE c.1 | SLE c.2
Plasma conc. (ng/mL) | 6 | 160 | 1600 | 6 | 160 | 1600 | 6 | 160 | 1600
Extraction recovery (%) (n = 6) | 69 | 73 | 78 | 89 | 83 | 79 | 104 | 97 | 103
Overall mean recovery (%) | 73.3 | 73.3 | 73.3 | 83.6 | 83.6 | 83.6 | 101.3 | 101.3 | 101.3

## Table 3
Calibrator (ng/mL) | Unnamed: 1 | Back-calculated conc.(mean ± SD) (ng/mL) | Precision (CV%, n = 3) | Accuracy(% nominal)
2 |  | 2.06 ± 0.06 | 2.7 | 103
4 |  | 3.72 ± 0.24 | 6.3 | 93
10 |  | 12.5 ± 0.06 | 0.5 | 104.2
50 |  | 50.1± 0.7 | 1.4 | 100.2
200 |  | 197 ± 3.2 | 1.6 | 98.5
800 |  | 842 ± 11.7 | 1.4 | 105.3
1800 |  | 1740 ± 5.6 | 0.3 | 96.7
2000 |  | 2070 ±76 | 3.8 | 103.5

## Table 4
QC sample (ng/mL)_QC sample (ng/mL) | Intra-day_LLOQ 2.00 | Intra-day_LQC6.00 | Intra-day_MQC160 | Intra-day_HQC1,600 | Intra-day_DiQC5,000 | Inter-day_LLOQ 2.00 | Inter-day_LQC6.00 | Inter-day_MQC160 | Inter-day_HQC1,600
N | 6 | 6 | 6 | 6 | 6 | 18 | 18 | 18 | 17
Mean | 1.96 | 6.04 | 159 | 1630 | 2.03 | 6.19 | 6.19 | 160 | 1640
CV (%) | 8.4 | 5.9 | 3.9 | 3.6 | 2.8 | 1.5 | 3.2 | 4 | 2.5
Accuracy (%) | 98 | 100.7 | 99.4 | 101.9 | 106 | 101.5 | 103.2 | 100 | 102.5

## Table 5
Sample Concentration | Mean conc.found (ng/mL) | Precision(CV%, n = 6) | % nominal conc.
Plasma sample ambient storage (24 h) | Plasma sample ambient storage (24 h) | Plasma sample ambient storage (24 h) | Plasma sample ambient storage (24 h)
Low-QC: 6 ng/mL | 6.29 | 2.0 | 104.8
High-QC: 1,600 ng/mL | 1600 | 1.3 | 100.0
Freeze-thaw cycles (n = 3) | Freeze-thaw cycles (n = 3) | Freeze-thaw cycles (n = 3) | Freeze-thaw cycles (n = 3)
Low-QC: 6 ng/mL | 6.27 | 3.5 | 104.5
High-QC: 1,600 ng/mL | 1600 | 1.7 | 100.0
Extracted samples at autosampler viability (72 h) | Extracted samples at autosampler viability (72 h) | Extracted samples at autosampler viability (72 h) | Extracted samples at autosampler viability (72 h)
Low-QC: 6 ng/mL | 5.95 | 2.7 | 99.2
High-QC: 1,600 ng/mL | 1530 | 2.2 | 95.6
Long-term stability in plasma at -70°C for 227 days | Long-term stability in plasma at -70°C for 227 days | Long-term stability in plasma at -70°C for 227 days | Long-term stability in plasma at -70°C for 227 days
Low-QC: 6 ng/mL | 5.67 | 5.7 | 94.5
High-QC: 1,600 ng/mL | 1470 | 4.4 | 91.9
Long-term stability in plasma at -20°C for 227 days | Long-term stability in plasma at -20°C for 227 days | Long-term stability in plasma at -20°C for 227 days | Long-term stability in plasma at -20°C for 227 days
Low-QC: 6 ng/mL | 5.60 | 2.6 | 93.3
High-QC: 1,600 ng/mL | 1501 | 2.2 | 93.8

## Table 6
Plasma Timepoint_Plasma Timepoint | Erlotinib Concentration in Human Plasma (ng/mL)_24 h post-dose | Erlotinib Concentration in Human Plasma (ng/mL)_24 h post-dose.1 | Erlotinib Concentration in Human Plasma (ng/mL)_15 min post-dose | Erlotinib Concentration in Human Plasma (ng/mL)_15 min post-dose.1
| Plasma sample | Standard addition: 20.0 ng/mL | Plasma sample | Standard addition: 800 ng/mL
Individual measurement | 20.1 | 39.6 | 779 | 1590
Individual measurement | 19.8 | 39.8 | 798 | 1660
Individual measurement | 20.0 | 38.6 | 750 | 1630
Individual measurement | 20.1 | 42.0 | 824 | 1650
Individual measurement | 20.1 | 41.6 | 734 | 1640
Individual measurement | 22.2 | 42.8 | 708 | 1560
Mean | 20.4 | 40.7 | 766 | 1622
%CV (n = 6) | 3.0 | 3.4 | 5.6 | 2.4
% recovery of standard addition | - | 101.5 | - | 107